3GPP TSG-RAN WG4 Meeting #82	R4-1701864

# Athens, Greece, 13 - 17 February, 2017

| CR-Form-v11.1 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | TR 37.842 | CR | 0014 | rev |  | Current version: | 13.1.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network | x | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | CR to TS 37.842: Weightening coefficient clarification for the uncertainty budget calculations |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Huawei, Ericsson |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R4 |  |  |  |  |  |  |  |  |  |
| Work item code: | AAS_BS_LTE_UTRA-Core |  |  |  |  |  |  | Date: |  | 2017-01-27 |
| Category: | F |  |  |  |  |  |  | Release: |  | Rel-13 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) Rel-12 (Release 12) Rel-13 (Release 13) Rel-14 (Release 14) |  |
| Reason for change: |  | During RAN4#80bis and RAN4#81 meetings, discussion on the correction of the weightening coefficient was held. This CR proposes to clarify this issues based on the alignment with TS 34.114 specification, additionally reusing terminology from the referred specification (“Evaluation of measurement data — Guide to the expression of uncertainty in measurement”). |  |  |  |  |  |  |  |  |
| Summary of change: |  | - Reference to the guide for MU derivation added - Additional symbols for MU derivation and uncertainties - Updates to the uncertainty budget calculation principle - Uncertainty assessment tables corections with additional equations - Minor editorials |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | Unclear symbol will cause further confusions, lack of alignment among specifications. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | 2, 3.2, 10.1.2, 10.3.1.1.2.5, 10.3.1.1.3.5, 10.3.1.1.4.5, 10.3.1.1.5.4, 10.3.1.2, 10.3.2.1.1.4, 10.3.2.1.2.4, 10.3.2.1.3.4, 10.3.2.1.4.5, 10.3.2.2 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | x | Other core specifications |  |  |  |  |  |  |
| affected: |  |  | x | Test specifications |  |  |  |  |  |  |
| (show related CRs) |  |  | x | O&M Specifications |  |  |  |  |  |  |
| Other comments: |  |  |  |  |  |  |  |  |  |  |

------------------------------ Modified section ------------------------------

# 2	References

The following documents contain provisions which, through reference in this text, constitute provisions of the present document.

-	References are either specific (identified by date of publication, edition number, version number, etc.) or non-specific.

-	For a specific reference, subsequent revisions do not apply.

-	For a non-specific reference, the latest version applies. In the case of a reference to a 3GPP document (including a GSM document), a non-specific reference implicitly refers to the latest version of that document in the same Release as the present document.

[1]	3GPP TR 21.905: "Vocabulary for 3GPP Specifications".

[2]	3GPP TR 37.840: "Study of Radio Frequency (RF) and Electromagnetic Compatibility (EMC) requirements for Active Antenna Array System (AAS) base station".

[3]	RP-130373: "Base Station (BS) RF requirements for Active Antenna System (AAS)".

[4]	3GPP TR 25.942: "UTRA Radio Frequency (RF) system scenarios".

[5]	3GPP TR 36.942: "E-UTRA Radio Frequency (RF) system scenarios".

[6]	3GPP TS 25.104: "Base Station (BS) radio transmission and reception (FDD)".

[7]	3GPP TS 36.104: "E-UTRA Base Station (BS) radio transmission and reception".

[8]	3GPP TS 37.104: E-UTRA, UTRA and GSM/EDGE Multi-Standard Radio (MSR) Base Station (BS) radio transmission and reception".

[9]	3GPP TS 25.105: "Base Station (BS) radio transmission and reception (TDD)".

[10]	3GPP TS 25.141: "Base Station (BS) conformance testing (FDD)".

[11]	3GPP TS 25.142: "Base Station (BS) conformance testing (TDD)".

[12]	3GPP TS 36.141: "E-UTRA Base Station (BS) conformance testing".

[13]	ANSI/IEEE Std 149-1979: "IEEE Standard Test Procedures for Antennas".

[14]	3GPP TR 25.914: "Technical Specification 3rd Generation Partnership Project; Technical Specification Group Radio Access Networks; Measurements of Radio Performances for UMTS Terminals in Speech Mode".

[15]	ETSI TS 102 273-1-1: "Electromagnetic compatibility and Radio spectrum Matters (ERM); Improvement on Radiated Methods of Measurement (using test site) and evaluation of the corresponding measurement uncertainties; Part 1: Uncertainties in the measurement of mobile radio equipment characteristics; Sub-part 1: Introduction".

[16]	ETSI TS 102 273-1-2: "Electromagnetic compatibility and Radio spectrum Matters (ERM); Improvement on Radiated Methods of Measurement (using test site) and evaluation of the corresponding measurement uncertainties; Part 1: Uncertainties in the measurement of mobile radio equipment characteristics; Sub-part 2: Examples and annexes".

[17]	3GPP TS 37.145-2: "Active Antenna System (AAS) Base Station (BS) conformance testing; Part 2: radiated conformance testing".

[18]	3GPP TS 37.114: "Active Antenna System (AAS) Base Station (BS) Electromagnetic Compatibility (EMC)".

[19]	3GPP TS 37.113: "E-UTRA, UTRA and GSM/EDGE; Multi-Standard Radio (MSR) Base Station (BS) Electromagnetic Compatibility (EMC)".

[20]	3GPP TS 25.113: "Base Station (BS) and repeater ElectroMagnetic Compatibility (EMC)".

[21]	3GPP TS 36.113: "Evolved Universal Terrestrial Radio Access (E-UTRA); Base Station (BS) and repeater ElectroMagnetic Compatibility (EMC)".

[22]	3GPP TS 37.105: "Active Antenna System (AAS) Base Station (BS) transmission and reception (Release 13)".

[23]	Recommendation ITU-R SM.329: "Unwanted emissions in the spurious domain".

[24]	3GPP TS 37.145-1: "Active Antenna System (AAS) Base Station (BS) conformance testing; Part 1: conducted conformance testing".

[25]	JCGM 100:2008: “Evaluation of measurement data — Guide to the expression of uncertainty in measurement”.

## ------------------------------ Next modified section ------------------------------

## 3.2	Symbols

For the purposes of the present document, the following symbols apply:

Composite antenna array pattern in dB

Array element pattern in dB

BeW 	Beam width in 

BeWϕ	Beam width in ϕ

ci	Sensitivity coefficient

EFF		Far field pattern

Emeas	Measured near field

Fbasis	Basis functions (near field to far field transformation)

Array factor

uc	Combined standard uncertainty

ue	Expanded uncertainty

ui	Individual, standard uncertainty contributor

Azimuth angle (defined between -180° and 180°)

Elevation angle of the signal direction (defined between -90° and 90°, 0° represents the direction perpendicular to the antenna array)

σ 	Standard uncertainty

Signal correlation coefficient

	Angle in the reference coordinate system between the x-axis and the projection of the radiation vector onto the x/y plane defined between -180° and 180°

	Angle in the reference coordinate system between the projection of the x/y plane and the radiation vector defined between -90° and 90°. 0° represents the direction perpendicular to the y/z plane. The angle is aligned with the down-tilt angle

## ------------------------------ Next modified section ------------------------------

### 10.1.2	Uncertainty budget calculation principle

Uncertainty contributions listed in subclause 10.3 with descriptions in the annex B need to be calculated to provide an overall total measurement uncertainty for each OTA test method for the accompanying conformance requirement.

The uncertainty tables are presented with two stages in mind:

-	Stage 1: the calibration of the absolute level of the DUT measurement results is performed by means of using a calibration antenna (for example a standard gain horn) whose absolute gain is known at the frequencies of measurement

-	Stage 2: the actual measurement with the DUT as either the transmitter or receiver (depending on EIRP or EIS measurement) is performed.

The final uncertainty budget should comprise of a minimum 5 headings:

1)	The uncertainty source,

2)	Uncertainty value,

3)	Distribution of the probability,

4)	Divisor based on distribution shape,

5)	Sensitivity coefficient ciand its calculated standard uncertainty ui (based on uncertainty value, divisor and sensitivity coefficient ci).

NOTE:	All measurement uncertainty contributions are assumed independent, log and small, hence the errors induced by taking values in log are negligible.

The calculation of the uncertainty contribution is based on the ISO Guide [25] to the expression of uncertainty in measurement. Each individual uncertainty is expressed by its standard deviation (termed as ‘standard uncertainty’) and represented by symbol ui.

The procedure for forming the uncertainty budget can be as follows:

1)	Compile lists of individual uncertainty contributions for EIRP/EIS measurement both in Stage 1 and Stage 2.

2)	Determine the standard uncertainty of each contribution ui by:

a)	Determining the distribution of the uncertainty (i.e. Gaussian, U-shaped, rectangular, or exponentially normal).

b)	Determining the maximum value of each uncertainty (unless the distributions is Gaussian).

c)	Calculating the standard uncertainty by dividing the uncertainty by  if the distribution is U-shaped, by 2 if the distribution exponentially normal and by  if the distribution is rectangular.

d)	Multiplying the standard uncertainty by the sensitivity coefficient ci.

3)	Convert the units into decibel, if necessary.

4)	Combine all the standard uncertainties by the Root of the Sum of the Squares (RSS) method.

5)	Combine the uncertainties in Stage 1 and Stage 2 into the combined uncertainty uc also by the RSS method:

.

6)	Multiply the result of the combined uncertainty uc by an expansion factor of 1.96 to derive expanded uncertainty ue at 95 % confidence level:

ue = 1.96 uc

## ------------------------------ Next modified section ------------------------------

10.3.1.1.2.5	Uncertainty assessment

Table 10.3.1.1.2.5-1.1: Indoor Anechoic Chamber uncertainty assessment for EIRP measurement

| UID | Uncertainty source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz < f ≦ 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | Rectangular | √3 | 1 | 0.02 | 0.02 |
| 2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | Rectangular | √3 | 1 | 0.17 | 0.17 |
| 3 | Quality of quiet zone | 0.10 | 0.10 | Gaussian | 1 | 1 | 0.10 | 0.10 |
| 4 | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 5 | Mutual coupling between the AAS BS and the receiving antenna | 0.00 | 0.00 | Rectangular | √3 | 1 | 0.00 | 0.00 |
| 6 | Phase curvature | 0.05 | 0.05 | Gaussian | 1 |  | 0.05 | 0.05 |
| 7 | Uncertainty of the RF Power Measurement Equipment | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 8 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | U-shaped | √2 | 1 | 0.10 | 0.23 |
| 9 | Random uncertainty | 0.1 | 0.1 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 10 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | U-shaped | √2 | 1 | 0.04 | 0.04 |
| 11 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 12 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | U-shaped | √2 | 1 | 0.04 | 0.04 |
| 13 | Quality of quiet zone | 0.10 | 0.10 | Gaussian | 1 | 1 | 0.10 | 0.10 |
| 14 | Polarization mismatch for reference antenna | 0.01 | 0.01 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 15 | Mutual coupling between the reference antenna and the receiving antenna | 0.00 | 0.00 | Rectangular | √3 | 1 | 0.00 | 0.00 |
| 16 | Phase curvature | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 |
| 17 | Uncertainty of the network analyzer | 0.13 | 0.20 | Gaussian | 1 | 1 | 0.13 | 0.20 |
| 18 | Influence of the reference antenna feed cable | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 19 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 |
| 20 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 21 | Uncertainty of the absolute gain of the reference antenna | 0.5 | 0.43 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 22 | Uncertainty of the absolute gain of the receiving antenna | 0.00 | 0.00 | Rectangular | √3 | 1 | 0.00 | 0.00 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.44 | 0.54 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 0.87 | 1.06 |

## ------------------------------ Next modified section ------------------------------

10.3.1.1.3.5	Uncertainty assessment

Table 10.3.1.1.3.5-1: CATR uncertainty assessment for EIRP measurement

| EIRP uncertainty budget |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 2 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 3 | Standing wave between DUT and test range antenna | 0.21 | 0.21 | U-shaped | √2 | 1 | 0.15 | 0.15 |
| 4 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 5 | QZ ripple with DUT | 0.0928 | 0.0928 | Normal | 1 | 1 | 0.0928 | 0.0928 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 6 | Network Analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 7 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 8 | Insertion loss variation in receiver chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 9 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 10 | Influence of the calibration antenna feed cable | 0.022 | 0.022 | U-shaped | √2 | 1 | 0.015 | 0.015 |
| 11 | SGH Calibration uncertainty | 0.50 | 0.433 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 13 | Misalignment SGH and pointing error | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 |
| 14 | Rotary joints | 0.048 | 0.048 | U-shaped | √2 | 1 | 0.034 | 0.034 |
| 15 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.06 | 0.06 |
| 16 | QZ ripple with SGH | 0.009 | 0.009 | Normal | 1 | 1 | 0.009 | 0.009 |
| 17 | Switching uncertainty | 0.26 | 0.26 | Rectangular | √3 | 1 | 0.15 | 0.15 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.512 | 0.60 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 0.99 | 1.17 |

## ------------------------------ Next modified section ------------------------------

10.3.1.1.4.5	Uncertainty assessment

Table 10.3.1.1.4.5-1: One Dimensional Compact Range Chamber uncertainty assessment for EIRP measurement

| EIRP uncertainty budget |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz < f ≦ 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT and pointing error | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 2 | Standing wave between DUT and test range antenna | 0.18 | 0.18 | U-shaped | √2 | 1 | 0.13 | 0.13 |
| 3 | Quiet zone ripple DUT | 0.0325 | 0.0325 | Gaussian | 1 | 1 | 0.03 | 0.03 |
| 4 | Phase curvature | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 |
| 5 | Polarization mismatch between DUT and receiving antenna | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 6 | Mutual coupling between DUT and receiving antenna | 0 | 0 | Rectangular | √3 | 1 | 0 | 0 |
| 7 | Measurement equipment | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 8 | Impedance mismatch in receiving chain | 0.0056 | 0.01 | U-shaped | √2 | 1 | 0 | 0.01 |
| 9 | RF leakage (DUT connector terminated and test range antenna connector cable terminated) | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 10 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 11 | Pointing error between reference antenna and test range antenna | 0 | 0 | Rectangular | √3 | 1 | 0 | 0 |
| 12 | Impedance mismatch in path to reference antenna | 0.05 | 0.05 | U-shaped | √2 | 1 | 0.04 | 0.04 |
| 13 | Impedance mismatch in path to compact probe | 0.03 | 0.03 | U-shaped | √2 | 1 | 0.02 | 0.02 |
| 14 | Standing wave between reference antenna and receiving antenna | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.06 | 0.06 |
| 15 | Quiet zone ripple reference antenna | 0.178 | 0.178 | Gaussian | 1 | 1 | 0.18 | 0.18 |
| 16 | Phase curvature | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 |
| 17 | Polarization mismatch between reference antenna and receiving antenna | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 18 | Mutual coupling between reference antenna and receiving antenna | 0 | 0 | Rectangular | √3 | 1 | 0 | 0 |
| 19 | Measurement equipment | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 20 | Influence of the reference antenna feed cable (flexing cables, adapters, attenuators, connector repeatability) | 0.082 | 0.082 | Rectangular | √3 | 1 | 0.05 | 0.05 |
| 21 | Mismatch of receiver chain | 0.2 | 0.3 | U-shaped | √2 | 1 | 0.14 | 0.21 |
| 22 | Insertion loss of receiver chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.1 | 0.1 |
| 23 | Uncertainty of absolute gain of reference antenna | 0.5 | 0.43 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 24 | RF leakage (SGH connector terminated and test range antenna connector cable terminated. | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.46 | 0.56 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 0.9 | 1.09 |

## ------------------------------ Next modified section ------------------------------

10.3.1.1.5.4	Uncertainty assessment

Table 10.3.1.1.5.4-1: Near field test range uncertainty assessment for EIRP measurement

| UID | Uncertainty source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Axes Intersection | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 2 | Axes Orthogonality | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 3 | Horizontal Pointing | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 4 | Probe Vertical Position | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 5 | Probe H/V pointing | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 6 | Measurement Distance | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 7 | Amplitude and Phase Drift | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 8 | Amplitude and Phase Noise | 0.02 | 0.02 | Gaussian | 1.00 | 1 | 0.02 | 0.02 |
| 9 | Leakage and Crosstalk | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 10 | Amplitude Non-Linearity | 0.04 | 0.04 | Gaussian | 1.00 | 1 | 0.04 | 0.04 |
| 11 | Amplitude and Phase Shift in rotary joints | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 12 | Channel Balance Amplitude and Phase | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 13 | Probe Polarization Amplitude and Phase | 0.0001 | 0.0001 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 14 | Probe Pattern Knowledge | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 15 | Multiple Reflections | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 16 | Room Scattering | 0.09 | 0.09 | Gaussian | 1.00 | 1 | 0.09 | 0.09 |
| 17 | DUT support Scattering | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 18 | Scan Area Truncation | 0.003 | 0.003 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 19 | Sampling Point Offset | 0.0058 | 0.0058 | Gaussian | 1.00 | 1 | 0.01 | 0.01 |
| 20 | Spherical Mode Truncation | 0.015 | 0.015 | Gaussian | 1.00 | 1 | 0.02 | 0.02 |
| 21 | Positioning | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.02 | 0.02 |
| 22 | Probe Array Uniformity | 0.055 | 0.055 | Gaussian | 1.00 | 1 | 0.06 | 0.06 |
| 23 | Mismatch of receiver chain | 0.284 | 0.284 | U-Shaped | 1.41 | 1 | 0.20 | 0.20 |
| 24 | Insertion loss of receiver chain | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 25 | Uncertainty of the absolute gain of the probe antenna | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 26 | Measurement Receiver | 0.14 | 0.26 | Gaussian | 1.00 | 1 | 0.14 | 0.26 |
| 27 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | Gaussian | 1.00 | 1 | 0.15 | 0.15 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 28 | Network analyzer | 0.13 | 0.20 | Gaussian | 1.00 | 1 | 0.13 | 0.20 |
| 29 | Mismatch of receiver chain | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 30 | Insertion loss of receiver chain | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 31 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | U-Shaped | 1.41 | 1 | 0.01 | 0.01 |
| 32 | Influence of the calibration antenna feed cable | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 33 | Influence of the probe antenna cable | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 34 | Reference antenna | 0.50 | 0.25 | Rectangular | 1.73 | 1 | 0.29 | 0.25 |
| 35 | Short term repeatability | 0.088 | 0.088 | Gaussian | 1.00 | 1 | 0.09 | 0.09 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.45 | 0.50 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 0.89 | 0.99 |

NOTE:	Standard uncertainty values for the signal generator, network analyzer and reference antenna are according to the test equipment vendors proposal, as captured in annex E.

## ------------------------------ Next modified section ------------------------------

#### 10.3.1.2	Maximum accepted test system uncertainty

Maximum test system uncertainties derivation methodology was described in subclause 10.1.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for the EIRP test can be derived from values captured in table 10.3.1.2-1, separately for each of the defined frequency ranges. The common maximum values are applicable for all test methods addressing EIRP test requirement. Based on the input values in table 10.3.2.2-1, the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values are derived for two frequency ranges (i.e. f ≦ 3GHz and 3GHz < f ≦ 4.2 GHz).

Table 10.3.1.2-1: Test system specific measurement uncertainty values for the EIRP test

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 0.87 | 1.06 |
| Compact Antenna Test Range | 1.00 | 1.17 |
| One Dimensional Compact Range Chamber | 0.9 | 1.09 |
| Near Field Test Range | 0.89 | 0.99 |
| Common maximum accepted test system uncertainty | 1.0 | 1.2 |

## ------------------------------ Next modified section ------------------------------

10.3.2.1.1.4	Uncertainty assessment

Table 10.3.2.1.1.4-1: Indoor Anechoic Chamber uncertainty assessment for EIS measurement

| UID | Uncertainty source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.02 | 0.02 |
| 2 | Pointing misalignment between the AAS BS and the transmitting antenna | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 | 0.17 |
| 3 | Quality of quiet zone | 0.1 | 0.1 | Gaussian | 1.00 | 1 | 0.10 | 0.10 |
| 4 | Polarization mismatch between the AAS BS and the transmitting antenna | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 |
| 5 | Mutual coupling between the AAS BS and the transmitting antenna | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 6 | Phase curvature | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 | 0.05 |
| 7 | Uncertainty of the RF signal generator | 0.46 | 0.46 | Gaussian | 1.00 | 1 | 0.46 | 0.46 |
| 8 | Impedance mismatch in the transmitting chain | 0.14 | 0.23 | U-shaped | 1.41 | 1 | 0.10 | 0.16 |
| 9 | Random uncertainty | 0.1 | 0.1 | Rectangular | 1.73 | 1 | 0.06 | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 10 | Impedance mismatch between the transmitting antenna and the network analyzer | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 | 0.04 |
| 11 | Positioning and pointing misalignment between the reference antenna and the transmitting antenna | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 |
| 12 | Impedance mismatch between the reference antenna and network analyzer | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 | 0.04 |
| 13 | Quality of quiet zone | 0.1 | 0.1 | Gaussian | 1.00 | 1 | 0.10 | 0.10 |
| 14 | Polarization mismatch for reference antenna | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 |
| 15 | Mutual coupling between the reference antenna and the transmitting antenna | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 16 | Phase curvature | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 | 0.05 |
| 17 | Uncertainty of the network analyzer | 0.13 | 0.20 | Gaussian | 1.00 | 1 | 0.13 | 0.20 |
| 18 | Influence of the reference antenna feed cable | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 |
| 19 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | Gaussian | 1.00 | 1 | 0.06 | 0.06 |
| 20 | Influence of the transmitting antenna feed cable | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 |
| 21 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.25 | Rectangular | 1.73 | 1 | 0.29 | 0.25 |
| 22 | Uncertainty of the absolute gain of the transmitting antenna | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.62 | 0.64 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.22 | 1.25 |

## ------------------------------ Next modified section ------------------------------

10.3.2.1.2.4	Uncertainty assessment

Table 10.3.2.1.2.4-1: CATR uncertainty assessment for EIS measurement

| EIS uncertainty budget |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz |  | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  |
| 2 | Standing wave between DUT and test range antenna | 0.21 | 0.21 | U-shaped | √2 | 1 | 0.15 | 0.15 |  |
| 3 | RF signal generator | 0.46 | 0.46 | Gaussian | 1 | 1 | 0.46 | 0.46 |  |
| 4 | RF leakage & dynamic range, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |  |
| 5 | QZ ripple with DUT | 0.0928 | 0.0928 | Normal | 1 | 1 | 0.0928 | 0.0928 |  |
| 6 | Miscellaneous uncertainty | 0 | 0 | Normal | 1 | 1 | 0 | 0 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |
| 7 | Network Analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |  |
| 8 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |  |
| 9 | Insertion loss of transmitter chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |  |
| 10 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.00156 | 0.00156 | Normal | 1 | 1 | 0.00156 | 0.00156 |  |
| 11 | Flexing cables & connector repeatability | 0.022 | 0.022 | U-shaped | √2 | 1 | 0.015 | 0.015 |  |
| 12 | SGH Calibration uncertainty | 0.502 | 0.433 | Rectangular | √3 | 1 | 0.29 | 0.25 |  |
| 13 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  |
| 14 | QZ ripple with SGH | 0.09 | 0.09 | Normal | 1 | 1 | 0.09 | 0.09 |  |
| 15 | Rotary joints | 0.048 | 0.048 | U-shaped | √2 | 1 | 0.034 | 0.034 |  |
| 16 | Misalignment SGH and pointing error | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 |  |
| 17 | Misalignment calibration system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  |
| 18 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.06 | 0.06 |  |
| 19 | Switching uncertainty | 0.26 | 0.26 | Rectangular | √3 | 1 | 0.15 | 0.15 |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.68 | 0.71 |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.33 | 1.40 |  |

## ------------------------------ Next modified section ------------------------------

10.3.2.1.3.4 	Uncertainty assessment

Table 10.3.2.1.3.4-1: One Dimensional Compact Range Chamber uncertainty assessment for EIS measurement

| EIS uncertainty budget |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz < f ≦ 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz |  | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT and pointing error | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  |
| 2 | Standing wave between DUT and test range antenna | 0.18 | 0.18 | U-shaped | √2 | 1 | 0.13 | 0.13 |  |
| 3 | Quiet zone ripple DUT | 0.0325 | 0.0325 | Gaussian | 1 |  | 0.03 | 0.03 |  |
| 4 | Signal generator | 0.46 | 0.46 | Gaussian | 1 | 1 | 0.46 | 0.46 |  |
| 5 | Phase curvature | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 |  |
| 6 | Polarization mismatch between DUT and transmitting antenna | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |  |
| 7 | Mutual coupling between DUT and transmitting antenna | 0 | 0 | Rectangular | √3 | 1 | 0 | 0 |  |
| 8 | Measurement equipment | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |  |
| 9 | Impedance mismatch in transmitter chain | 0.0056 | 0.01 | U-shaped | √2 | 1 | 0 | 0.01 |  |
| 10 | RF leakage and dynamic range a) feeder package terminated b) DUT terminated | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |
| 11 | Misalignment positioning system | 0 | 0 | Exp. Normal | 2 | 1 | 0 | 0 |  |
| 12 | Pointing error between reference antenna and test range antenna | 0 | 0 | Rectangular | √3 | 1 | 0 | 0 |  |
| 13 | Impedance mismatch in path to reference antenna | 0.05 | 0.05 | U-shaped | √2 | 1 | 0.04 | 0.04 |  |
| 14 | Impedance mismatch in path to compact probe | 0.03 | 0.03 | U-shaped | √2 | 1 | 0.02 | 0.02 |  |
| 15 | Standing wave between reference antenna and test range antenna | 0.15 | 0.15 | U-shaped | √2 | 1 | 0.11 | 0.11 |  |
| 16 | Quiet zone ripple reference antenna | 0.178 | 0.178 | Gaussian | 1 | 1 | 0.18 | 0.18 |  |
| 17 | Phase curvature | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 |  |
| 18 | Polarization mismatch between reference antenna and transmitting antenna | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |  |
| 19 | Mutual coupling between reference antenna and transmitting antenna | 0 | 0 | Rectangular | √3 | 1 | 0 | 0 |  |
| 20 | Measurement equipment | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |  |
| 21 | Influence of reference antenna feed cable (flexing cables, adapters, attenuators, connector repeatability) | 0.082 | 0.082 | Rectangular | √3 | 1 | 0.05 | 0.05 |  |
| 22 | Mismatch of transmitter chain | 0.2 | 0.3 | U-shaped | √2 | 1 | 0.14 | 0.21 |  |
| 23 | Insertion loss of transmitter chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.1 | 0.1 |  |
| 24 | Uncertainty of absolute gain of reference antenna | 0.5 | 0.43 | Rectangular | √3 | 1 | 0.29 | 0.25 |  |
| 25 | RF leakage (SGH connector terminated and test range antenna connector cable terminated) | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.66 | 0.73 |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.29 | 1.43 |  |

## ------------------------------ Next modified section ------------------------------

10.3.2.1.4.5	Uncertainty assessment

Table 10.3.2.1.4.5-1: Near field test range uncertainty assessment for EIS measurement

| UID | Uncertainty source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Axes Intersection | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 2 | Axes Orthogonality | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 3 | Horizontal Pointing | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 4 | Probe Vertical Position | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 5 | Probe H/V pointing | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 6 | Measurement Distance | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 7 | Amplitude and Phase Drift | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 8 | Amplitude and Phase Noise | 0.02 | 0.02 | Gaussian | 1.00 | 1 | 0.02 | 0.02 |
| 9 | Leakage and Crosstalk | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 10 | Amplitude Non-Linearity | 0.04 | 0.04 | Gaussian | 1.00 | 1 | 0.04 | 0.04 |
| 11 | Amplitude and Phase Shift in rotary joints | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 12 | Channel Balance Amplitude and Phase | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 13 | Probe Polarization Amplitude and Phase | 0.0001 | 0.0001 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 14 | Probe Pattern Knowledge | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 15 | Multiple Reflections | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 16 | Room Scattering | 0.09 | 0.09 | Gaussian | 1.00 | 1 | 0.09 | 0.09 |
| 17 | DUT support Scattering | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 18 | Scan Area Truncation | 0.003 | 0.003 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 19 | Sampling Point Offset | 0.0058 | 0.0058 | Gaussian | 1.00 | 1 | 0.01 | 0.01 |
| 20 | Spherical Mode Truncation | 0.015 | 0.015 | Gaussian | 1.00 | 1 | 0.02 | 0.02 |
| 21 | Positioning | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.02 | 0.02 |
| 22 | Probe Array Uniformity | 0.055 | 0.055 | Gaussian | 1.00 | 1 | 0.06 | 0.06 |
| 23 | Mismatch of transmitter chain | 0.284 | 0.284 | U-Shaped | 1.41 | 1 | 0.20 | 0.20 |
| 24 | Insertion loss of transmitter chain | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 25 | Uncertainty of the absolute gain of the probe antenna | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 26 | Vector Signal Generator | 0.46 | 0.46 | Gaussian | 1.00 | 1 | 0.46 | 0.46 |
| 27 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | Gaussian | 1.00 | 1 | 0.15 | 0.15 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 28 | Uncertainty of network analyser | 0.13 | 0.20 | Gaussian | 1.00 | 1 | 0.13 | 0.20 |
| 29 | Mismatch of transmitter chain | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 30 | Insertion loss of transmitter chain | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 31 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | U-Shaped | 1.41 | 1 | 0.01 | 0.01 |
| 32 | Influence of the calibration antenna feed cable | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 33 | Influence of the probe antenna cable | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 34 | Reference antenna | 0.50 | 0.25 | Rectangular | 1.73 | 1 | 0.29 | 0.25 |
| 35 | Short term repeatability | 0.088 | 0.088 | Gaussian | 1.00 | 1 | 0.09 | 0.09 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.63 | 0.63 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.24 | 1.24 |

NOTE:	Standard uncertainty values for the signal generator, network analyzer and reference antenna are according to the test equipment vendors proposal, as captured in annex E.

## ------------------------------ Next modified section ------------------------------

#### 10.3.2.2	Maximum accepted test system uncertainty

Maximum test system uncertainties derivation methodology was described in subclause 10.1.1. The maximum accepted test system uncertainty values was derived based on test system specific values, separately for EIRP and for EIS tests.

According to the methodology referred above, the common maximum accepted test system uncertainty values derivation for EIS are captured in table 10.3.2.2-2. According to the methodology in subclause 10.1.1, the common maximum accepted test system uncertainty values for the EIS test can be derived from table 10.3.2.2-2, for each of the defined frequency ranges. The common maximum values are applicable for all test methods addressing EIS test requirement. Based on the input values in table 10.3.2.2-2, the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values are derived for two frequency ranges (i.e. f ≦ 3GHz and 3GHz < f ≦ 4.2 GHz).

Table 10.3.2.2-2 Test system specific measurement uncertainty values for the EIS test

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 1.22 | 1.25 |
| Compact Antenna Test Range | 1.33 | 1.40 |
| One Dimensional Compact Range Chamber | 1.29 | 1.43 |
| Near Field Test Range | 1.24 | 1.24 |
| Common maximum accepted test system uncertainty | 1.3 | 1.4 |

## ------------------------------ End of modified sections ------------------------------