3GPP TSG-RAN WG4 Meeting # 103-e	R4-2208661
Electronic Meeting, May 09 – May 20, 2022

Agenda item:	9.2.2

Source:	SRTC, Bluetest

Title:	TP to TR 38.834: addition of RC in test methodology

Document for:	Approval

1	Introduction

In this contribution, we propose the text descriptions of RC based methodology to [1] aligning with [2-3]. The Background for this addition is introduced in [4].

2	Conclusion

The following proposal is made in this contribution:

Proposal 1: Approve the text proposal below.

3	References

- 3GPP TR 38.834 V17.0.0  (2022-03)

- 3GPP TS 37.544 V16.2.0 (2021-09)

- 3GPP TR 37.902 V16.0.0 (2020-06)

- R4-2208659, “Discussion on the addition of RC in test methodology,” SRTC, Bluetest, 3GPP RAN4#103-bis-e May 2022

4	Appendix –Text Proposal to TS 38.834 v17.0.0

--------------Start of text proposal 1-------------

3.3	Abbreviations

For the purposes of the present document, the abbreviations given in 3GPP TR 21.905 [1] and the following apply. An abbreviation defined in the present document takes precedence over the definition of the same abbreviation, if any, in 3GPP TR 21.905 [1].

Abbreviation format (EW)

DUT	Device Under Test

EIRP 	Effective Isotropic Radiated Power

EUT	Equipment Under Test

FR1	Frequency Range 1

FS	Free Space

LEE	Laptop Embedded Equipment

LME	Laptop Mounted Equipment

NSA	Non-Standalone, a mode of operation where operation of an other radio is assisted with an other radio

OTA	Over The Air

QZ	Quiet Zone

SA	Standalone

SS	System Simulator

TRP	Total Radiated Power

TRS	Total Radiated Sensitivity

UE	User Equipment

RC	Reverberation Chamber

--------------End of text proposal 1-------------

--------------Start of text proposal 2-------------

5	Performance metrics

Transmitter power measurements shall be performed using the Total Radiated Power (TRP) as the measurement metric.

This definition will be used to calculate the Total Radiated Power (TRP) value of NR FR1 DUT.

The TRP with Anechoic Chamber method is defined as:

TRP= 14πθ=0πϕ=02πEIRPθθ,ϕ+EIRPϕθ,ϕsinθdϕdθ		(5.1)

Where the effective isotropic radiated power (EIRP) is defined as

EIRPθ,ϕ=PTGTθ,ϕ		   	           	 (5.2)

Where PTGT is the product of the power delivered to the antenna and the antenna’s power gain.

Where EIRPθ and EIRPϕ are the EIRP in the corresponding θ and ϕ polarizations.

The summation form based on the sin weights of TRP with Anechoic Chamber method is defined as:

TRP≈ π2NMn=0N-1m=0M-1EIRPθθn,ϕm+EIRPϕθn,ϕmsinθn            (5.3)

Where N and M are the number of sampling intervals for θ and ϕ. θn and ϕm are the measurement angles.

The summation form based on the Clenshaw-Curtis quadrature integral approximation of TRP with Anechoic Chamber method is defined as:

TRP≈ 12Mn=0Nm=0M-1EIRPθθn,ϕm+EIRPϕθn,ϕmWθn            (5.4)

Where the value of W(θn) follows Table 5.1-1.

Table 5.1-1 Weights for Clenshaw-Curtis Quadrature with 15o

| Clenshaw-Curtis |  |
| --- | --- |
|  [deg] | Weights |
| 0 | 0.007 |
| 15 | 0.0661 |
| 30 | 0.1315 |
| 45 | 0.1848 |
| 60 | 0.227 |
| 75 | 0.2527 |
| 90 | 0.262 |
| 105 | 0.2527 |
| 120 | 0.227 |
| 135 | 0.1848 |
| 150 | 0.1315 |
| 165 | 0.0661 |
| 180 | 0.007 |

The TRP can also be calculated from Rayleigh faded samples of the total power transmitted from the UE. The measurement of transmitter performance in an isotropic Rayleigh fading environment is based on sampling the radiated power of the UE for a discrete number of field combinations in the chamber. The average value of these statistically distributed samples is proportional to the TRP and by calibrating the average power transfer function, an absolute value of the TRP can be obtained. Thus

(5.3)

where  is the reference power transfer function for fixed measurement antenna n,  is the reflection coefficient for fixed measurement antenna n and  is the path loss in the cables connecting the measurement receiver to fixed measurement antenna n. These parameters are calculated from the calibration measurement and are further discussed in Clause 7.2.  is the average power measured by fixed measurement antenna n and can be calculated using the following expression:

(5.4)

where  is sample number m of the complex transfer function measured with fixed measurement antenna n and  is the total number of samples measured for each fixed measurement antenna.

Note that all averaging must be performed using linear power values (e.g. measurements in Watts).

5.2	Definition of Total Radiated Sensitivity (TRS)

Receiver sensitivity measurements shall be performed using data throughput as the measurement metric. The DUT’s receiver sensitivity corresponds to the minimum downlink signal power required to provide a data throughput rate greater than or equal to 95% of the maximum throughput of the reference measurement channel (RMC).

This definition will be used to calculate the Total Radiated Sensitivity (TRS) value of NR FR1 DUT.

The TRS with Anechoic Chamber method is defined as:

TRS=4πθ=0πϕ=02π1EISθθ,ϕ+1EISϕθ,ϕsinθdϕdθ     	   (5.5)

Where the effective isotropic sensitivity (EIS) is defined as the minimum power level at which the throughput exceeds or equal to 95% of the maximum throughput of the specified RMC, at each sampling point.

Where EISθ and EISϕ are the EIS in the corresponding θ and ϕ polarizations.

The summation form based on the sin weights of TRS with Anechoic Chamber method defined as:

TRS≈2NMπn=0N-1m=0M-11EISθθn,ϕm+1EISϕθn,ϕmsinθn                  (5.6)

Where N and M are the number of sampling intervals for θ and ϕ. θn and ϕm are the measurement angles.

The summation form based on the Clenshaw-Curtis quadrature integral approximation of TRS with Anechoic Chamber method is defined as:

TRS≈2Mn=0Nm=0M-11EISθθn,ϕm+1EISϕθn,ϕmWθn                  (5.7)

Where the value of W(θn) follows Table 5.2-1.

Table 5.2-1 Weights for Clenshaw-Curtis Quadrature with 30o

| Clenshaw-Curtis |  |
| --- | --- |
|  [deg] | Weights |
| 0 | 0.007 |
| 30 | 0.1315 |
| 60 | 0.227 |
| 90 | 0.262 |
| 120 | 0.227 |
| 150 | 0.1315 |
| 180 | 0.007 |

The TRS can also be calculated from measurements in a Rayleigh fading 3 dimensional isotropic environment with in average uniform elevation and azimuth distribution. The calculation of the TRS is in this case based on searching for the lowest power received by the UE/MS for a discrete number of field combinations in the chamber that gives a BER that is better than the specified target BER level. By calibrating the average power transfer function, an absolute value of the TRS can be obtained. The following expression can be used to find the TRS.

(5.6)

where  is the reference power transfer function for fixed measurement antenna n,  is the reflection coefficient for fixed measurement antenna n and  is the path loss in the cables connecting the measurement receiver to fixed measurement antenna n. These parameters are calculated from the calibration measurement and are further discussed in Clause 7.2.  is calculated by using the following equation:

(5.7)

where  is the m:th value of the transfer function for fixed measurement antenna n, which gives the BER threshold.  is the total number of values of the BER threshold power measured for each fixed measurement antenna.

--------------End of text proposal 2-------------

--------------Start of text proposal 3-------------

# 7	Test setup and calibration

7.1		General

7.1.1	Minimum Range Length for anechoic chamber method

This sub-section specifies the minimum range lengths for NR FR1 TRP-TRS OTA systems. The range length is defined as the distance from the centre of the quiet zone to the aperture of the measurement probes/antennas, as illustrated in Figure 7.1.1-1.

Figure 7.1.1-1: Illustration of range length definition

The minimum range length shall be the maximum of the following three limits

-	The phase uncertainty limit: RQZ+2Drad2/

-	The amplitude uncertainty limit: 3D

-	The reactive Near-Field limit: RQZ+2

where RQZ is defined as the radius of the quiet zone, i.e., RQZ=D/2, and Drad is the diameter of the effective radiating aperture. The minimum range length calculations for D=30cm quiet zone size TRP-TRS OTA test systems shall assume that Drad is 30cm below 1GHz and decrease linearly from 30cm to 5cm from 1GHz to 7.125GHz, respectively. The last column of Table 7.1.1-1 shall be considered the minimum range length for NR FR1 TRP-TRS OTA systems with 30cm quiet zone size.

Table 7.1.1-1: Minimum Range Length for NR FR1 TRP-TRS OTA systems with 30cm quiet zone size.

| F [GHz] | Drad [m] | RQZ+2Drad²/λ | 3D = 6RQZ | RQZ+2λ | max(RQZ+2λ,3D,RQZ+2D²/λ) |
| --- | --- | --- | --- | --- | --- |
| 0.41 | 0.30 | 0.40 | 0.9 | 1.61 | 1.61 |
| 0.6 | 0.30 | 0.51 | 0.9 | 1.15 | 1.15 |
| 0.7 | 0.30 | 0.57 | 0.9 | 1.01 | 1.01 |
| 0.8 | 0.30 | 0.63 | 0.9 | 0.90 | 0.90 |
| 1 | 0.30 | 0.75 | 0.9 | 0.75 | 0.90 |
| 1.2 | 0.29 | 0.83 | 0.9 | 0.65 | 0.90 |
| 1.4 | 0.28 | 0.90 | 0.9 | 0.58 | 0.90 |
| 1.6 | 0.28 | 0.96 | 0.9 | 0.52 | 0.96 |
| 1.8 | 0.27 | 1.01 | 0.9 | 0.48 | 1.01 |
| 2 | 0.26 | 1.05 | 0.9 | 0.45 | 1.05 |
| 2.2 | 0.25 | 1.07 | 0.9 | 0.42 | 1.07 |
| 2.4 | 0.24 | 1.09 | 0.9 | 0.40 | 1.09 |
| 2.6 | 0.23 | 1.11 | 0.9 | 0.38 | 1.11 |
| 2.8 | 0.23 | 1.11 | 0.9 | 0.36 | 1.11 |
| 3 | 0.22 | 1.10 | 0.9 | 0.35 | 1.10 |
| 4 | 0.18 | 0.99 | 0.9 | 0.30 | 0.99 |
| 5 | 0.14 | 0.77 | 0.9 | 0.27 | 0.90 |
| 6 | 0.10 | 0.52 | 0.9 | 0.25 | 0.90 |
| 7 | 0.06 | 0.29 | 0.9 | 0.24 | 0.90 |
| 7.125 | 0.05 | 0.27 | 0.9 | 0.23 | 0.90 |

7.1.2	Test volume for Reverberation chamber

The reverberation chamber shall be large enough to contain the largest EUT to be tested, when the EUT is physically located within the valid test volume. No EUT and supporting equipment (including RF absorber used to load the chamber, RF absorber, and test fixtures) shall occupy more than 8% of the chamber volume. The valid test volume is the volume for which uncertainties have been established that satisfied, and that in other regards is suitable to hold a device under test. Valid test volume boundaries shall be located at least 0.5λ at the frequency of operation from the walls, mechanical mode-stirrers (if used), and the antennas. The test volume shall encompass both the reference antenna and the EUT. The chamber must support a mode-stirring sequence that shall be chosen such that the measurement uncertainty is below the threshold value specified.

## 7.2	Test setup

For FR1 TRP TRS testing, both Single-antenna and multiple-antennas anechoic chambers can both be applied. In Figure 7.2-1, an example TRP TRS test system with combined axes system is presented.

Figure 7.2-1: Example of a FR1 TRP TRS anechoic chamber OTA test system with combined axis

In Figure 7.2-2, an example TRP TRS test system with distributed axes system is presented.

Figure 7.2-2: Example of a FR1 TRP TRS anechoic chamber OTA test system with combined axis

For the reverberation chamber method, a schematic picture of the test setup is provided in Figure 7.2-3.

Figure 7.2-3: Example of a FR1 TRP TRS reverberation chamber OTA Test System

7.3	Calibration procedure

7.3.1	Calibration procedure for anechoic chamber method

The relative power values of the measurement points will be transformed to absolute radiated power values (in dBm) by performing a range path loss calibration measurement. The system needs to be calibrated by using a reference calibration antenna with known gain values. In the range path loss calibration measurement, the reference antenna is measured in the same place as the DUT, i.e. the center of the QZ, and the attenuation of the complete transmission path (Ltotal) from the DUT to the measurement receiver/BS simulator is calibrated out.

Figure 7.3-1: Example FR1 TRP TRS anechoic chamber calibration setup

The calibration measurement is repeated for each measurement path (two orthogonal polarizations and each signal path). The range path loss calibration measurement is performed in a two-step process including total path loss measurement and cable calibration.

Step 1: Cable calibration: the measurement of path loss LDE LDE, by connecting the cable from D to E to the two ports of VNA, and measure the cable path loss.

Step 2: Total path loss measurement: the measurement of total path loss LBC;

1. Place the reference calibration antenna (e.g. reference dipole) in the center of the test zone aligned with θ polarization of the measurement antenna, connected to a VNA port E, with the other VNA port C connected to the input of the Switch box – in Figure 2.

2. Configure the proper output power of VNA.

3. Measure the response LCE LCE of each path from each θ polarization of the measurement antenna to the reference antenna in the center of QZ.

4. Repeat the steps 1 to 3 with the reference antenna aligned with the ϕ polarization of the measurement antenna.

Then, the  Ltotal=LCE-LDE+Gcal, Where LDE LDE is cable loss from D to E. Gcal is the gain or efficiency of the calibration antenna at the frequency of interest. In TRP and TRS measurements point C is connected to the calibrated input/output port of measurement receiver.

This range path loss calibration procedure is common to both SA and EN-DC measurements.

7.3.2	Calibration procedure for reverberation chamber method

For the reverberation chamber method, all polarizations and transmission paths are included in one calibration measurement. The calibration antenna can be place in an arbitrarily position, as long as it is placed 0.5 wavelengths from other metallic objects and 0.7 wavelengths from absorbing objects. An illustration of the measurement setup for this procedure is shown in Figure 7.3-2.

The purpose of the calibration measurement is to determine the average power transfer function in the chamber, mismatch of fixed measurement antennas and path losses in cables connecting the power sampling instrument and the fixed measurement antennas. Preferably a network analyzer is used for these measurements. Recommended calibration antennas are dipoles tuned to the frequency band of interest.

In general, the calibration of a reverberation chamber is performed in three steps:

  - 1.	Measurement of S-parameters through the reverberation chamber for a complete stirring sequence

- 2.	Calculation of the chamber reference transfer function

- 3.	Measurement of connecting cable insertion loss

If several setups are used (e.g. empty chamber, chamber with head or hand phantom, etc.), steps 1 and 2 must be repeated for each configuration. The calibration measurement setup is shown in Figure 7.3-2.

Figure 7.3-2: Example FR1 TRP TRS reverberation chamber calibration setup

#### 7.3.2.1	Measurement of S-parameters through the chamber for a complete stirring sequence

- This step will measure S-parameters through the reverberation chamber through a complete stirring sequence. This information is required to determine the chamber’s reference transfer function. The procedure must be performed separately for each measurement setup of which the loading of the chamber has been changed. The calibration procedure must be repeated for each frequency as defined above. Therefore, it is advantageous if the network analyzer can be set to a frequency sweep covering the defined frequencies, so that all frequencies of interest can be measured with a minimal number of measurement runs.

- i.	Place all objects into the RC which will be used during TRP or TRS measurements, including a head phantom, hand phantom and fixture for the EUT. This ensures that the loss in the chamber, which determines the average power transfer level, is the same during both calibration and test measurements. Also, if the EUT is large or contains many antennas, it may represent a noticeable loading of the chamber. It should then be present in the chamber and turned on during the calibration.

- ii.	Place the calibration antenna inside the chamber. The calibration antenna is preferably mounted on a low-loss dielectric fixture, to avoid effects from the fixture itself which may affect the EUT’s radiation efficiency and mismatch factor. The calibration antenna must be placed in the chamber in such a way that it is far enough from any walls, mode-stirrers, head phantom, hand phantom, or other object, such that the environment for the calibration antenna (taken over the complete stirring sequence) resembles a free space environment. “Far enough away” depends on the type of calibration antenna used. For low gain nearly omni-directional antennas like dipoles, it is normally sufficient to ensure that this spacing is larger than 0.5 wavelengths. More directive calibration antennas should be situated towards the centre of the chamber. The calibration antenna should be present in the chamber during the TRP/TRS measurements.

- iii.	Calibrate the network analyzer with a full 2-port calibration in such a way that the vector S-parameters between the ports of the fixed measurement antenna and the calibration antenna can be accurately measured. Preferably, the network analyzer is set to perform a frequency sweep at each stirrer position. This will enable calibration of several frequency points during the same stirring sequence, thereby reducing calibration time. This will also enable frequency stirring, i.e., averaging the measured power transfer function over a small frequency bandwidth around each measured frequency point (moving frequency window). This will increase accuracy at the expense of frequency resolution.

- iv.	Connect the antennas and measure the S-parameters for each stirrer position and each fixed measurement antenna.

- The number of stirrer positions in the chosen stirring sequence, i.e. the number of S-parameter samples at each frequency point, should be chosen in such a way that it is large enough to yield an acceptable statistical contribution to the total measurement uncertainty. As a guideline it should be larger than 100, preferable 200 or 400 to ensure that the number of independent samples is not severely limited by the total number of samples measured. The number of independent samples, which is a subset of all samples, determines the statistical contribution to the expanded accuracy (which is two times the standard deviation). This should be not less than 100 to ensure an expanded accuracy better than 0.5 dB. The number of independent samples depends on the operating frequency, volume of the chamber, efficacy of the chamber’s stirrers, the level of loading by absorbing objects, and whether or not frequency stirring is used.

- The sequence of moving the stirrers to different positions may be either step-wise (stopping stirrer for each sample) or continuous (sampling on-the-fly). With continuous stirring it may not be possible to characterize the chamber over a wide frequency band at the same time.

7.3.2.2	Calculation of the chamber reference transfer function

From the S-parameters obtained in the calibration measurement, the chamber reference transfer function for fixed antenna n can be calculated. The reflection coefficient for fixed antenna n can be calculated by the following equation,

(7.1)

Thus, the chamber reference transfer function can be calculated by the following equation,

(7.2)

where M is the total number of samples of the transfer function measured for each fixed measurement antenna and     is sample number m of the transfer function for measurement antenna n. Moreover,  is the complex average of the calibration antenna reflection coefficient. Finally,  is the radiation efficiency of the calibration antenna.

- NOTE:	The radiation efficiency of the fixed antenna is not corrected for because it will be the same both during calibration and measurements. Therefore, the fixed antenna’s radiation efficiency will not affect the final results. The same can be said about the mismatch factor of the fixed measurement antennas, but it is still advantageous to correct for this factor if frequency stirring is applied to improve accuracy.

7.3.2.3	Cable calibration

This measurement step will calibrate the power loss of the cable needed to connect the instrument used to measure the received power at the fixed measurement antenna during TRP measurements, and to generate the power radiated by the fixed antenna during TRS measurements.

i.	Disconnect the cables between the VNA and the chamber.

ii.	Connect the cables one-by-one between the two ports of the network analyzer. The VNA must be calibrated at its own two ports.

iii.	Measure the frequency response of the transmission S-parameter ( or ) of the cable.

iv.	Save the power transfer values () of the frequency response curve for the test frequencies and cables positions, etc).

Calibration shall be performed yearly or if any equipment in the measurement system is changed.

--------------End of text proposal  3-------------

--------------Start of text proposal 4-------------

8	SA test methodology

8.1	General

For UE radiated conformance testing P-MPRc shall be 0 dB.

8.2	Total Radiated Power (TRP)

8.2.1	Test Conditions

For the anechoic chamber based measurement procedures the TRP of the DUT is measured by sampling the radiated transmit power of the DUT with three-dimensional scan at various locations surrounding the device. The measurement is performed with a constant sampling step of 15 degrees in both theta () and phi () axes for TRP measurement. This accounts for a total of 266 measurements for each of two orthogonal polarizations since measurements at theta = 0 and 180 degrees only require one measurement each. For some test system can not measure 180º EIRP, then the extrapolation approach can be adopted when generating the 3D antenna pattern. All of the measured power values will be integrated to TRP, as defined in Clause 5.1.

When measuring TRP in an isotropic Rayleigh fading environment, 100 independent samples can be considered sufficient for ensuring an expanded accuracy better than 0.5dB. Observe that it might be necessary to collect more samples than this during a measurement sequence in order to ensure a sufficient number of independent samples, since all collected samples might not be independent. The number of samples to be collected during a measurement sequence for each chamber should be chosen so that the overall measurement uncertainty is within the specified limits.

8.2.2	UE configurations

For devices containing multiple Tx antennas, the Tx Antenna Switching (TAS) function should be OFF, and the TRP should be measured for each Tx antenna individually. The antenna with better TRP is identified as the primary antenna, and the corresponding TRP result will be used to determine the pass/fail compliance. Otherwise the primary antenna should be selected based on manufacturer declaration. To ensure the TAS OFF testing, the manufacture should provide either software/guidance to lab to control which Tx antenna is used, or the pre-configured DUT locked at primary antenna.

The NR System Simulator (SS) and DUT shall be configured per TS 38.521-1 [5], section 6.2 (Transmitter Power) using the default settings specified in TS 38.521-1 [5] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Table 4.3-1.

The NR SS should send continuous uplink power control “up” commands to the DUT to ensure the DUT's transmitter is at maximum output power during the TRP test.

8.2.3	Test procedure

For TRP measurement, the evaluations shall be performed at maximum transmit power.

For the anechoic chamber method the measurement procedure includes the following steps:

1) Place the DUT inside the QZ following the positioning guideline defined in Clause 6.

2) Connect the SS with the DUT through the link antenna.

3) Measure the power at each measurement point, and calculate EIRPθ,ϕ by adding the composite loss of the entire transmission path.

The TRP value is calculated using the TRP integration approaches outlined in Clause 5.1.

For the reverberation chamber method, the measurement procedure includes the following steps:

1) Place the DUT inside the RC.

2) Connect the SS with the DUT through the fixed measurement antenna.

3) Measure a sufficient number of independent samples.

The TRP value is calculated using the equations outlined in Clause 5.1.

--------------End of text proposal 4-------------

--------------Start of text proposal 5------------

8.3	Total Radiated Sensitivity (TRS)

8.3.1	Test Conditions

The TRS of the DUT is measured by sampling effective isotropic sensitivity (EIS) of the DUT with three-dimensional scan at various locations surrounding the device. The measurement is performed with a constant sampling step of 30 degrees in both theta () and phi () axes for TRS measurement.

EIS, or receiver sensitivity measurements, is defined as the minimum downlink signal power received at the UE antenna input required to provide a data throughput rate greater than or equal to 95% of the maximum throughput of the reference measurement channel (RMC) (the maximum throughput is per Appendix A of TS 38.521-1 [5]).

When measuring TRS in an isotropic Rayleigh fading environment, 100 independent samples can be considered sufficient for ensuring an expanded accuracy better than 0.5dB. Observe that it might be necessary to collect more samples than this during a measurement sequence in order to ensure a sufficient number of independent samples, since all collected samples might not be independent. The number of samples to be collected during a measurement sequence for each chamber should be chosen so that the overall measurement uncertainty is within the specified limits.

8.3.2	UE configurations

For TRS measurement, no specific setting is needed for Rx antennas. By default, the maximum number of Rx antennas supported at each band should be enabled during the TRS test.

The NR System Simulator (SS) and DUT shall be configured per section 7.3 (Reference Sensitivity) of TS 38.521-1 [5] using the defaults specified in TS 38.521-1 [5] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Table 4.3-2.

The NR SS should send continuous uplink power control “up” commands to the DUT to ensure the DUT's transmitter is at maximum output power during the TRS test.

8.3.3	Test procedure

For the anechoic chamber method the measurement procedure includes the following steps:

1) Place the DUT inside the QZ following the positioning guideline defined in Clause 6.

2) Connect the SS with the DUT through the measurement antenna.

3) Determine each EIS, i.e., the minimum power level at which the throughput exceeds or equal to 95% of the maximum throughput of the specified RMC, at each sampling point. The downlink power step size shall be no more than 0.5 dB when the RF power level is near the NR sensitivity level.

The TRS value is calculated using the equation outlined in Clause 5.2.

For the reverberation chamber method, the measurement procedure includes the following steps:

1) Place the DUT inside the RC.

2) Connect the SS with the DUT through the fixed measurement antenna.

3) Determine each received power, i.e., the minimum power level at which the throughput exceeds or equal to 95% of the maximum throughput of the specified RMC, at each independent sample. The downlink power step size shall be no more than 0.5dB when the RF power level is near the NR sensitivity level.

The TRS value is calculated using the equations outlined in Clause 5.2.

--------------End of text proposal 5-------------

--------------Start of text proposal 6------------

9	EN-DC test methodology

9.1	General

For UE radiated conformance testing P-MPRc shall be 0 dB.

9.2	Total Radiated Power (TRP)

9.2.1	Test Conditions

For UE TRP measurement in EN-DC mode, the TRP measurement should be only performed at NR carrier. An LTE link antenna is used to provide a stable LTE link to the DUT without precise path loss or polarization control.

For anechoic chamber and reverberation chamber test method, the sampling step for EN-DC TRP measurement is the same as SA mode, defined in Clause 8.2.1.

9.2.2	UE configurations

For UE TRP measurement in EN-DC mode, the UL output power of LTE carrier should be set as a constant power of 10dBm, while measuring NR at maximum output power, i.e., with fixed p-MaxEUTRA-r15=10 dBm, and p-NR-FR1 not configured.

For devices supporting multiple Tx antennas, the guidelines specified in Clause 8.2.2 shall be used.

For EN-DC, the SS and DUT shall be configured per TS 38.521-3 [6], Section 6.2B.1 (UE Maximum Output Power for EN-DC) using the default settings specified in TS 38.521-3 [6] and TS 38.508 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Table 4.3-3 for EN-DC TRP.

The NR SS should send continuous uplink power control “up” commands to the DUT to ensure the DUT’s transmitter is at maximum output power during the TRP test.

9.2.3	Test procedure

For TRP measurement, the evaluations shall be performed at 10dBm LTE output power and NR maximum transmit power, i.e., with fixed p-MaxEUTRA-r15=10 dBm, and p-NR-FR1 not configured.

The measurement procedure for EN-DC TRP is identical to that for SA mode specified in Clause 8.2.3, as long as the specific SS and UE configuration in Clause 9.2.2 is adopted.

9.3	Total Radiated Sensitivity (TRS)

9.3.1	Test Conditions

For UE TRS measurement in EN-DC mode, the TRS measurement should be only performed at NR carrier. When NR is under test, the LTE carrier is used only to provide a stable connection, and therefore a calibrated downlink power level to the DUT is not required. It is recommended that the DL power on the untested LTE carrier be limited in order to minimize the potential interference for NR.

The definition of EN-DC TRS measurement and the sampling step is identical to that for SA mode, defined in Clause 8.3.1.

9.3.2	UE configurations

For EN-DC TRS measurement, the UL power configuration for LTE and NR is 50%-50% power splitting, i.e.,

- For PC3, p-MaxEUTRA-r15=20 dBm, and p-NR-FR1= 20dBm;

- For PC2, p-MaxEUTRA-r15=23 dBm, and p-NR-FR1= 23dBm.

For devices supporting multiple Rx antennas, the guidelines specified in Clause 8.3.2 shall be used. It should be noted that all receivers on both LTE and NR shall be enabled during EN-DC test.

The EN-DC SS and DUT shall be configured per section 7.3B.2 (Reference Sensitivity for EN-DC) of TS 38.521-3 [6], using the defaults specified in TS 38.521-3 [6] and TS 38.508 [7], as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Table 4.3-3 for EN-DC TRS.

The NR SS should send continuous uplink power control “up” commands to the DUT to ensure the DUT’s transmitter is at maximum output power during the TRP test.

9.3.3	Test procedure

For EN-DC TRS test, only NR carrier measurement is required. The test procedure for EN-DC TRS is identical to that defined in Clause 8.3.3, as long as the specific SS and UE configuration in Clause 9.3.2 is adopted.

--------------End of text proposal 6-------------

--------------Start of text proposal 7------------

Annex B (informative):
Estimation of measurement uncertainty

B.1	General

- Individual uncertainty contributions in the TRP and TRS measurements are discussed and evaluated in this Annex. A technique for calculating the total measurement uncertainty is also presented.

- An important part of a standard measurement procedure is the identification of uncertainty sources and the evaluation of the overall measurement uncertainty. There are various individual uncertainty sources in the measurement procedure that introduce a certain uncertainty contribution to the final measurement result. The approach in this standard test procedure is that the test laboratories are not limited to using some specific instruments and antenna positioners, for example.

- The TRP/TRS measurement procedure can be considered to include two stages. In Stage 1 the calibration of the absolute level of the DUT measurement results is performed by means of using a calibration antenna whose absolute gain/radiation efficiency is known at the frequencies of interest. In Stage 2 the actual measurement of the 3-D pattern of the Device Under Test (DUT) is performed. The uncertainty contributions are analysed in clause B.2 while the uncertainty budget and example tables related to TRP and TRS are listed in clauses B.3 and B.4 respectively.

- The calculation of the uncertainty contribution is based on the Guide to the expression of uncertainty in measurement [10]. Each individual uncertainty is expressed by its Standard Deviation (termed here as ‘standard uncertainty’) and represented by symbol U. The uncertainty contributions can be classified to two categories: Type-A uncertainties, which are statistically determined e.g. by repeated measurements, and Type-B uncertainties, which are derived from existing data e.g. data sheets. Several individual uncertainties are common in Stage 1 and Stage 2 and therefore cancel.

- The procedure of forming the uncertainty budget is:

1)	Compile lists of individual uncertainty contributions for TRP or TRS measurement in both Stage 1 and Stage 2.

2)	Determine the standard uncertainty of each contribution by

a)	Determining the distribution of the uncertainty (Actual, U-shaped, rectangular, etc.)

b)	Determining the maximum value of each uncertainty (unless the distribution is Actual)

c)	Calculating the standard uncertainty by dividing the uncertainty by 2 if the distribution is U-shaped, and by 3 if the distribution is rectangular.

3)	Convert the units (if necessary) of each uncertainty element into the chose unit, i.e. dB.

4)	Combine all the standard uncertainties by the root-sum-squares (RSS) method to derive the ‘combined standard uncertainty’.

5)	Multiply the result by an expansion factor of 1.96 to derive expanded uncertainty at 95% confidence level: 1.96·Uc.

The combination of uncertainties is performed using dB values for simplicity. It has been shown that using dB uncertainty values gives a slightly worse combined uncertainty result than using linear values for the uncertainties. The analysis method therefore errs on the safe side.

B.2	Measurement uncertainty contribution descriptions

B.2.1	Mismatch uncertainty

If the same chain configuration (e.g. including the measurement receiver; the measurement antenna and other elements) is used in both stages, the uncertainty is considered systematic and constant  0.00dB value.

If it is not the case, this uncertainty contribution has to be considered and determined by the following methods.

B.2.1.1	Mismatch uncertainty between measurement receiver / communication tester and the measurement antenna

In a measurement configuration, when two elements (devices, networks…) are connected, if the matching is not ideal, there is an uncertainty in the RF level signal passing through the connection. The magnitude of the uncertainty depends on the VSWR at the junction of the two connectors. In practical measurement system there are probably several connections in a test set-up, they will all interact and contribute to the combined mismatch uncertainty.

The total combined mismatch uncertainty is composed of 2 parts:

1)	The mismatch through the connector between two elements.

2)	The mismatch due to the interaction between two elements.

B.2.1.1.1	Mismatch uncertainty through the connector between two elements

Hereunder, a measurement configuration:

Figure B.2.1.1.1-1: Mismatch uncertainty through the connector

ΓMR is the complex reflection coefficient of the Measurement Receiver. This term is also applicable to the communication tester.

Γcable4 is the complex reflection coefficient of the cable4.

S21 is the forward gain in the network between the two reflection coefficients of interest.

S12 is the backward gain in the network between the two reflection coefficients of interest.

Note that S21 and S12 are set to 1 if the two parts are directly connected.

The uncertainty limits of the mismatch are calculated by means of the following formula (equation 6.1 of [11]):

Mismatch limits% voltage=ΓMR·Γcable4·S21·S12·100

These mismatch limits are divided by 2 (equation 6.2 of [11]) because of the U-shaped distribution of the mismatch uncertainty and give the following standard uncertainty:

Umismatch% voltage=ΓMR·Γcable4·S21·S12·1002

To convert this standard uncertainty in dB, we divide it by the standard uncertainty conversion factor (table 1 of [11]):

UmismatchdB=ΓMR·Γcable4·S21·S12·1002·11.5

B.2.1.1.2	Mismatch uncertainty due to the interaction of several elements

Previously, we presented how to determine the mismatch uncertainty between two elements through the junction (connector). Now, we introduce the other type of mismatch uncertainty, which is a result of the interaction between several elements.

Hereunder, a measurement configuration:

Figure B.2.1.1.1.2-1: Mismatch uncertainty due to the interaction of several elements

Firstly, we determine the mismatch uncertainty between junctions of the elements:

Between the MR and the cable3:

Umismatch1dB=ΓMR·Γcable3·S21·S12·1002·11.5

Between the cable3 and the cable4:

Umismatch2dB=Γcable3·Γcable4·S21·S12·1002·11.5

S21 and S12 are set to 1 because there is no element between cable3 and cable 4.

Umismatch1dB=ΓMR·Γcable3·1002·11.5

Umismatch2dB=Γcable3·Γcable4·1002·11.5

Each mismatch uncertainty due to the interaction between the measurement receiver and the cable4 is determined by means of the following formula:

Umismatch_interaction1dB=ΓMR·Γcable4·S21Cable3·S12Cable3·1002·11.5

S21 and S12 are equal and correspond to the cable3 attenuation.

Umismatch_interaction1dB=ΓMR·Γcable4·S21Cable32·1002·11.5

We consider in the general case, the following measurement configuration:

Figure B.2.1.1.1.2-2: Mismatch uncertainty measurement configuration

In the general case, this uncertainty contribution can be calculated by:

Umismatch_interactionNdB=ΓMR·Γantenna·S21cable1·S12cable1·…·S21cableN·S12cableN·1002·11.5

S21=S12 for passive elements (cables…)

Umismatch_interactionNdB=ΓMR·Γantenna·S21cable12·…·S21cableN2·1002·11.5

B.2.1.1.3	Total combined mismatch uncertainty

The two kinds of mismatch uncertainty contributions are combined by the root-sum-squares (RSS) method to derive the total combined mismatch uncertainty.

The total combined mismatch uncertainty is equal to:

Umismatch12+…+UmismatchN2+Umismatch_interaction12+…+Umismatch_interactionN2

This formula shows that the uncertainty is frequency dependent by the way of the forward and the backward gains in the network between the two components. The uncertainty upon S21 and S12 increases with frequency.

Note that for an anechoic chamber, horn antennas are frequently used as measurement antennas. There are two kinds of horn antennas: single-polarized and dual-polarized. With the second one, it is possible to measure the co-polarized and cross-polarized signals without any movement of the measurement antenna, which reduces the cable antenna uncertainty contribution and improves the measurement stability.

To conduct the signals to the measurement receiver, the measurement system configuration using a dual-polarized horn antenna has to be completed with an RF Relay. This device will include new mismatch uncertainty contributions, which have to be determined with the previously presented calculation methods, completed by the RF relay parameters contributions, and described in the following.

B.2.1.2	Mismatch uncertainty of the RF relay

If the same receiver chain configuration (including the measurement receiver; the measurement antenna and other elements) is used in both stages, the uncertainty is considered systematic and constant  0.00dB value.

If it is not the case, this uncertainty contribution has to be considered and determined by the following method.

The following figure describes the RF Relay with its S-parameters and the complex reflection coefficient of the inputs and output:

Figure B.2.1.2-1: Mismatch uncertainty of the RF relay

The RF relay is used to switchover the cross and direct polarization signals from the measurement antenna. To determine RF Relay mismatch uncertainty contributions, reflection coefficients for each port and the cross-talk attenuation have to be known.

The total combined mismatch uncertainty is composed of two parts:

1)	The mismatch uncertainty contributions when the RF Relay switches on the direct polarization signal

2)	The mismatch uncertainty contributions when the RF Relay switches on the cross-polarization signal

Each part is composed of two types of uncertainties introduced in the previous paragraph: the mismatch through the connector between two elements and the mismatch due to the interaction between several elements.

B.2.1.2.1	First part: RF Relay switched on the co-polarized signal

B.2.1.2.1.1	The mismatch through the connector between two elements

Between the Input1 and the port1:

Umismatch1dB=Γ1·S11·1002·11.5

Between the port3 and the Output:

Umismatch2dB=Γ3·S33·1002·11.5

Between the Input2 and the port2:

The RF Relay switchovers on the direct polarization signal. As a result, there is no mismatch uncertainty contribution.

B.2.1.2.1.2	Mismatch due to the interaction between two elements or more

Between the Input1 and the Output:

Umismatch_interaction1dB=Γ1·Γ3·S31·S13·1002·11.5

Between the Input1 and the Input2:

Umismatch_interaction2dB=Γ1·Γ2·S21·S12·1002·11.5

The RF Relay switchovers on the cross-polarization signal. As a result; this uncertainty contribution is usually disregarded because of the high crosstalk attenuation which is characterized by S21 and S12 S-parameters. If the crosstalk attenuation is low, this uncertainty contribution has to be considered.

Between the Input2 and the Output:

Umismatch_interaction3dB=Γ2·Γ3·S23·S32·1002·11.5

The RF Relay switchovers on the cross polarization signal. As a result; this uncertainty contribution is usually disregarded because of the high cross-talk attenuation, which is characterized by S23 and S32 S-parameters. If the crosstalk attenuation is low, this uncertainty contribution has to be considered.

B.2.1.2.2	Second part: RF relay switched on the cross-polarized signal

B.2.1.2.2.1	The mismatch through the connector between two elements

Between the Input1 and the port1: the RF Relay switchovers on the direct polarization signal. As a result, there is no mismatch uncertainty contribution.

Between the port3 and the Output:

Umismatch3dB=Γ3·S33·1002·11.5

Between the Input2 and the port2:

Umismatch4dB=Γ2·S22·1002·11.5

B.2.1.2.2.2	Mismatch due to the interaction between two elements or more

Between the Input1 and the Output:

Umismatch_interaction4dB=Γ1·Γ3·S31·S13·1002·11.5

The RF Relay switchovers on the cross-polarization signal. As a result; this uncertainty contribution is usually disregarded because of the high crosstalk attenuation which is characterized by S31 and S13 S-parameters. If the crosstalk attenuation is low, this uncertainty contribution has to be considered.

Between the Input1 and the Input2:

Umismatch_interaction5dB=Γ1·Γ2·S21·S12·1002·11.5

The RF Relay switchovers on the cross-polarization signal. As a result; this uncertainty contribution is usually disregarded because of the high crosstalk attenuation which is characterized by S21 and S12 S-parameters. If the crosstalk attenuation is low, this uncertainty contribution has to be considered.

Between the Input2 and the Output:

Umismatch_interaction6dB=Γ2·Γ3·S23·S32·1002·11.5

B.2.1.2.3	Total combined mismatch uncertainty

Each non-zero mismatch uncertainty contribution from both parts (RF Relay switched on the cross and direct polarization signal) are combined by the root-sum-squares (RSS) method to derive the total combined mismatch uncertainty.

The total combined mismatch uncertainty is equal to:

Umismatch12+…+UmismatchN2+Umismatch_interaction12+…+Umismatch_interactionN2

If a RF Relay is used to drive the cross and direct polarization signals from the dual-polarized antenna, this total combined mismatch uncertainty has to be added with all the uncertainty measurement contributions for the total combined measurement uncertainty.

B.2.2	Insertion loss

B.2.2.1 Insertion loss of the measurement antenna cable

If the measurement antenna cable does not move between the calibration and the DUT measurement stage, the uncertainty due to the insertion loss of the cable is assumed to be systematic. Moreover, this uncertainty is common and constant in both stages and that is why this leads to 0.00dB value.

If a different cable is used in the calibration measurement and in the DUT measurement, and the difference of the insertion loss is used in the calculations, then the overall combined standard uncertainty of the insertion loss measurement should be used in the uncertainty budget. The distribution of this uncertainty is assumed to be rectangular, in which case the standard uncertainty can be calculated as the maximum value divided by 3.

B.2.2.2	Insertion loss of the measurement antenna attenuator (if used)

See Insertion loss of the measurement antenna cable

If the measurement antenna attenuator is used in both stages, the uncertainty is considered systematic and constant  0.00dB value.

B.2.2.3	Insertion loss of the RF relays (if used)

See Insertion loss of the measurement antenna cable.

If the RF relay is used in both stages, the uncertainty is considered systematic and constant  0.00dB value.

B.2.2.4	Insertion loss: calibration antenna feed cable

The feed cable of the calibration antenna only appears in Stage 1. As a result, this uncertainty has to be considered.

This uncertainty will be measured or calculated from the manufacturer’s data in logs with a rectangular distribution (see clause 5.1.2 in [11]).

B.2.2.5	Insertion loss: calibration antenna attenuator (if used)

If a calibration antenna attenuator is used, it only appears in Stage 1. As a result, this uncertainty has to be considered.

This uncertainty will be calculated from the manufacturer’s data in logs with a rectangular distribution (see clause 5.1.2 in [11]).

B.2.3	Influence of the antenna cable

B.2.3.1	Measurement antenna cable

If the measurement antenna is directional (i.e. peak gain >+5dBi e.g. horn, LPDA, etc.) and the same measurement antenna cable configuration is used for both stages, the uncertainty is considered systematic and constant  0.00dB value.

B.2.3.2	Calibration antenna cable

If an efficiency calibration is performed, influence of the calibration antenna feed cable can be assumed to be negligible, due to data averaging.

In the case of gain calibration, the influence of the calibration antenna feed cable must be assessed by measurements. A gain calibration measurement is repeated with a reasonably differing routing of the feed cable. Largest difference between the results is entered to the uncertainty budget with a rectangular distribution.

B.2.4	Measurement receiver: uncertainty of the absolute level

The receiving device is used to measure the received signal level in TRP tests either as an absolute level or as a relative level. Receiving device used is typically a communication tester, spectrum nalyser (SA), or power meter (PM). Generally, there occurs an uncertainty contribution from limited absolute level accuracy and non-linearity.

This uncertainty will be determined from the manufacturer’s datasheet and the distribution used (see clause 5.1.2 in [11]) shall match that provided in the datasheet. In the absence of a declared distribution in the datasheet, the rectangular distribution should be used.

B.2.5	Communication tester: uncertainty of the absolute level

The transmitter device (typically a communication tester or BS simulator) is used to drive a signal to the measurement antenna in sensitivity tests either as an absolute level or as a relative level. Receiving device used is the UE. Generally, there occurs uncertainty contribution from limited absolute level accuracy and non-linearity of the communication tester.

For practical reasons, the calibration measurement (Stage 1) should be only performed with the measurement antenna as a receiver. Hence, the uncertainty on the absolute level of the transmitter device cannot be assumed as systematic. This uncertainty will be determined from the manufacturer’s datasheet and the distribution used (see clause 5.1.2 in [11]) shall match that provided in the datasheet. In the absence of a declared distribution in the datasheet, the rectangular distribution should be used. Furthermore, the uncertainty of the non-linearity of the device is included in the absolute level uncertainty.

B.2.6	Sensitivity measurement: output level step resolution

When output power of the communication tester is swept to reach the throughput target that defines the sensitivity threshold, used power step resolution creates this uncertainty.  Output power step used in the sensitivity measurement is divided by factor 2 and then a rectangular distribution applied to obtain the uncertainty.

B.2.7	Measurement distance

The uncertainty contribution from a finite measurement distance is estimated in three parts. The two three elements of mismatch uncertainty contributions are combined by the root-sum-squares (RSS) method to derive the total combined mismatch uncertainty.

B.2.7.1	Offset of phase centre from axis(es) of rotation

B.2.7.1.1	Offset of DUT phase centre from axis(es) of rotation

In all the measurements defined in this test procedure the DUT and phantom combination is rotated about the ear reference point of SAM phantom, which is also assumed to be the location of the phase centre in both angular directions of the measurements.

For some positioning systems this may be practically impossible in which case a measurement uncertainty contribution can arise because the phase centre will rotate on a non-zero radius about the centre of rotation, thereby giving a variable measurement distance. Data averaging process may lead to a partial self-cancel of this uncertainty.

The uncertainty limits of this effect are calculated by means of the following formula (uj22 of [12]):

Phase centre limits%=± estimated offset from the axis of rotationrange length·100

Because of the phase center can be anywhere between the offset limits, the distribution is assumed to have a rectangular and the phase center limit is divided by 3giving the following standard uncertainty:

Uphase centre offset%=± estimated offset from the axis of rotationrange length·3·100

To convert this standard uncertainty in dB, we divide it by the standard uncertainty conversion factor (table 1 of [11]):

Uphase centre offsetdB=± estimated offset from the axis of rotationrange length·3·11.5·100

B.2.7.1.2	Offset of calibration antenna phase centre from axis(es) of rotation

If a gain calibration is performed in Stage 1 with a directive antenna (e.g. horn antenna), the uncertainty contribution of calibration antenna’s phase centre displacement is estimated by means of the following formula (uj21 of [12]):

Phase centre limits%=± maximum dimension of the antenna2·range length·100

Because of the phase center can be anywhere between the offset limits, the distribution is assumed to have a rectangular distribution and the phase center limit is divided by 3giving the following standard uncertainty:

Uphase centre offset%=± maximum dimension of the antenna2·range length·3·100

To convert this standard uncertainty in dB, we divide it by the standard uncertainty conversion factor (table 1 of [11]):

Uphase centre offsetdB=± maximum dimension of the antenna2·range length·3·11.5·100

If a gain calibration is performed in Stage 1 with omnidirectional calibration antenna (e.g. sleeve dipoles), uncertainty should be 0.00 dB provided that care is taken in their positioning since the phase centre are easily identifiable.

For an efficiency calibration with an omnidirectional calibration antenna, the Uphase centre offset is calculated similarly as for gain calibration but the uncertainty may be divided by factor 2. This is due to correcting impact of data averaging in this type of calibration.

B.2.7.2	Mutual coupling

In measurement of radio performances of UMTS mobile phones in speech mode, the mutual coupling uncertainty for this frequency band is a 0.00dB value (see annex A.2 in [13]).

The 0.00dB value can be extended for NR FR1 band frequencies.

B.2.7.3	Phase curvature

This uncertainty originates from the finite far-field measurement distance, which causes phase curvature across the DUT. If the minimum measurement distance is respected, this error is assumed to be negligible..

B.2.8	Quality of quiet zone

The uncertainty contribution due to unwanted reflections and obstructions within the anechoic chamber, including imperfect absorber treatments and the impact of positioning equipment support structure, is determined from data acquired using the ripple test methodology in clause 7.4. This data consists of single-axis pattern cuts that represent the sum of direct and reflected rays from a highly symmetrical omnidirectional radiation pattern measured at various points throughout the test volume. The data must be measured in sufficient spatial or angular resolution to accurately capture the peaks and nulls of the pattern to within a small fraction of the overall ripple contribution. In general, the worst-case peak-to-null ripple will reflect the potential error in a peak EIRP or EIS measurement for an omnidirectional DUT pattern located anywhere within the test volume. Note however that nulls in the pattern can exhibit considerably larger errors due to reflected signals being stronger than the line-of-sight signal from the null.

When measuring the range path loss in Stage 1 using a dipole pattern, the associated measurement uncertainty may be determined from the peak-to-null ripple, after relative path loss compensation, of a single radial offset ripple test, where the range is configured as for range calibration (e.g. with any extraneous support structure removed) using a rectangular distribution.

For spherically integrated quantities such as TRP and TRS, the peak-to-null ripple would overestimate the measurement uncertainty due to the inherent averaging of the various peaks and nulls as the spherical pattern is integrated. In this case, the surface standard deviation (SSD) [15] is used to obtain a statistical representation of the expected impact of ripple on the integrated power from an isotropic radiator placed anywhere within the test volume. Due to the impracticality of maintaining a constant path loss reference between individual ripple test cuts, each resultant pattern is treated individually and then the worst case SSD result is chosen as the standard uncertainty of the quiet zone.

For the phi-axis ripple tests, the pattern can be considered an equatorial (theta = 90°) cut of the isotropic pattern where every point has equal weighting on an evenly spaced spherical surface.  Thus, the standard deviation of the single cut should be equivalent to the standard deviation of the entire spherical surface.  Defining pk as the kth ripple measurement point in linear power units pk = 10PkdB10, and p as the average of all  pk values in the associated ripple test, then the standard deviation of the corresponding cut is given by:

spk=1N-1k-0N-1pkp-12

On the theta-axis ripple test each ripple test cut can be considered as a great circle cut through an isotropic pattern with the symmetrical distortions that would be produced by revolving the pattern about the phi (0-180°) axis. Thus, it becomes apparent that ripple near the poles impacts a smaller total surface area on the sphere than that near theta = 90 and 270°. In this case, sin(theta) weighting is used to generate the spherical surface weighted standard deviation as:

spk = 1N-1k=0N-1pkp-1sinθk2

Note that this equation simplifies to the previous equation when theta = 90°, so the two formulations are in fact the same, regardless of which orientation of the ripple test is used.

The standard uncertainty for the quiet zone ripple contribution to the TRP/TRS measurement is then given by the maximum SSD from all of the ripple test measurements.

B.2.9	DUT Tx-power drift

A single point power reference measurement in the beginning and at the end of the measurement procedure is recommended to monitor the power drift of the DUT. Based on TX-power drift measurements for typical 3G UE, an uncertainty of 0.2 dB shall be entered to uncertainty budget with a rectangular distribution. If the drift measurement indicates larger drift, the actual drift shall be included to uncertainty.

In order to minimize Tx-power drift error it is recommended to interleave sensitivity and power measurement of multiple channels. This spreads the measurements over a longer period, which helps to average the drift of the TX-power.

Typical TX-power drifts of 3G UE, measured in a single angular point, DUT placed against phantom head are shown in Figure B.2.9-1.

Figure B.2.9-1: Output power variation of typical 3G UE during battery life

B.2.10	DUT sensitivity drift

Due to statistical uncertainty of sensitivity measurement, drift in the TRS cannot be monitored similarly to TRP. An uncertainty value of 0.2dB can be used with a rectangular distribution, or the TRS drift should be measured, with a setup corresponding to the actual TRS measurement.

B.2.11	Uncertainty related to the use of phantoms

B.2.11.1	Uncertainty from using different types of SAM phantom

This uncertainty contribution originates from the fact that different laboratories may use the two different versions of SAM head: the SAM head phantom or the SAM phantom including the head and the shoulders. The standard SAM head is the specified phantom. However, the use of the other type of SAM is also allowed with the requirement that the resulting uncertainty contribution is considered in the uncertainty budget.

B.2.11.2	Simulated tissue liquid uncertainty

This uncertainty will occur, if the laboratory uses a liquid which has dielectric parameters deviating more than ±15% of the target parameters.

B.2.11.3	Uncertainty of dielectric properties and shape of the hand phantom

The hand phantom contributes to OTA measurement uncertainty due to the manufacturing tolerances of its dielectric properties and shape. The dielectric properties on the surface of the hand may differ from those of its interior, so both are included in the evaluation. The moulded exterior surface of the hand shall be measured directly with an open-ended coaxial probe. The interior hand material is evaluated indirectly, by substituting a cube-shaped sample moulded from the same material and having some exterior surfaces removed. Following procedure will be used to evaluate the dielectric properties of the hand phantom;

1.	Each hand shall be manufactured together with a reference cube of the same material. The sides of the reference cube shall be not less than 40 mm in length.

2.	The moulded surface on three orthogonal sides of the cube shall be sliced away to a depth of at least 3 mm, in order to expose interior material for evaluation. The remaining three sides of the cube shall be left untreated.

3.	Relative permittivity and conductivity shall be measured at ten different points on each of the three cut, exposed surfaces of the reference cube, and the combined interior averages (εintavg, σintavg, 30 points) and standard deviations (εintstd,σintstd,30 points) shall be calculated. Individual interior averages for each of these three sides (εinti,σinti,10 points) shall also be calculated.

4.	Relative permittivity and conductivity shall be measured at ten points on the hand phantom exterior. A measurement point shall be located to each fingertip or as close to the tip as applicable. One measurement point shall be located to the back of the hand and one to the inner surface of wrist area. The exterior averages (εext_avg, σext_avg, 10 points) and standard deviations (εext_std, σext_std, 10 points) calculated accordingly.

5.	The total averages (εavg , σavg) shall be calculated as the average of exterior and interior values by either evaluating all data points or using equations :

εavg=εext_avg+3·εint_avg4

σavg=σext_avg+3·σint_avg4

6.	The total standard deviations (εstd , σstd) shall be calculated as the statistical combination of exterior and interior values by either evaluating all data points or using equations:

εstd=14εext_std2+εext_avg2+3*εint_std2+εint_avg2-εavg2

σstd=14σext_std2+σext_avg2+3*σint_std2+σint_avg2-σavg2

7.	The hands are acceptable for radiated performance testing, i.e., meet the minimal requirements, if

a.	εavg deviates by less than 15% from the target values

b.	σavg deviates by less than 25% from the target values

c.	the difference between the averaged permittivity of each 10-point interior surface εint_avg deviates by less than 10% and εext_avg by less than 20% from the total average εavg

d.	the difference between the averaged conductivity of each 10-point interior surfaceσint_avg deviates by less than 20% and σext_avg by less than 30% from the total average σavg

e.	the standard deviation of the combined measurements (30 interior points and 10 exterior points) is less than 20% for permittivity εstd and less than 40% for conductivity σstd

8.	For the hands meeting the minimal requirements of step 7, the following approximations shall be used to determine the hand uncertainty due to dielectric properties.

UεdB=c1·10log101+∆εavg2+εunc2+a1εstd2ε

UσdB=c2·10log101+∆σavg2+σunc2+a1σstd2ε

∆εavg, ∆σavg, εstd, σstd are the values determined as defined above and εunc and σunc are expanded measurement uncertainties (k = 2) of the dielectric parameter measurement method. The cube will be provided together with the hand such that the user can evaluate if the interior (cube) properties of the hand has degenerated over time by performing the test above. Coefficient c1=0.78, c2=0.39 and a1=0.50 were determined by numeric simulations.

In case the hand phantoms are manufactured within CAD models, the tolerance is 2% and therefore the effects shape errors are negligible. If the tolerance is larger, a numerical study must be conducted.

B.2.11.4	Uncertainty from using different types of Laptop Ground Plane phantom

This uncertainty contribution originates from the fact that different laboratories may use different variations of Laptop Ground Plane phantom. The standard Laptop Ground Plane is the specified phantom.

B.2.12	Coarse sampling grid

Decreasing of sampling density to finite number of samples affects the measurement uncertainty by two different errors. First is due to inadequate number of samples and second is a systematic discrimination approximation error in TRP and TRS equations.

Figure B.2.12-1 shows simulated sampling grid errors for typical 3G UE. Approximation error is not included. Simulations are based on thin plate surface interpolation of real radiation patterns, measured beside a phantom head.

Figure B.2.12-1: Simulated TPR/TRS error as a function of sampling grid

The offset of systematic approximation error can be expressed by using formula

Offset=10log10π2Nn-1Nsinθn

where

N is number of angular intervals in elevation,

θn is elevation.

Figure B.2.12-2: Approximation error of TRP/TRS

The 15 sampling grid used in TRP measurements has been shown to introduce only very small differences as compared to the results obtained with denser grids, so with that sampling grid the uncertainty contribution can assumed negligible.

When using sample step size of 15 - 30, standard uncertainty of 0.15dB can be assumed to cover errors. If step size >30 is used, larger uncertainty should be considered.

B.2.13	Random uncertainty

The random uncertainty characterizes the undefined and miscellaneous effects which cannot be forecasted. One can estimate this type of uncertainty with a repeatability test by making a series of repeated measurement with a reference DUT without changing anything in the measurement setup.

The random uncertainty differs from one laboratory to another. Moreover, each DUT has its own electromagnetic behaviour and random uncertainty. Some uncertainty also occurs from the positioning of the DUT against the SAM phantom, as the DUT cannot be attached exactly in the same way every time. This uncertainty depends on how much the DUT’s positioning against the SAM phantom and hand phantoms varies from the specified testing positions. It is noted that the uncertainty of the phone positioning depends on the phone holder and the measurement operator and is in fact difficult to distinguish from random uncertainty. Some uncertainty also occurs from the positioning of the DUT plugged into the Laptop Ground Plane phantom, as the DUT may not be plugged into the USB connector and positioned exactly in the same way every time. This uncertainty depends on how much the DUT’s position plugged into the Laptop Ground Plane phantom varies from the specified plug-in position. Therefore, the positioning uncertainty is included in random uncertainty.

To estimate this uncertainty for the SAM phantom, it is suggested to perform at least five evaluations of TRP/TRS whereby the device shall be dismounted and newly positioned with a fully charged battery before each test. This measurement set has to be carried out in mid channel of lowest and highest frequency bands utilized by the testing lab, for at least three phones with different type of mechanical design. The values have to be normalized by the mean for each measurement set. As a result, the uncertainty contribution entered to uncertainty budget is the difference between the maximum and minimum normalized values.

With head and hand phantoms, random uncertainty evaluation may be done separately for each measurement configuration i.e. head only, browsing mode or speech mode. A speech mode random uncertainty evaluation, were both head and hand phantoms are used, can reasonably be considered to be the worst-case scenario and thus random uncertainties in other configurations to be less.

To estimate this uncertainty for the Laptop Ground Plane phantom, it is suggested to perform at least five evaluations of TRP/TRS for the plug-in position whereby the device shall be dismounted and newly positioned before each test. This measurement set has to be carried out in mid channel of lowest and highest frequency bands utilized by the testing lab, for at least three USBs with different type of mechanical design. The values have to be normalized by the mean for each measurement set. As a result, the uncertainty contribution entered to uncertainty budget is the difference between the maximum and minimum normalized value.

B.2.14	Frequency response

Test systems might not be able to ensure flat frequency response across the entire channel bandwith required for testing (e.g. up to 100MHz for NR FR1). When a frequency response correction based on the results from the system calibration measurements in Stage 1 is not possible or practical, this uncertainty has to be considered.

This uncertainty term can be estimated as described in [14] using the following formula:

εj=10logk=j-N2j+N2PLkN+1PLj

where εj is the expected relative error in the average power result for a given channel in dB, PLj, is the linear path loss at the center frequency of the given channel, PLk, is the linear path loss at each frequency point across the corresponding channel, and N is the number of frequency steps across a given channel bandwidth. The maximum deviation εj across all of the possible channels in a band shall be used to estimate the required frequency response uncertainty contribution with a rectangular distribution.

This error may be removed directly at each frequency, fj, by using the average path loss across the channel as the range loss correction rather than the path loss at the center frequency as described in [14].

For sensitivity measurements, this effect is included in the output level step resolution.

B.2.15	Uncertainty of network nalyser

This uncertainty includes the all uncertainties involved in the S21 measurement with a network nalyser, and will be determined from the manufacturer’s datasheet and the distribution used (see clause 5.1.2 in [11]) shall match that provided in the datasheet. In the absence of a declared distribution in the datasheet, the rectangular distribution should be used.

B.2.16	Uncertainty of the gain/efficiency of the calibration antenna

The calibration antenna only appears in Stage 1. Therefore, the gain/efficiency uncertainty has to be considered.

This uncertainty will be determined from the manufacturer’s datasheet and the distribution used (see clause 5.1.2 in [11]) shall match that provided in the datasheet. In the absence of a declared distribution in the datasheet, the rectangular distribution should be used. Alternatively, the uncertainty could come from a calibration report with traceability to a National Metrology Institute with measurement uncertainty budgets generated following the guidelines outlined in internationally accepted standards

If the manufacturer’s data do not give the information, the value has to be checked, see annex A.12 in [13].

B.3	Total Radiated Power (TRP)

The uncertainty contributions related to TRP are listed in Table B.3-1. A preliminary example uncertainty budget is presented in Table B.3-2.

Table B.3-1 Uncertainty contributions in TRP measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure 7.2-1, Figure 7.2-2) |  |  |
| 1 | Mismatch of receiver chain | B.2.1 |
| 2 | Insertion loss of receiver chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Measurement Receiver: uncertainty of the absolute level | B.2.4 |
| 5 | Measurement distance | B.2.7 |
| 6 | Quality of quiet zone | B.2.8 |
| 7 | DUT Tx-power drift | B.2.9 |
| 8 | Uncertainty related to the use of phantoms | B.2.11 |
| 9 | Coarse sampling grid | B.2.12 |
| 10 | Random uncertainty | B.2.13 |
| 11 | Frequency Response | B.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure 7.3-1) |  |  |
| 12 | Uncertainty of network analyzer | B.2.15 |
| 13 | Mismatch of receiver chain | B.2.1 |
| 14 | Insertion loss of receiver chain | B.2.2 |
| 15 | Mismatch in the connection of calibration antenna | B.2.1 |
| 16 | Influence of the calibration antenna feed cable | B.2.3 |
| 17 | Influence of the measurement antenna cable | B.2.3 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | B.2.16 |
| 19 | Measurement distance | B.2.7 |
| 20 | Quality of the Quiet Zone | B.2.8 |

Table B.3-2 Preliminary example of uncertainty budget for TRP hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гpower meter <0.05 Гmeasurement antenna <0.16 | 0.07 | U-shaped | 1.41 | 1 | 0.05 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Power Meter | 0.06 | Rectangular | 1.73 | 1 | 0.03 |
| 5 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | Rectangular | 1.73 | 1 | 0.16 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | Actual | 1 | 1 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | Rectangular | 1.73 | 1 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | 0.32 | Rectangular | 1.73 | 1 | 0.18 |
| 9 | Coarse sampling grid | Negligible 15° sampling grid | 0 | Actual | 1 | 1 | 0.00 |
| 10 | Random Uncertainty | Monoblock, clamshell and PDA design used for testing | 0.81 | Rectangular | 1.73 | 1 | 0.47 |
| 11 | Frequency Response | Average path loss corrected | 0 | Rectangular | 1.73 | 1 | 0.00 |
|  | Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | Manufacturer’s uncertainty | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | Rectangular | 1.73 | 1 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| 19 | Measurement distance | Dipole: aligned with phase center | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| Combined standard uncertainty |  |  |  |  |  |  | 0.91 |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | 1.78 |

B.4	Total Radiated Sensitivity (TRS)

The uncertainty contributions related to TRS are listed in Table B.4-1. A preliminary example uncertainty budget is presented in Table B.4-2.

| Table B.4-1 Uncertainty contributions in TRS measurement for anechoic chamber methodUID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure 7.2-1, Figure 7.2-2) |  |  |
| 1 | Mismatch of transmitter chain | B.2.1 |
| 2 | Insertion loss of transmitter chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Communication Tester: uncertainty of the absolute output level | B.2.5 |
| 5 | Sensitivity measurement: output level step resolution | B.2.6 |
| 6 | Measurement distance | B.2.7 |
| 7 | Quality of quiet zone | B.2.8 |
| 8 | DUT sensitivity drift | B.2.10 |
| 9 | Uncertainty related to the use of phantoms | B.2.11 |
| 10 | Coarse sampling grid | B.2.12 |
| 11 | Random uncertainty | B.2.13 |
| 12 | Frequency Response | B.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure 7.3-1) |  |  |
| 13 | Uncertainty of network analyzer | B.2.15 |
| 14 | Mismatch of transmitter chain | B.2.1 |
| 15 | Insertion loss of transmitter chain | B.2.2 |
| 16 | Mismatch in the connection of calibration antenna | B.2.1 |
| 17 | Influence of the calibration antenna feed cable | B.2.3 |
| 18 | Influence of the measurement antenna cable | B.2.3 |
| 19 | Uncertainty of the absolute gain/radiation efficiency of the calibration antenna | B.2.16 |
| 20 | Measurement distance | B.2.7 |
| 21 | Quality of quiet zone | B.2.8 |

Table B.4-2: Preliminary example of uncertainty budget for TRS hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | ГCommTester <0.13 Г antenna connection <0.03 | 0.07 | U-shaped | 1.41 | 1 | 0.05 |  |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 4 | Communication Tester: uncertainty of the absolute output level | Manufacturer’s data sheet | 1 | Rectangular | 1.73 | 1 | 0.58 |  |
| 5 | Sensitivity measurement: output level step resolution | Step of 0.5 dB | 0.25 | Rectangular | 1.73 | 1 | 0.14 |  |
| 6 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | Rectangular | 1.73 | 1 | 0.16 |  |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | Actual | 1 | 1 | 0.5 |  |
| 8 | DUT sensitivity drift | Drift measurement | 0.2 | Rectangular | 1.73 | 1 | 0.12 |  |
| 9 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | 0.32 | Rectangular | 1.73 | 1 | 0.18 |  |
| 10 | Coarse sampling grid | 30° sampling grid | 0.15 | Actual | 1 | 1 | 0.15 |  |
| 11 | Random uncertainty | Monoblock, clamshell and PDA used for testing | 0.91 | Rectangular | 1.73 | 1 | 0.53 |  |
| 12 | Frequency Response | Included in the output level step resolution | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers NA setup | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| 14 | Mismatch of transmitter chain | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |  |
| 15 | Insertion loss of transmitter chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |  |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with dipole | 0.3 | Rectangular | 1.73 | 1 | 0.17 |  |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| 20 | Measurement distance | Dipole: aligned with phase center | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 21 | Quality of quiet zone | Peak-to-null ripple | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| Combined standard uncertainty |  |  |  |  |  |  | 1.12 |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | 2.20 |

--------------End of text proposal 7------