3GPP TSG-RAN5 Meeting#82	R5-192659

# Athens, Greece, 25th Feb. – 1st Mar. 2019

Title:							On the QoQZ impact for ACLR

Source:			Anritsu

Agenda Item:			5.3.8.17

Document for:			Endorsement

# 1.	Introduction

# The discussion for FR2 ACLR MU has been provided in [1]-[2]. In this paper, further analysis for FR2 ACLR QoQZ MU is provided.

# 2.	Discussion

2.1.	Experiments

# In [2], the Quality of quiet zone value for the ACLR is provided. The one discussion there was whether the ACLR-QoQZ should be defined with difference of In-band QoQZ of two adjacent frequency or STD(standard deviation) of difference of KPI(TRP for ACLR). The factor which impacts the ACLR measurement is the correlation of MU(error) between two adjacent frequencies, hence STD of difference(ratio) of KPI should be used to evaluate the impact for ACLR.

# Proposal 1 : QoQZ impact for ACLR is evaluated with STD(standard deviation) of difference of TRP for 2 frequencies

# The exact definition of STD of TRP difference can be written as follows.

STD of TRP difference=pos=1Nposdir=1NdirTRPf1,pos,dir-TRPf2,pos,dir-pos=1Nposdir=1NdirTRPf1,pos,dir-TRPf2,pos,dirNpos∙Ndir2Npos∙Ndir-1

Table 1 shows the result of in-band absolute QoQZ and Table 2 shows STD of KPI(EIRP/TRP) difference. Reference antenna position, directions are same as defined in 38.810 for in-band absolute QoQZ.

Table 1: In-band absolute QoQZ Performance

| Positions | Quiet Zone Size | EIRP |  | TRP |  |
| --- | --- | --- | --- | --- | --- |
|  |  | 40.000 GHz | 40.800 GHz | 40.000 GHz | 40.800 GHz |
| P1-P7 | 30cm | +/- 0.50 dB | +/- 0.52 dB | +/- 0.52 dB | +/- 0.53 dB |
| P1 only |  | +/- 0.28 dB | +/- 0.45 dB | +/- 0.35 dB | +/- 0.43 dB |

Table 2: STD of KPI difference

| Positions | Quiet Zone Size | Std. Deviation of Difference in EIRPs between 40.0 and 40.8 GHz | Std. Deviation of Difference in TRPs between 40.0 and 40.8 GHz |
| --- | --- | --- | --- |
| P1-P7 | 30cm | +/- 0.52dB | +/- 0.47dB |
| P1 only |  | +/- 0.32dB | +/- 0.27dB |

2.2.	Consideration

Though the MU for relative measurement tends to be thought to be smaller than the absolute MUs, but it can only be said when there is some level of correlation between absolute MUs. If two absolute MUs are 100% uncorrelated, the MU for relative measurement can be even bigger up to RSS of two MUs i.e. the MU for relative measurement could be in the range 0 to 2 times of absolute MU depending on the correlation of two MUs. The result in Table 1 indicates that the correlation of QoQZ is in the middle of 100% correlated and 100% uncorrelated.

Observation 1 : Table 1 indicates that the correlation of QoQZ of 800MHz separation is in the middle of 100% correlated and 100% uncorrelated.

One main factor is the reflection/scattering within the chamber due to the “non-ideal” anechoicity such as existence of positioner, link antennas, etc…  Assuming 800MHz frequency separation, c/f  = 3.0e+8[m/sec] / 800e+6[Hz] = 37.5 [cm] of path length difference will correspond to the 2π phase difference. Hence, the reflected/scattered signal inside the 30cm QZ can have the 2π phase difference. Therefore it is understandable that with the condition of 800MHz separation and QZ size 30cm, the correlation of QoQZ MU will be reduced at some extent.

Observation 2 : It is understandable that with the condition of 800MHz separation and QZ size 30cm, the correlation of QoQZ will be reduced at some extent

Considering above experiments and analysis,  it is reasonable to adopt the same level of MU as for in-band QoQZ for relative power measurement of ACLR.

# Proposal 2 : For QoQZ impact for ACLR, use 0.5dB as uncertainty value for measurement stage and 0.2dB for calibration stage for Quiet Zone size ≤ 30cm.

# 4.	Conclusion

# Proposal 1 : QoQZ impact for ACLR is evaluated with STD(standard deviation) of difference of TRP for 2 frequencies

Observation 1 : Table 1 indicates that the correlation of QoQZ of 800MHz separation is in the middle of 100% correlated and 100% uncorrelated.

Observation 2 : It is understandable that with the condition of 800MHz separation and QZ size 30cm, the correlation of QoQZ will be reduced at some extent

# Proposal 2 :  For QoQZ impact for ACLR, use 0.5dB as uncertainty value for measurement stage and 0.2dB for calibration stage for Quiet Zone size ≤ 30cm.

# 5.	References

[1] R5-188189, “On the FR2 MU for occupied BW and ACLR”, Anritsu, RAN5#81

[2] R5-190391, “On QoQZ MU Impact on ACLR”, Keysight, RAN5 NR#4 AdHoc