3GPP TSG-RAN5 Meeting #108	R5-254071

# Bengaluru, India, 25th - 29th August, 2025

Agenda item:	5.4.20

Source:	Keysight Technologies UK Ltd

Title:	Discussion on FR1 MUs up to 7.125GHz

Document for:	Discussion and Endorsement

# Introduction

This contribution aims to continue the MU and TT analysis initiated in [1] for FR1 RF test cases operating at frequencies up to 7.125GHz.

# Discussion

## Background

The analysis made in [1] and associated CRs [2] to [4] for MU and TT up to 7.125GHz was limited to the NR-U test cases covered in the NR_unlic-UEConTest Rel-16 work plan because band n104 was not introduced at that moment in RAN5 specs. Those NR-U test cases are only single carrier, single UL (SISO) transmission. Therefore, MU and TT for carrier aggregation or UL MIMO/TxD tests has never been done for frequencies over 6GHz.

The endorsed proposal in [1], replicated here for convenience, did not do a detailed analysis per test case and for that reason it was not precluding revisiting the general proposal in the future if more detailed analysis is provided.

Proposal 1. As a general rule, for the range 6.0GHz < f ≤ 7.125GHz, define the same MTSU already defined for f = 6.0GHz in the given test case type. Exceptions in MTSU definition for specific test cases are not precluded.

In the associated and agreed CRs [2] to [4], MU and TT were defined for all the NR-U test cases except for the EVM test case. In addition, in every test case and in annex F, the following editor’s note was captured to highlight that current MU and TT analysis does not come from a full consensus among TE vendors.

- MU and TT for >6GHz (band n96) are working assumption based on analysis of single TE vendor. Values will be revisited once analysis from other TE vendors is available.

Licensed band n104 is now introduced to RAN5 specs so MU and TT analysis up to 7.125GHz needs to be extended for all the test cases.

Observation 1. Existing MU and TT analysis up to 7.125GHz is limited to NR-U single carrier, single UL (SISO) transmission test cases, excluding the EVM test.

Observation 2. Existing MU and TT analysis up to 7.125GHz is generic, not precluding detailed updates per test case if further analysis is provided.

Observation 3. MU and TT analysis up to 7.125GHz needs to be extended for all the test cases now licensed band n104 is introduced to RAN5 specs.

## RF Single carrier tests

Based on the general rule proposed and endorsed for NR-U test cases in [1], same can be applied to the equivalent single carrier, single UL (SISO) transmission test cases.

Proposal 1. As a general rule, for the range 6.0GHz < f ≤ 7.125GHz, define the same MTSU already defined for f = 6.0GHz in the single carrier, single UL (SISO) transmission test cases. Exceptions in MTSU definition for specific test cases are not precluded.

In the following subsections, certain specific analysis will be done for the OFF power measurement, the EVM measurements and the maximum input level test.

#### OFF power measurement

- We have conducted a more detailed analysis of the OFF power test case and concluded that MU and TT at f=7.125GHz and bandwidth up to 100MHz should be increased by 0.3dB compared to the values currently defined for f=6GHz. If this proposal is endorsed, current MU and TT defined in the NR-U test cases 6.3F.2 (OFF) and 6.3F.3.2 (ON/OFF) should be updated accordingly.

Proposal 2. For the OFF power measurement, for 6.0GHz < f ≤ 7.125GHz and CBW ≤ 100MHz, define MU/TT increased by 0.3dB versus the value defined for f = 6GHz and CBW ≤ 100MHz.

#### EVM Measurement

- MU and TT for EVM test have never been defined for f > 6GHz. After evaluation of the performance in our test system, we propose the MU defined in Table 2.2.2-1 below. The proposed TT there are just calculated accordingly following the TT formula defined for the EVM test case in [5], Annes F.3.2.

| - Maximum Test System Uncertainty | - Test Tolerance (TT) |
| --- | --- |
| For up to 256QAM: f ≤ 6.0GHz, BW ≤ 100MHz 15 dBm < PUL PUSCH, PUCCH, PRACH: ±1.5 % -25 dBm < PUL ≤ 15 dBm PUSCH, PUCCH, PRACH: ±2.5 % -40dBm ≤ PUL ≤ -25dBm PUSCH, PUCCH, PRACH: ±3.0 % For up to 256QAM: 6GHz < f ≤ 7.125GHz, BW ≤ 100MHz 15 dBm < PUL PUSCH, PUCCH, PRACH: ±1.5 % -25 dBm < PUL ≤ 15 dBm PUSCH, PUCCH, PRACH: ±3.0 % -40dBm ≤ PUL ≤ -25dBm PUSCH, PUCCH, PRACH: ±3.9 % Absolute Uplink power measurement same as 6.3.1. Relative Uplink power measurement same as 6.3.4.3. | For up to 64QAM f ≤ 6.0GHz, BW ≤ 100MHz 0% 6GHz < f ≤ 7.125GHz, BW ≤ 100MHz 0%, -25dBm < PUL, 0.9%, -40dBm ≤ PUL ≤ -25dBm, 64QAM For 256QAM f ≤ 6.0GHz, BW ≤ 100MHz 0.3%, 15dBm < PUL 0.8%, -25dBm < PUL ≤ 15dBm, - 1.1%, -40dBm ≤ PUL ≤ -25dBm 6GHz < f ≤ 7.125GHz, BW ≤ 100MHz 0.3%, 15dBm < PUL 1.1%, -25dBm < PUL ≤ 15dBm, 1.74%, -40dBm ≤ PUL ≤ -25dBm |

- Table 2.2.2-1. MU and TT proposal for general EVM measurement.

Previous proposal and analysis are limited for the general EVM measurement. For the EVM measurement in transient periods, more time is required to complete the analysis.

Proposal 3. For general EVM measurement, for 6.0GHz < f ≤ 7.125GHz and CBW ≤ 100MHz, define MU and TT as captured in Table 2.2.2-1. For EVM measurement in transient periods, for 6.0GHz < f ≤ 7.125GHz MU and TT are FFS.

#### Maximum input level

- We have conducted a more detailed analysis of the maximum input level test case and concluded that MU and TT at f=7.125GHz and bandwidth up to 100MHz should be increased by 0.5dB compared to the values currently defined for f=6GHz. If this proposal is endorsed, current MU and TT defined in the NR-U test case 7.4F should be updated accordingly.

Proposal 4. For the maximum input level test, for 6.0GHz < f ≤ 7.125GHz and CBW ≤ 100MHz, define MU/TT increased by 0.5dB versus the value defined for f = 6GHz and CBW ≤ 100MHz.

## RF UL MIMO / TxD tests

More time is required to complete the MU and TT analysis for UL MIMO/TxD test cases for 6.0GHz < f ≤ 7.125GHz, noting that recently 4Tx UL MIMO/TxD tests are being added and that their MU/TT analysis even for f ≤ 6GHz is not complete yet.

## RF CA tests

More time is also required to conduct the MU/TT analysis for the carrier aggregation (CA) test cases for 6.0GHz < f ≤ 7.125GHz, noting that in case of CA the MU/TT analysis for intra-band aggregated bandwidth greater than 100MHz even for f ≤ 6GHz is not complete yet.

# Conclusion

The following observations and proposals were made in this contribution.

Observation 1. Existing MU and TT analysis up to 7.125GHz is limited to NR-U single carrier, single UL (SISO) transmission test cases, excluding the EVM test.

Observation 2. Existing MU and TT analysis up to 7.125GHz is generic, not precluding detailed updates per test case if further analysis is provided.

Observation 3. MU and TT analysis up to 7.125GHz needs to be extended for all the test cases now licensed band n104 is introduced to RAN5 specs.

Proposal 1. As a general rule, for the range 6.0GHz < f ≤ 7.125GHz, define the same MTSU already defined for f = 6.0GHz in the single carrier, single UL (SISO) transmission test cases. Exceptions in MTSU definition for specific test cases are not precluded.

Proposal 2. For the OFF power measurement, for 6.0GHz < f ≤ 7.125GHz and CBW ≤ 100MHz, define MU/TT increased by 0.3dB versus the value defined for f = 6GHz and CBW ≤ 100MHz.

Proposal 3. For general EVM measurement, for 6.0GHz < f ≤ 7.125GHz and CBW ≤ 100MHz, define MU and TT as captured in Table 2.2.2-1. For EVM measurement in transient periods, for 6.0GHz < f ≤ 7.125GHz MU and TT are FFS.

Proposal 4. For the maximum input level test, for 6.0GHz < f ≤ 7.125GHz and CBW ≤ 100MHz, define MU/TT increased by 0.5dB versus the value defined for f = 6GHz and CBW ≤ 100MHz.

# References

- R5-234170, “Derive MU for FR1 bands above 6GHz - AP97.21”, Keysight, RAN5#100

- R5-235850, “TT and editor note update in NR-U Tx test cases for FR1 bands above 6GHz”, Keysight, RAN5#100

- R5-235851, “TT and editor note update in NR-U Rx test cases for FR1 bands above 6GHz”, Keysight, RAN5#100

- R5-235852, “MU and TT definition for FR1 bands above 6GHz - Annex F update”, Keysight, RAN5#100

- TS 38.521-1, “User Equipment (UE) conformance specification; Radio transmission and reception; Part 1: Range 1 Standalone”, v19.0.0 (2025-06)

- TS 38.101-1, “User Equipment (UE); Radio transmission and reception; Part 1: Range 1 Standalone”, v18.10.0 (2025-06)