| 3GPP TR 26.921 V0.3.0 (2019-07) |  |
| --- | --- |
| Technical Report |  |
| 3rd Generation Partnership Project; Technical Specification Group SA; Investigations on ambient noise reproduction systems for acoustic testing of terminals (Release 16) |  |
| The present document has been developed within the 3rd Generation Partnership Project (3GPP TM) and may be further elaborated for the purposes of 3GPP. The present document has not been subject to any approval process by the 3GPP Organizational Partners and shall not be implemented. This Specification is provided for future development work within 3GPP only. The Organizational Partners accept no liability for any use of this Specification. Specifications and Reports for implementation of the 3GPP TM system should be obtained via the 3GPP Organizational Partners' Publications Offices. |  |

3GPP Postal address 3GPP support office address 650 Route des Lucioles - Sophia Antipolis Valbonne - FRANCE Tel.: +33 4 92 94 42 00 Fax: +33 4 93 65 47 16 Internet http://www.3gpp.org

Copyright Notification No part may be reproduced except as authorized by written permission. The copyright and the foregoing restriction extend to reproduction in all media. © 2019, 3GPP Organizational Partners (ARIB, ATIS, CCSA, ETSI, TSDSI, TTA, TTC). All rights reserved. UMTS™ is a Trade Mark of ETSI registered for the benefit of its members 3GPP™ is a Trade Mark of ETSI registered for the benefit of its Members and of the 3GPP Organizational Partners LTE™ is a Trade Mark of ETSI registered for the benefit of its Members and of the 3GPP Organizational Partners GSM® and the GSM logo are registered and owned by the GSM Association

Contents

Foreword	6

1	Scope	7

2	References	7

3	Definitions, symbols and abbreviations	7

3.1	Definitions	7

3.2	Symbols	8

3.3	 Abbreviations	8

4	Ambient Noise Testing in Hands-free Mode	8

4.1	 Introduction	8

4.2	 Results of Round Robin Test on Different Background Noise Simulation Techniques	8

4.2.1	Introduction	8

4.2.2	Test Setup	9

4.2.2.1	Laboratories	9

4.2.2.2	Noise field simulations	9

4.2.2.3	Noise Types	11

4.2.2.4	Devices	11

4.2.2.5	Speech Quality Test Methodology	12

4.2.3	Results for Handheld Hands-free	12

4.2.3.1	Introduction	12

4.2.3.2	Comparison of inter-lab accuracy for the different background noise simulation methods	12

4.2.3.2.1	Outliers	12

4.2.3.2.2	Wideband	13

4.2.3.2.2.1	No background noise	13

4.2.3.2.2.2	Background Noises & Simulation acc. to ES 202 396-1	15

4.2.3.2.2.3	Simulation acc. to ES 202 396-1 with binaural noises from TS 103 224	17

4.2.3.2.2.4	Background Noises & Simulation acc. to TS 103 224	19

4.2.3.2.3	Narrowband	21

4.2.3.2.3.1	No background noise	21

4.2.3.2.3.2	Background Noises & Simulation acc. to ES 202 396-1	23

4.2.3.2.3.3	Simulation acc. to ES 202 396-1 with binaural noises from TS 103 224	25

4.2.3.2.3.4	Background Noises & Simulation acc. to TS 103 224	27

4.2.3.3	Comparison of measurement results between ES 202 396-1 and TS 103 224 background noise simulation technique	29

4.2.3.3.1	Introduction	29

4.2.3.3.2	Wideband	29

4.2.3.3.2.1	Background Noises & Simulation acc. to ES 202 396-1 compared to TS 103 224 simulation	29

4.2.3.3.2.2	Simulation acc. to ES 202 396-1 with noises from TS 103 224 (binaural) compared to TS 103 224 simulation	31

4.2.3.3.3	Narrowband	33

4.2.3.3.3.1	Background Noises & Simulation acc. to ES 202 396-1 compared to TS 103 224 simulation	33

4.2.3.3.3.2	Simulation acc. to ES 202 396-1 with noises from TS 103 224 (binaural) compared to TS 103 224 simulation	35

4.2.3.4	Analysis results according to TS 26.132	37

4.2.3.4.1	Introduction	37

4.2.3.4.2	Wideband	37

4.2.3.4.2.1	Comparison of absolute Results	37

4.2.3.4.2.2	Differences between TS 103 224 and ES 202 396-1	38

4.2.3.4.2.3	Differences between TS 103 224 and ES 202 396-1 with noises from TS 103 224	38

4.2.3.4.3	Narrowband	39

4.2.3.4.3.1	Comparison of absolute Results	39

4.2.3.4.3.2	Differences between TS 103 224 and ES 202 396-1	40

4.2.3.4.3.3	Differences between TS 103 224 and ES 202 396-1 with noises from TS 103 224	41

4.2.3.5	Analyses of the noise spectra reproduced at the reference microphone	42

4.2.3.5.1	Introduction	42

4.2.3.5.2	Simulation acc. to ES 202 396-1 with recordings from TS 103 224	43

4.2.3.5.3	Simulation and Recordings acc. to TS 103 224	46

4.2.4	Results for Desktop Hands-free	50

4.2.4.1	Introduction	50

4.2.4.2	Comparison of inter-lab accuracy for the different background noise simulation methods	50

4.2.4.2.1	Wideband	50

4.2.4.2.1.1	No background noise	50

4.2.4.2.1.2	Simulation acc. to ES 202 396-1 with noises from TS 103 224 (binaural)	52

4.2.4.2.1.3	Background Noises & Simulation acc. to TS 103 224	54

4.2.4.2.2	Narrowband	56

4.2.4.2.2.1	No background noise	56

4.2.4.2.2.2	Simulation acc. to ES 202 396-1 with noises from TS 103 224 (binaural)	58

4.2.4.2.2.3	Background Noises & Simulation acc. to TS 103 224	60

4.2.4.3	Comparison between ES 202 396-1 and TS 103 224 background noise simulation technique	62

4.2.4.3.1	Introduction	62

4.2.4.3.2	Wideband	62

4.2.4.3.3	Narrowband	63

4.2.4.4	Analysis results according to TS 26.131 & TS 26.132	64

4.2.4.4.1	Introduction	64

4.2.4.4.2	Wideband	64

4.2.4.4.2.1	Comparison of absolute Results	64

4.2.4.4.2.2	Differences between TS 103 224 and ES 202 396-1 with noises from TS 103 224	65

4.2.4.4.3	Narrowband	66

4.2.4.4.3.1	Comparison of absolute Results	66

4.2.4.4.3.2	Differences between TS 103 224 and ES 202 396-1 with noises from TS 103 224	67

4.2.4.5	Analyses of the noise spectra reproduced at the reference microphone	68

4.2.4.5.1	Introduction	68

4.2.4.5.2	Simulation acc. to ES 202 396-1 with recordings from TS 103 224, reference recording at place of DUT-microphone	68

4.2.4.5.3	Simulation & Recordings acc. to TS 103 224, reference recording at place of DUT-microphone	69

5	Ambient Noise Testing in Handset Mode	70

5.1	 Introduction	70

5.2	Analysis of modified ambient noise playback systems	70

5.2.1	Motivation	70

5.2.2	Test Setup	71

5.2.2.1	Configurations	71

5.2.2.2	Evaluation of sound field	72

5.2.3	Noise Types	73

5.2.4	Measurement rooms	73

5.2.5	Results	73

5.2.6	Discussion	75

5.2.7	Conclusions	76

5.3	 Noise Field Simulations in different Labs	76

5.3.1	Introduction	76

5.3.2	Noise Field Simulations	76

5.3.2.1	System according to ES 202 396-1 (binaural noise sources)	76

5.3.2.2	System according to TS 103 224 (8-channel sources)	77

5.3.2.3	System according to TS 103 224 (binaural noise sources)	77

5.3.3	Noise Types	78

5.3.4	Measurement rooms	80

5.3.5	Test Methodology	86

5.3.5.1	Spectral Analysis	86

5.3.5.2	Error Metrics	86

5.3.6	Measurement Results	87

5.3.6.1	Overall Results	87

5.3.6.2	Results per Room	88

5.3.6.3	Results per Background Noise	91

5.3.6.4	Impact of Bandwidth	92

5.3.7	Discussion	95

6	Conclusions	96

6.1	Conclusions on handheld and desktop hands-free mode	96

6.2	Conclusions on handset mode	96

Annex A: Collection of reports of round robin test in HHHF mode	97

A.1	Introduction	97

A.2	List of documents	97

Annex B: Detailed results for modified ambient noise playback	98

B.1	Introduction	98

B.2	Configuration A	99

B.3	Configuration B	100

B.4	Configuration C	101

B.5	Configuration D	102

Annex <X> (informative): Change history	103

# Foreword

This Technical Report has been produced by the 3rd Generation Partnership Project (3GPP).

The contents of the present document are subject to continuing work within the TSG and may change following formal TSG approval. Should the TSG modify the contents of the present document, it will be re-released by the TSG with an identifying change of release date and an increase in version number as follows:

Version x.y.z

where:

x	the first digit:

1	presented to TSG for information;

2	presented to TSG for approval;

3	or greater indicates TSG approved document under change control.

y	the second digit is incremented for all changes of substance, i.e. technical enhancements, corrections, updates, etc.

z	the third digit is incremented when editorial only changes have been incorporated in the document.

# 1	Scope

The present document provides a collection of investigations on aspects of UE performance in the presence of ambient noise related to the 3GPP terminal testing specifications TS 26.132 [3] and/or performance requirements per TS 26.131 [2].

The evaluation of terminals in conjunction with speech signals and realistic background noise tests as well as performance requirements were added in Release 11 of these specifications. Increased signal processing capabilities facilitated more sophisticated noise reduction functionality in terminals already before this release. However, an incorrect or mistuned device may substantially impact speech quality, intelligibility/listening effort and user experience in general.

The present document focuses on the evaluation of terminals in handset and hands-free mode, but also other related aspects like, e.g., accuracy of noise field simulations are taken into account.

# 2	References

The following documents contain provisions which, through reference in this text, constitute provisions of the present document.

-	References are either specific (identified by date of publication, edition number, version number, etc.) or non-specific.

-	For a specific reference, subsequent revisions do not apply.

-	For a non-specific reference, the latest version applies. In the case of a reference to a 3GPP document (including a GSM document), a non-specific reference implicitly refers to the latest version of that document in the same Release as the present document.

[1]	3GPP TR 21.905: "Vocabulary for 3GPP Specifications".

[2]	3GPP TS 26.131: "Terminal acoustic characteristics for telephony; Requirements".

[3]	3GPP TS 26.132: "Speech and video telephony terminal acoustic test specification".

[4]	ETSI ES 202 396-1: "Speech quality performance in the presence of background noise; Part 1: Background noise simulation technique and background noise database".

[5]		ETSI TS 103 224: "A sound field reproduction method for terminal testing including a background noise database".

[6]	ETSI TS 103 106: "Speech Quality performance in the presence of background noise: Background noise transmission for mobile terminals - Objective test methods".

# 3	Definitions, symbols and abbreviations

## 3.1	Definitions

For the purposes of the present document, the terms and definitions given in 3GPP TR 21.905 [1] and the following apply. A term defined in the present document takes precedence over the definition of the same term, if any, in 3GPP TR 21.905 [1].

Definition format (Normal)

<defined term>: <definition>.

example: text used to clarify abstract rules by applying them literally.

## 3.2	Symbols

For the purposes of the present document, the following symbols apply:

Symbol format (EW)

<symbol>	<Explanation>

## 3.3		Abbreviations

For the purposes of the present document, the abbreviations given in 3GPP TR 21.905 [1] and the following apply. An abbreviation defined in the present document takes precedence over the definition of the same abbreviation, if any, in 3GPP TR 21.905 [1].

DUT	Device Under Test

G-MOS	Global MOS (overall quality)

HHHF	Handheld Hands-free

N-MOS	Noise MOS (intrusiveness of noise)

NB	Narrowband telephonometry

MOS	Mean Opinion Score

RMSE	Root-Mean-Square Error

S-MOS	Speech MOS (distortion of speech)

SWB	Super-wideband telephonometry

WB	Wideband telephonometry

# 4	Ambient Noise Testing in Hands-free Mode

Editor’s Note: Some further information about the Round Robin results is expected.

## 4.1		Introduction

The investigations described in the following clauses refer to ambient noise testing of terminals in handheld and desktop hands-free mode.

## 4.2		Results of Round Robin Test on Different Background Noise Simulation Techniques

### 4.2.1	Introduction

The following clauses summarize the results of a Round Robin test conducted between April and August 2015. Besides providing the results, the reproduction accuracy of the different simulation methods in different labs is analysed. Furthermore, the impact of the different analysis methods on the results measured with different mobile terminals is shown. Corresponding meeting temporary documents are provided in Annex A.

The aim of the Round Robin test conducted under the ATeMPO_SPINE work item was mainly to answer the following questions:

	How good is the reproducibility of the different methods?

	Are there differences in the measured performance when using the different devices with different setups?

The following labs were participating in the test:

	Audience Inc. / Knowles Inc.

	HEAD acoustics GmbH

	Orange

	Sony Mobile Communications

In the Round Robin test, seven different devices were used, some labs tested only a subset of six devices as agreed upon in advance. Not all labs conducted the full set of tests in narrowband (NB) and wideband (WB). At HEAD acoustics premises, two rooms were used to conduct the tests; in all other labs, the tests were conducted in one room.

The two noise field simulations according to ETSI ES 202 396-1 [4] and ETSI TS 103 224 [5] were investigated for each lab.

For measurements using the background playback according to ES 202 396-1 [4], occasional playback problems were observed in lab 2.1 (see also clause 4.2.3.2.1), which explains the outliers. The overall tendencies observed in the complete data set are not strongly affected by these issues. On the contrary, it can be noted that for this lab, MOS results were in general consistent between playback methods when using the same noise scenarios, while other labs results showed a slight unexpected offset between the methods. Occasional playback problems were also observed in lab 3.1 (see attached Annex A).

### 4.2.2	Test Setup

#### 4.2.2.1	Laboratories

The labs according to Table 1 participated in the round robin test. Typical room acoustic parameters are provided as well here.

Table 1: Laboratories of the round robin test

| Label | Lab | HATS Rotation [°] | Rev. Time (RT60) | Clarity Index (C80) |
| --- | --- | --- | --- | --- |
| Lab 1.1 | HEAD acoustics Room 1 | 0 | 125 ms | 37.1 dB |
| Lab 1.2 | HEAD acoustics Room 2 | 0 | 240 ms | 20.5 dB |
| Lab 2.1 | Sony | 22.5 | 139 ms | 30.0 dB |
| Lab 3.1 | Audience | 22.5 | 117 ms | 29.2 dB |
| Lab 4.1 | Orange | 0 | 89 ms | 40 dB |
| Lab 4.2 | Orange | 22.5 | 89 ms | 40 dB |

NOTE 1:	For some labs, a rotation of 22.5° of the HATS was chosen in order to avoid obstructions of the sound (at rotation 0°, the HATS is exactly between a certain loudspeaker and the DUT).

NOTE 2:	Lab 4.1 and 4.2 represent identical measurement chambers and describe different rotations of the HATS. In some of the following clauses, results may be presented either for both or only one of the setups.

#### 4.2.2.2	Noise field simulations

In order to minimize all variabilities coming from other sources than the different setups and rooms used, the test material and the test procedure was prepared by HEAD acoustics and provided to all test labs. All labs used the HEAD acoustics software ACQUA with the HAE-BGN and 3PASS background noise systems with the identical database. All measurements were collected in a final database containing the results of all labs. Based on this procedure, all types of analyses as needed for the Round Robin experiment could be performed.

The selection of the rooms was up to the test labs. However, the requirements as found in [4] and [5] were followed. The detailed test setup is described in the individual reports of the test labs (see Annex A).

Figure 1 and Figure 2 show the general test setup for the two background noise simulation methods.

Figure 1: General setup of the 4+1 loudspeaker arrangement according to ES 202 396-1

Figure 2: General setup of the 8-loudspeaker arrangement according to TS 103 224

The calibration of the setups was made by each lab individually. The calibration procedures as described in [4] and [5] are implemented in the HAE-BGN and 3PASS software and were followed by each lab. The calibration results can be found in the individual test reports of the test labs (see Annex A)..

#### 4.2.2.3	Noise Types

For testing with the noise field simulation according to ETSI ES 202 396-1 [4], the six binaural noises according to Table 2 were used.

Table 2: Noise types for playback acc. to ES 202 396-1

| Alias | Filename in database of [4] |
| --- | --- |
| Pub | Pub_Noise_binaural_V2 |
| Crossroad | Outside_Traffic_Crossroads_binaural |
| Trainstation | Train_Station_binaural |
| Inside Car | Fullsize_Car1_130Kmh_binaural |
| Cafeteria | Cafeteria_Noise_binaural |
| Office | Work_Noise_Office_Callcenter_binaural |

For testing with the noise field simulation according to ETSI TS 103 224 [5], the six noises according to Table 3 were used. The corresponding eight-channel recordings are also binaurally available in the noise database, which allows the usage of both systems with the identical noise type. Each of these noises was selected as a counterpart to a binaural recording listed in Table 2. Comparisons between the counterparts can be made based on the alias columns in both tables.

Table 3: Noise types for playback acc. to TS 103 224

| Alias | Filename in database of [5] |  |
| --- | --- | --- |
|  | 8-channel | Binaural |
| Pub | Pub_handsfree | Pub_bin |
| Crossroad | Crossroadnoise_handsfree | Crossroadnoise_bin |
| Trainstation | Trainstation_handsfree | Trainstation_bin |
| Inside Car | FullSizeCar_130_handsfree | FullSizeCar_130_bin |
| Cafeteria | Cafeteria_handsfree | Cafeteria_bin |
| Office | Callcenter2_handsfree | Callcenter2_bin |

#### 4.2.2.4	Devices

Seven devices under test (DUT) according to Table 4 were available for the round robin test. Six phones were tested in all labs.

Table 4: Devices under test

| Name | Size |
| --- | --- |
| DUT1 | 138.1 x 67 x 6.9 mm |
| DUT2 | 143.4 x 70.5 x 6.8 mm |
| DUT3 | 138.5 x 70.9 x 8.9 mm |
| DUT4 | 162.8 x 85.4 x 8.7 mm |
| DUT5 | 127.3 x 64.9 x 8.6 mm |
| DUT6 | 150.1 x 72.7 x 9.6 mm |
| DUT7 | 157.7 x 78.7 x 7.7 mm |

#### 4.2.2.5	Speech Quality Test Methodology

As source speech material for the evaluation of speech quality in the presence of background noise, 16 American English speech samples (fullband) according to ETSI TS 103 106 [6] were used for both bandwidth modes.

For terminals evaluated in NB mode, the analysis according to ETSI TS 103 106 [6] in NB-mode was carried out for each of the 16 samples and the results were averaged per background noise condition.

For terminals evaluated in WB mode, the analysis according to ETSI TS 103 106 [6] in WB-mode was carried out for each of the 16 samples and the results were averaged per background noise condition.

This usage of the instrumental assessment method provides S-MOS, N-MOS and G-MOS for each lab, bandwidth mode, device, and background noise.

### 4.2.3	Results for Handheld Hands-free

#### 4.2.3.1	Introduction

In the following clauses, results of several measurements and analyses are presented for terminals in handheld hands-free mode. Without loss of generality, the results of lab 2.1 (one of the labs that conducted all experiments with all devices) are always compared to the results for the other labs. The test setup complies with clause 5.1.3.3 of [3]. Possible differences in results across labs may be influenced and explained by:

	Calibration differences

	Setup differences

	Room differences

	Time variant behaviour of the device under test

#### 4.2.3.2	Comparison of inter-lab accuracy for the different background noise simulation methods

##### 4.2.3.2.1	Outliers

Four outliers could be observed which could clearly be traced back to an error in the measurement setup and therefore do not represent any characteristics of the background noise generation method used. All diagrams in clause 4.2.3.1 (and corresponding sub-clauses) still contain these outliers, while the RMSE-values are calculated without outliers. The outliers are listed in Table 5.

Table 5: Outliers removed from analysis

| Bandwidth | System | Noise Type | Lab | DUT |
| --- | --- | --- | --- | --- |
| NB | ES 202 396-1 | Inside Car | Lab 2.1 | DUT2 |
| NB | ES 202 396-1 with noises from TS 103 224 | Inside Car | Lab 2.1 | DUT2 |
| WB | ES 202 396-1 with noises from TS 103 224 | Inside Car | Lab 3.1 | DUT2 |
| WB | TS 103 224 | Office | Lab 4.2 | DUT7 |

##### 4.2.3.2.2	Wideband

###### 4.2.3.2.2.1	No background noise

In this clause, results under silent conditions in WB mode are presented, as shown in Figure 3 to Figure 5. Basically, the variance to be expected in different labs with no background noise simulation present can be observed.

It seems that the parameters described in clause 4.2.3.1 may impact the results to a certain extent providing the basis for the interpretation of the experiments with the background noise simulation methodologies. The main impact is on S-MOS resulting in somewhat scattered G-MOS results as well.

Figure 3: Correlation between G-MOS (WB) results from Lab 2.1 and other labs (Silence)

Figure 4: Correlation between S-MOS (WB) results from Lab 2.1 and other labs (Silence)

Figure 5: Correlation between N-MOS (WB) results from Lab 2.1 and other labs (Silence)

###### 4.2.3.2.2.2	Background Noises & Simulation acc. to ES 202 396-1

The results shown in Figure 6 to Figure 8 are based on using the ES 202 396-1 noise field simulation, using the binaurally recorded background noises from the associated noise database. The following observations can be made:

	Compared to Lab 2.1 the results measured in the other labs are in general about 0.2 to 0.3 MOS higher for N-MOS and G-MOS

	The results are somewhat scattered leading to RMSE values between 0.13 and 0.22

	N-MOS shows the highest variation in results up to about 0.5 MOS

Figure 6: Correlation between G-MOS (WB) between Lab 2.1 and other labs

Figure 7: Correlation between S-MOS (WB) between Lab 2.1 and other labs

Figure 8: Correlation between N-MOS (WB) between Lab 2.1 and other labs

###### 4.2.3.2.2.3	Simulation acc. to ES 202 396-1 with binaural noises from TS 103 224

The results shown in Figure 9 to Figure 11 are based on using the ES 202 396-1 noise field simulation, but using the binaurally recorded background noises from TS 103 224. The following observations can be made:

	Compared to Lab 2.1, the results measured in the other labs are in general about 0.3 to 0.4 MOS higher for S-MOS, N-MOS and G-MOS

	The results are somewhat scattered leading to RMSE values between 0.17 and 0.24

	N-MOS shows the highest variation in results up to about 0.5 MOS

Please note that the indicated outlier for test lab 3.1 is excluded from the RMSE calculations.

Outlier in Lab 3.1

Figure 9: Correlation between G-MOS (WB) between Lab 2.1 and other labs

Figure 10: Correlation between S-MOS (WB) between Lab 2.1 and other labs

Outlier in Lab 3.1

Figure 11: Correlation between N-MOS (WB) between Lab 2.1 and other labs

###### 4.2.3.2.2.4	Background Noises & Simulation acc. to TS 103 224

The results shown in Figure 12 to Figure 14 are based on using the TS 103 224 noise field simulation, using the 8-channel noise recording from the associated noise database. For this setup, the following observations can be made:

	All results line up fairly well, no offset between the labs can be observed

	The results are less scattered leading to low RMSE values of 0.1

	Especially N-MOS is measured very consistent between labs

Figure 12: Correlation between G-MOS (WB) between Lab 2.1 and other labs

Figure 13: Correlation between S-MOS (WB) between Lab 2.1 and other labs

Figure 14: Correlation between N-MOS (WB) between Lab 2.1 and other labs

##### 4.2.3.2.3	Narrowband

###### 4.2.3.2.3.1	No background noise

In this clause, results under silent conditions in NB mode are presented, as shown in Figure 15 to Figure 17. Basically, the variance to be expected in different labs with no background noise simulation present can be observed.

It seems that the parameters described in clause 4.2.3.1 may impact the results to a certain extent providing the basis for the interpretation of the experiments with the background noise simulation methodologies. The impact on S-MOS and G-MOS is less pronounced than in wideband mode.

Figure 15: Correlation between G-MOS (NB) between Lab 2.1 and other labs (Silence)

Figure 16: Correlation between S-MOS (NB) between Lab 2.1 and other labs (Silence)

Figure 17: Correlation between N-MOS (NB) between Lab 2.1 and other labs (Silence)

###### 4.2.3.2.3.2	Background Noises & Simulation acc. to ES 202 396-1

The results shown in Figure 18 to Figure 20 are based on using the ES 202 396-1 noise field simulation, using the binaurally recorded background noises from the associated noise database. The following observations can be made:

	Compared to Lab 2.1 the results measured in the other labs are in general about 0.2 to 0.3 MOS higher for N-MOS and G-MOS

	The results are somewhat scattered leading to RMSE values between 0.15 and 0.21

	N-MOS shows the highest variation in results up to about 0.5 MOS

Note that the indicated outlier for test lab 2.1 is excluded from the RMSE calculations.

Figure 18: Correlation between G-MOS (NB) results from Lab 2.1 and other labs

Figure 19: Correlation between S-MOS (NB) results from Lab 2.1 and other labs

Outlier in Lab 2.1

Figure 20: Correlation between N-MOS (NB) results from Lab 2.1 and other labs

###### 4.2.3.2.3.3	Simulation acc. to ES 202 396-1 with binaural noises from TS 103 224

The results shown Figure 21 to Figure 23 are based on using the ES 202 396-1 noise field simulation, but using the binaurally recorded background noises from TS 103 224. The following observations can be made:

	Compared to Lab 2.1, the results measured in the other labs are in general about 0.3 to 0.4 MOS higher for S-MOS, N-MOS and G-MOS

	The results are somewhat scattered leading to RMSE values between 0.22 and 0.25

	N-MOS shows the highest variation in results up to about 0.5 MOS

	S-MOS shows a high variation of the results for Pub-noise and Train station noise

Note that the indicated outlier for test lab 2.1 is excluded from the RMSE calculations.

Figure 21: Correlation between G-MOS (NB) between Lab 2.1 and other labs

Figure 22: Correlation between S-MOS (NB) between Lab 2.1 and other labs

Outlier in Lab 2.1

Figure 23: Correlation between N-MOS (NB) between Lab 2.1 and other labs

###### 4.2.3.2.3.4	Background Noises & Simulation acc. to TS 103 224

The results shown in this section are based on using the TS 103 224 noise field simulation, using the 8-channel noise recording from the associated noise database. For this setup, the following observations can be made:

	All results line up fairly well, no offset between the labs can be observed

	The results are less scattered leading to low RMSE values between 0.11 and 0.14

	Especially the measured N-MOS values are very consistent

Figure 24: Correlation between G-MOS (NB) results from Lab 2.1 and other labs

Figure 25: Correlation between S-MOS (NB) results from Lab 2.1 and other labs

Figure 26: Correlation between N-MOS (NB) results from Lab 2.1 and other labs

#### 4.2.3.3	Comparison of measurement results between ES 202 396-1 and TS 103 224 background noise simulation technique

##### 4.2.3.3.1	Introduction

In this section, S-, N-, and G-MOS as measured in the different rooms for different terminals are analysed. The comparison is focused on the differences between background noise simulation systems. The comparison is carried out between each method and the different sets of background noises. The relation between the results measured using the ETSI ES 202 396-1 background noise simulation system and the results obtained using the ETSI TS 103 224 simulation system are presented for all noises, all labs, and all terminals.

The results of the different devices are colour-coded as indicated in the following figures.

##### 4.2.3.3.2	Wideband

###### 4.2.3.3.2.1	Background Noises & Simulation acc. to ES 202 396-1 compared to TS 103 224 simulation

While for crossroad, inside car office and pub noise the measurements lead to quite comparable results, deviations are visible for the cafeteria and train station noise. The reason for these differences is the different nature of the noises in the different noise databases. E.g. the cafeteria noise is in average about 6 dB lower in level for the recordings in ES 202 396-1, the train station noise is about 7 dB lower in level.

Figure 27: Correlation of G-MOS (WB) results between TS 103 224 and ES 202 396-1 methodology

Figure 28: Correlation of S-MOS (WB) results between TS 103 224 and ES 202 396-1 methodology

Figure 29: Correlation of N-MOS (WB) results between TS 103 224 and ES 202 396-1 methodology

###### 4.2.3.3.2.2	Simulation acc. to ES 202 396-1 with noises from TS 103 224 (binaural) compared to TS 103 224 simulation

When choosing the same type of recordings by taking the binaurally recorded noises from TS 103 224, the differences between the methods become smaller and results line up in general. More variation can be seen for the N-MOS results than for the S-MOS results. This is clearly indicating the difference in the background noise simulation method. The S-MOS seems to be mainly determined by the speech signal which is unchanged in the experiments.

Figure 30: Correlation of G-MOS (WB) results between TS 103 224 and ES 202 396-1 methodology (noises from TS 103 224)

Figure 31: Correlation of S-MOS (WB) results between TS 103 224 and ES 202 396-1 methodology (noises from TS 103 224)

Figure 32: Correlation of N-MOS (WB) results between TS 103 224 and ES 202 396-1 methodology (noises from TS 103 224)

##### 4.2.3.3.3	Narrowband

###### 4.2.3.3.3.1	Background Noises & Simulation acc. to ES 202 396-1 compared to TS 103 224 simulation

While for crossroad, inside car office and pub noise the measurements lead to quite comparable results, deviations are visible for the cafeteria and train station noise. The reason for these differences is the different nature of the noises in the different noise databases. E.g. the cafeteria noise is in average about 6 dB lower in level for the recordings in ES 202 396-1, the train station noise is about 7 dB lower in level.

Figure 33: Correlation of G-MOS (NB) results between TS 103 224 and ES 202 396-1 methodology

Figure 34: Correlation of S-MOS (NB) results between TS 103 224 and ES 202 396-1 methodology

Figure 35: Correlation of N-MOS (NB) results between TS 103 224 and ES 202 396-1 methodology

###### 4.2.3.3.3.2	Simulation acc. to ES 202 396-1 with noises from TS 103 224 (binaural) compared to TS 103 224 simulation

When choosing the same type of recordings by taking the binaurally recorded noises from TS 103 224, the differences between the methods become smaller and results line up in general. More variation can be seen for the N-MOS results than for the S-MOS results. This is clearly indicating the difference in the background noise simulation method. The S-MOS seems to be mainly determined by the speech signal which is unchanged in the experiments.

Figure 36: Correlation of G-MOS (NB) results between TS 103 224 and ES 202 396-1 methodology (noises from TS 103 224)

Figure 37: Correlation of S-MOS (NB) results between TS 103 224 and ES 202 396-1 methodology (noises from TS 103 224)

Figure 38: Correlation of N-MOS (NB) results between TS 103 224 and ES 202 396-1 methodology (noises from TS 103 224)

#### 4.2.3.4	Analysis results according to TS 26.132

##### 4.2.3.4.1	Introduction

In this clause, the speech quality analysis as specified in [3] and [2] is carried out. Averaging the results across all noise conditions provides one single S-MOS, N-MOS and G-MOS for each device.

##### 4.2.3.4.2	Wideband

###### 4.2.3.4.2.1	Comparison of absolute Results

The results in Figure 39 show the same tendency for all phones when measured using the different background noise simulation setups. The rank order in performance remains unchanged.

Figure 39: S-/N-/G-MOS (WB) averaged across background noises and all labs

###### 4.2.3.4.2.2	Differences between TS 103 224 and ES 202 396-1

Differences in the absolute values when comparing ES 202 396-1 with TS 103 224 can be seen for all terminals (see Figure 40). The offset is not constant but depending on the type of terminal. The main difference is found in the N-MOS values indication the different sensitivity of individual terminals to the different types of noise. The performance for all terminals is better when using the ES 202 396-1. Again, this is due to the less stressing noises in this standard.

| G-MOS | S-MOS |
| --- | --- |
| N-MOS |  |

Figure 40: Differences (WB) between TS 103 224 and ES 202 396-1 methodology

###### 4.2.3.4.2.3	Differences between TS 103 224 and ES 202 396-1 with noises from TS 103 224

When using the binaurally recorded background noises from TS 103 224 for testing, the difference between the two systems gets much smaller (see Figure 41). Again, the offset is not constant, but depends on the type of terminal. The performance for all terminals is slightly better when using the ES 202 396-1 simulation method compared to the sound field simulation technique described in TS 103 224.

| G-MOS | S-MOS |
| --- | --- |
| N-MOS |  |

Figure 41: Differences (WB) between TS 103 224 and ES 202 396-1 methodology 
(binaural noises from TS 103 224 database)

##### 4.2.3.4.3	Narrowband

###### 4.2.3.4.3.1	Comparison of absolute Results

The results in Figure 42 show the same tendency for all phones when measured using the different background noise simulation setups. The rank order in performance remains unchanged.

Figure 42: S-/N-/G-MOS (NB) averaged across background noises and all labs

###### 4.2.3.4.3.2	Differences between TS 103 224 and ES 202 396-1

Differences in the absolute values when comparing ES 202 396-1 with TS 103 224 can be seen for all terminals (see Figure 43). The offset is not constant but depending on the type of terminal. The differences are found in S-, N-, and G-MOS values. As in wideband the performance for all terminals is better when using the ES 202 396-1. Again, this is due to the less stressing noises in this standard.

| G-MOS | S-MOS |
| --- | --- |
| N-MOS |  |

Figure 43: Differences (NB) between TS 103 224 and ES 202 396-1 methodology

###### 4.2.3.4.3.3	Differences between TS 103 224 and ES 202 396-1 with noises from TS 103 224

When using the binaurally recorded background noises from TS 103 224 for testing, the difference between the two systems gets much smaller (see Figure 44). Again, the offset is not constant, but depends on the type of terminal and the performance for all terminals is slightly better when using the ES 202 396-1 simulation method compared to the sound field simulation technique described in TS 103 224.

| G-MOS | S-MOS |
| --- | --- |
| N-MOS |  |

Figure 44: Differences (NB) between TS 103 224 and ES 202 396-1 methodology 
(binaural noises from TS 103 224 database)

#### 4.2.3.5	Analyses of the noise spectra reproduced at the reference microphone

##### 4.2.3.5.1	Introduction

The background noise spectra measured at the reference microphone (input signal for the analysis according to TS 103 106 [6]) of all devices under test are plotted into one diagram for all background noises which were used in this test. Each diagram contains background noise spectra from

	6 DUTs from Lab 1 (DUT7 was not measured)

o	Included twice, since two different rooms were available

	7 DUTs from Lab 2

	7 DUTs from Lab 3

	7 DUTs from Lab 4

o	Included twice, since two different setups were evaluated

##### 4.2.3.5.2	Simulation acc. to ES 202 396-1 with recordings from TS 103 224

For all investigated background noises, quite large differences can be noticed in the spectra reproduced by the ES 202 396-1 simulation method. The differences depend to some extent on the noise type, as shown in Figure 45 to Figure 50. More uniform background noises, such as e.g., car, show less spectral variance than e.g., train station (which performs worst).

The differences in magnitude range from 5 dB up to 15 dB. The larger differences are mostly located in the low frequency domain.

Figure 45: 1/3rd octave spectra at reference microphone and Cafeteria noise from TS 103 224 
(ES 202 396-1 methodology)

Figure 46: 1/3rd octave spectra at reference microphone and Crossroad noise from TS 103 224 
(ES 202 396-1 methodology)

Figure 47: 1/3rd octave spectra at reference microphone and Inside Car noise from TS 103 224 
(ES 202 396-1 methodology)

Figure 48: 1/3rd octave spectra at reference microphone and Office noise from TS 103 224 
(ES 202 396-1 methodology)

Figure 49: 1/3rd octave spectra at reference microphone and Pub noise from TS 103 224 
(ES 202 396-1 methodology)

Figure 50: 1/3rd octave spectra at reference microphone and Train Station noise from TS 103 224 
(ES 202 396-1 methodology)

##### 4.2.3.5.3	Simulation and Recordings acc. to TS 103 224

The differences in the spectra reproduced by the TS 103 224 simulation method are significantly lower compared to the ones of ES 202 396-1, as shown in Figure 51 to Figure 56. The differences seem also to be independent of the noise type.

The differences range from 1 dB up to 10 dB. As expected, the sound field reproduction is highly accurate and consistent across labs in the low frequency domain up to about 2 kHz (where most energy of the noise is found). Here the spectral differences are within a range of 1-2 dB. Up to about 8 kHz, the differences are still less than 3 dB and from 10 - 20 kHz, the differences mostly remain below 5 dB.

Figure 51: 1/3rd octave spectra at reference microphone and Cafeteria noise from TS 103 224 
(TS 103 224 methodology)

Figure 52: 1/3rd octave spectra at reference microphone and Crossroad noise from TS 103 224 
(TS 103 224 methodology)

Figure 53: 1/3rd octave spectra at reference microphone and Inside Car noise from TS 103 224 
(TS 103 224 methodology)

Figure 54: 1/3rd octave spectra at reference microphone and Office noise from TS 103 224 
(TS 103 224 methodology)

Figure 55: 1/3rd octave spectra at reference microphone and Pub noise from TS 103 224 
(TS 103 224 methodology)

Figure 56: 1/3rd octave spectra at reference microphone and Train Station noise from TS 103 224 
(TS 103 224 methodology)

### 4.2.4	Results for Desktop Hands-free

#### 4.2.4.1	Introduction

In the following clauses, results of several measurements and analyses are presented for terminals in desktop hands-free mode. The test setup complies with clause 5.1.3.2 of [3]. Similar as for the handheld hands-free setup, possible differences in results across labs may be influenced and explained by:

	Calibration differences

	Setup differences

	Room differences

	Time variant behaviour of the device under test

It seems that these parameters may have impact on the results in a similar or even bigger range than the experiments which include the background noise simulation.

#### 4.2.4.2	Comparison of inter-lab accuracy for the different background noise simulation methods

##### 4.2.4.2.1	Wideband

###### 4.2.4.2.1.1	No background noise

In this clause, results under silent conditions in WB mode are presented, as shown in Figure 57 to Figure 59. Basically, the variance to be expected in different labs with no background noise simulation present can be observed, as shown in Figure 57 to Figure 59.

It seems that these parameters described in clause 4.2.4.1 may have impact on the results in a similar or even bigger range than the experiments which include the background noise simulation.

Figure 57: Correlation between G-MOS (WB) results from Lab 2.1 and other labs (Silence)

Figure 58: Correlation between S-MOS (WB) results from Lab 2.1 and other labs (Silence)

Figure 59: Correlation between N-MOS (WB) results from Lab 2.1 and other labs (Silence)

###### 4.2.4.2.1.2	Simulation acc. to ES 202 396-1 with noises from TS 103 224 (binaural)

The results shown in Figure 66 to Figure 68 are based on using the ES 202 396-1 noise field simulation, but using the binaurally recorded background noises from TS 103 224. The following observations can be made:

	No systematic differences in results of other labs compared to Lab 2.1 can be observed.

	The RMSE values for S-, N- and G-MOS are 0.12.

Figure 60: Correlation between G-MOS (WB) results from Lab 2.1 and other labs

Figure 61: Correlation between S-MOS (WB) results from Lab 2.1 and other labs

Figure 62: Correlation between N-MOS (WB) results from Lab 2.1 and other labs

###### 4.2.4.2.1.3	Background Noises & Simulation acc. to TS 103 224

The results shown in Figure 63 to Figure 65 are based on using the TS 103 224 noise field simulation, using the 8-channel noise recording from the associated noise database. For this setup, the following observations can be made:

	No systematic differences in results of other labs compared to Lab 2.1 can be observed.

	The RMSE values for S-, N- and G-MOS are 0.07 – 0.17

	The lowest variation can be seen for N-MOS, the results are quite consistent

Figure 63: Correlation between G-MOS (WB) between Lab 2.1 and other labs

Figure 64: Correlation between S-MOS (WB) between Lab 2.1 and other labs

Figure 65: Correlation between N-MOS (WB) between Lab 2.1 and other labs

##### 4.2.4.2.2	Narrowband

###### 4.2.4.2.2.1	No background noise

In this clause, results under silent conditions in NB mode are presented, as shown in Figure 66 to Figure 68. Basically, the variance to be expected in different labs with no background noise simulation present can be observed, as shown in Figure 66 to Figure 68.

It seems that these parameters described in clause 4.2.4.1 may have impact on the results in a similar or even bigger range than the experiments which include the background noise simulation. The RMSE of the results is quite high, mainly influences by S-MOS.

Figure 66: Correlation between G-MOS (NB) results from Lab 2.1 and other labs (Silence)

Figure 67: Correlation between G-MOS (NB) results from Lab 2.1 and other labs (Silence)

Figure 68: Correlation between G-MOS (NB) results from Lab 2.1 and other labs (Silence)

###### 4.2.4.2.2.2	Simulation acc. to ES 202 396-1 with noises from TS 103 224 (binaural)

The results shown in Figure 69 to Figure 71 are based on using the ES 202 396-1 noise field simulation, but using the binaurally recorded background noises from TS 103 224. The following observations can be made:

	No systematic differences in results of other labs compared to Lab 2.1 can be observed.

	The RMSE values for S-, N- and G-MOS are ranging from 0.17 to 0.22.

	The results are less consistent between labs than in wideband.

Figure 69: Correlation between G-MOS (NB) results from Lab 2.1 and other labs

Figure 70: Correlation between S-MOS (NB) results from Lab 2.1 and other labs

Figure 71: Correlation between N-MOS (NB) results from Lab 2.1 and other labs

###### 4.2.4.2.2.3	Background Noises & Simulation acc. to TS 103 224

The results shown in Figure 72 to Figure 74 are based on using the TS 103 224 noise field simulation, using the 8-channel noise recording from the associated noise database. For this setup, the following observations can be made:

	No systematic differences in results of other labs compared to Lab 2.1 can be observed.

	The RMSE values for S-, N- and G-MOS are in the range of 0.13 – 0.2.

	The lowest variation can be seen for N-MOS, the results are quite consistent.

Figure 72: Correlation between G-MOS (NB) between Lab 2.1 and other labs

Figure 73: Correlation between S-MOS (NB) between Lab 2.1 and other labs

Figure 74: Correlation between N-MOS (NB) between Lab 2.1 and other labs

#### 4.2.4.3	Comparison between ES 202 396-1 and TS 103 224 background noise simulation technique

##### 4.2.4.3.1	Introduction

In this section, S-, N-, and G-MOS as measured in the different rooms for different terminals is analysed. The comparison is focused on the differences between background noise simulation systems. The comparison is carried out between each method and the different sets of background noises. The results measured using the ETSI ES 202 396-1 background noise simulation system are used as the reference (all noises, all labs, and all terminals). The results obtained using the ETSI TS 103 224 simulation system are plotted on the abscissa (x-axis) for comparison.

In contrast to the related analysis for handheld hands-free devices (see clause 4.2.3.3), only one noise type according to TS 103 224 is used (Office/Callcenter). The results of the different devices are colour-coded as indicated in the following figures.

##### 4.2.4.3.2	Wideband

The results are shown in Figure 75. In general, a fairly good agreement between the background noise simulations can be noticed for S-MOS. Some deviations of the averaged results for N-MOS leads to subsequent deviations in G-MOS as well.

| G-MOS | S-MOS |
| --- | --- |
| N-MOS |  |

Figure 75: Correlation of G-MOS (WB) results between TS 103 224 and ES 202 396-1 methodology (noise from TS 103 224)

##### 4.2.4.3.3	Narrowband

Similar as in wideband mode, a fairly good agreement between the background noise simulations can be observed for S-MOS in Figure 76. Some deviation of the averaged results in N-MOS again leads to subsequent deviations in G-MOS as well.

| G-MOS | S-MOS |
| --- | --- |
| N-MOS |  |

Figure 76: Correlation of G-MOS (NB) results between TS 103 224 and ES 202 396-1 methodology (noise from TS 103 224)

#### 4.2.4.4	Analysis results according to TS 26.131 & TS 26.132

##### 4.2.4.4.1	Introduction

In this clause, the speech quality analysis as specified in [3] and [2] is carried out. Averaging the results across all noise conditions provides one single S-MOS, N-MOS and G-MOS for each device.

##### 4.2.4.4.2	Wideband

###### 4.2.4.4.2.1	Comparison of absolute Results

The results in Figure 77 show the same tendency for all phones when measured using the different background noise simulation setups. The rank order in performance remains unchanged.

Figure 77: S-/N-/G-MOS (WB) averaged across background noises and all labs

###### 4.2.4.4.2.2	Differences between TS 103 224 and ES 202 396-1 with noises from TS 103 224

When using the binaurally recorded background noises from TS 103 224 for the tests the differences when comparing ES 202 396-1 with TS 103 224 (see Figure 78) are slightly higher than the deviations observed in HHHF mode (see Figure 41). The offset is not constant, but depends on the type of terminal. Except for DUT2, the performance for all terminals is slightly better when using the ES 202 396-1 simulation method compared to the sound field simulation technique described in TS 103 224.

Figure 78: Differences (WB) between TS 103 224 and ES 202 396-1 methodology 
(binaural noises from TS 103 224 database)

##### 4.2.4.4.3	Narrowband

###### 4.2.4.4.3.1	Comparison of absolute Results

The results in Figure 79 show the same tendency for all phones when measured using the different background noise simulation setups. The rank order in performance remains unchanged.

Figure 79: S-/N-/G-MOS (NB) averaged across background noises and all labs

###### 4.2.4.4.3.2	Differences between TS 103 224 and ES 202 396-1 with noises from TS 103 224

When using the binaurally recorded background noises from TS 103 224 for the tests, the differences when comparing ES 202 396-1 with TS 103 224 (see Figure 80) are slightly higher than the deviations observed in HHHF mode (see Figure 44). The offset is not constant, but depends on the type of terminal. Except for DUT2, the performance for all terminals is slightly better when using the ES 202 396-1 simulation method compared to the sound field simulation technique described in TS 103 224.

Figure 80: Differences (NB) between TS 103 224 and ES 202 396-1 methodology 
(binaural noises from TS 103 224 database)

#### 4.2.4.5	Analyses of the noise spectra reproduced at the reference microphone

##### 4.2.4.5.1	Introduction

The background noise spectra measured at the reference microphone (input signal for the analysis according to TS 103 106 [6]) of all devices under test are plotted into one diagram. In contrast to the corresponding analysis for HHHF (see clause 4.2.3.5), only one noise type according to TS 103 224 (Office/Callcenter) was evaluated.

##### 4.2.4.5.2	Simulation acc. to ES 202 396-1 with recordings from TS 103 224, reference recording at place of DUT-microphone

As already seen for the HHHF experiments (see clause 4.2.3.5.2), quite large differences can be observed in Figure 81 for the spectra reproduced by the ES 202 396-1 simulation method in the different labs.

The differences range from 5 dB up to 15 dB. The largest differences are located in the low frequency range.

Figure 81: 1/3rd octave spectra at reference microphone and Office noise from TS 103 224 
(ES 202 396-1 methodology)

##### 4.2.4.5.3	Simulation & Recordings acc. to TS 103 224, reference recording at place of DUT-microphone

The differences in the spectra reproduced by the TS 103 224 simulation method are significantly lower compared to the ES 202 396-1 simulation method, as shown in Figure 82.

The differences range from 1 dB up to 8 dB. As expected, the sound field reproduction is highly accurate and consistent across labs in the low frequency domain up to about 2 kHz (where most energy of the noise is found). Here the spectral differences are within a range of 1-2 dB. Up to about 8 kHz, the differences are still less than 5 dB and from 10 - 20 kHz, the differences mostly remain below 5 dB.

Figure 82: 1/3rd octave spectra at reference microphone and Office noise from TS 103 224 
(TS 103 224 methodology)

# 5	Ambient Noise Testing in Handset Mode

## 5.1		Introduction

The investigations described in the following sections refer to ambient noise testing of terminals in handset mode.

## 5.2	Analysis of modified ambient noise playback systems

### 5.2.1	Motivation

The noise field simulation system according to ETSI ES 202 396-1 [4] is currently used for handset UE testing in TS 26.132 [3]. Recent investigations (see clause 4.2.3.5.2) indicated drawbacks regarding the reproducibility across different labs, presumably due to manual steps in the equalization process. On the other hand, the system according to ETSI TS 103 224 [5] provides an improved reproduction accuracy by using automated equalization, as e.g. described in clause 4.2.3.5.3. Thus, it seems reasonable to investigate this noise field simulation also for handset testing.

However, one of the issues to resolve is the usage of background noise scenarios for testing. The simulation according TS 103 224 requires suitable eight-channel signals, recorded at the defined positions of reproduction. ES 202 396-1 utilizes two-channel/binaural recordings, which are not available in the aforementioned eight-channel format. For handset UE testing according to TS 26.132, two possible approaches for the usage of ETSI TS 103 224 may be considered:

- Usage of the noise field simulation and eight-channel background noise scenarios from TS 103 224.

- Another solution is to retain the binaural scenarios from ES 202 396-1, while using the noise field simulation TS 103 224.

The latter solution may be realized with the recently introduced extension of TS 103 224 (clause 7). This study provides some initial results of this flexible noise field simulation, which utilizes various microphone-/loudspeaker combinations. To investigate the consistency of the reproduced noise field between measurement rooms, identical measurements were carried out in two measurement rooms with several configurations. Comparisons of the resulting sound fields as well as deviation metrics are presented in the following.

### 5.2.2	Test Setup

#### 5.2.2.1	Configurations

The recently introduced clause 7 of TS 103 224 provides a flexible and more generic method for noise field simulations. An almost arbitrary number of microphones and loudspeakers can be combined here. For the current evaluation, the two HATS microphones (type 3.3 ears, ID equalization) are used as equalization points. For the playback, the loudspeaker setups 4.1 and 4.0 (four loudspeakers with/without subwoofer) according to ES 202 396-1 are considered. In addition, the default 8.0-setup as described in TS 103 224 is evaluated. Table 6 summarizes the configurations used for this investigation.

Table 6: Noise field simulation configurations

| ID | Noise simulation spec. | Loudspeaker setup |
| --- | --- | --- |
| A | ES 202 396-1 | 4.1 |
| B | TS 103 224 | 4.0 |
| C | TS 103 224 | 4.1 |
| D | TS 103 224 | 8.0 |

Figure 83 illustrates the principle of the different two-point equalizations according to [5] more in detail. In general, all investigated measurement rooms are equipped with eight loudspeakers and a sub-woofer. Depending on the selected configuration, either four (labelled as 1-4 in Figure 83) or eight (labelled as 1-8 in Figure 83) channels are used for equalization and playback. For configuration C, the subwoofer is used as a fifth channel.

Figure 83: Possible flexible equalization configurations

For configuration A, the noise field equalization according to ETSI ES 202 396-1 (4.1 setup) is applied according to clause 6.3 of [4]. In a first step, each loudspeaker is individually compensated regarding level and frequency response to the corresponding ear in a range of 120 Hz and 20 kHz. Then the two left-handed (1 and 4 in Figure 83) and two right-handed (2 and 3 in Figure 83) pre-compensated loudspeakers are jointly equalized to the left and right ear. The compensation for level and frequency response is then repeated for the subwoofer in a frequency range between 30 Hz and 120 Hz. The four loudspeaker channels are then delayed by manually determined values. In the two rooms, different delay values were used, but are close to the recommendations as described in [4]. Finally, the overall frequency responses (left output to left ear, right output to right ear) are manually adjusted with IIR filters to comply with the specified tolerance of ±3 dB in the range of 50 Hz and 10 kHz.

For the configuration B, C and D, sweep-based impulse responses are measured from each used loudspeaker to the two ear microphones. Figure 83 depicts two examples: the two impulse responses h3,L and h3,R are measured between loudspeaker #3 (which is used for configurations B, C and D) and left/right ears. Another example in Figure 83 illustrates the impulse responses h5,L and h5,R between loudspeaker #5 and left/right ears, which are only used for configuration D.

After pre-processing of the impulse responses (clause 6.2.3 of [5]), the inversion filters for each output channel are calculated via matrix inversion (see clause 6.2.4 of [5]) between 50 Hz and 20 kHz, including the modifications for the flexible setups as described in clause 7.4 of [5]. For configuration C, the inversion filters of the sub-woofer are determined in the range of 40 Hz and 120 Hz.

In a final filter adjustment step, all loudspeakers for the configurations B, C and D are active and a suitable real noise recording is used as the so-called “reference signal”, as described in clause 7.3 of [5]. For the current investigation, the first 10.0s of Pub Noise according to ETSI ES 202 396-1 were found to be appropriate for this purpose.

NOTE:	Configuration D seems to be heavily overdetermined: In total, 16 impulse responses are collected for an equalization of only two microphones. However, in this and all other configurations, the accuracy and tolerances as specified in clause 7.5 in [5] were met for both rooms.

#### 5.2.2.2	Evaluation of sound field

In order to investigate the reproduction accuracy across measurement rooms, all introduced equalization methods are analysed at microphone positions, which were not part of the equalization procedure. For this purpose, the eight positions of the fixed array for handset-type and headset terminals as shown in Figure 84 are recorded. These represent relevant positions, where the microphones of the test devices are usually located.

Figure 84: Microphone positions according to [5]

### 5.2.3	Noise Types

The four noise types as shown in Table 7 are used for the evaluation of the reproduction accuracy.

NOTE:	The test procedures as described in clauses 7/8/9/10.12.1 of [3] (speech quality in the presence of ambient noise) are specified with eight noises. The four noise types of Table 7 were selected as an arbitrary, but also representative subset.

Table 7: Investigated noise types acc. to ETSI ES 202 396-1

| Filename | Duration | Level |
| --- | --- | --- |
| Pub_Noise_binaural_V2 | 30 s | L: 75,0 dB(A) R: 73,0 dB(A) |
| Outside_Traffic_Road_binaural | 30 s | L: 74,9 dB(A) R: 73,9 dB(A) |
| Train_Station_binaural | 30 s | L: 68,2 dB(A) R: 69,8 dB(A) |
| Fullsize_Car1_130Kmh_binaural | 30 s | L: 69,1 dB(A) R: 68,1 dB(A) |

### 5.2.4	Measurement rooms

The dimensions of the measurement rooms evaluated in this study are described in Table 8.

NOTE:	For the present investigation, two rooms (room 1 and 4) were considered. In order to keep the naming / numbering convention consistent across related contributions and studies, the same names as in previous work are used.

Table 8: Measurement rooms

| Name | Length [m] | Width [m] | Height [m] | Comment |
| --- | --- | --- | --- | --- |
| Room 1 | 2.40 | 3.40 | 2.05 | Semi-anechoic |
| Room 4 | 1.80 | 2.40 | 2.05 | Semi-anechoic, rather small chamber |

### 5.2.5	Results

For each measurement room and each background noise, the transfer function in 1/3rd octave bands between the binaural source file and the eight recorded positions was evaluated. Each spectrum of the eight positions is then referenced to the corresponding spectrum of the binaural reference signal. Channel numbers 1 to 3 of the recordings are referenced to the left ear signal, while 4 to 8 are referenced to the right ear signal. Note that this procedure does not change the relation between the different configurations, as the reference signal is identical in all cases.

For the sake of clarity, two simplifications are made for illustration and discussion of the results:

- Only the two microphone positions 5 and 8 according to Figure 84 are considered in the following. These two positions represent typical locations of primary/bottom (pos. 5) and secondary/top (pos. 8) microphones of a device under test.

- Only the background noise Road is presented in this clause. For information, Annex B provides all result curves from all combinations of noise field configurations and background noises.

Figure 85 and Figure 86 show the referenced spectra for each configuration (A, B, C and D according to Table 6) in the two investigated rooms 1 and 4. Each plot contains separate curves for the positions / array channels 5 and 8. Mainly the consistency between rooms is of interest. It is not expected that curves are close to 0 dB for all frequencies, since the curves represent transfer functions between different points in space. Referencing the spectra are mainly intended to perform an energy normalization for better comparison of the different channels.

| Figure 85: Referenced spectra of microphones 5 and 8 for configuration A and B, Road noise |  |
| --- | --- |

| Figure 86: Referenced spectra of microphones 5 and 8 for configuration C and D, Road noise |  |
| --- | --- |

In order to quantify the observed differences between the rooms, the maximum and average deviation across frequencies is calculated for each configuration, background noise and microphone position. Table 9 shows the deviation metrics, which are averaged across all eight microphone positions. Additionally, Table 10 provides the deviation metrics averaged across all noise types.

Table 9: Deviations between two rooms, per setup and per background noise

| BGN | Setup | Max. Deviation [dB] | Avg. Deviation [dB] |
| --- | --- | --- | --- |
| Car 130kmh | A | 11.2 | 3.2 |
|  | B | 4.6 | 1.6 |
|  | C | 4.4 | 1.5 |
|  | D | 3.8 | 1.2 |
| Pub | A | 7.8 | 2.8 |
|  | B | 5.5 | 1.6 |
|  | C | 5.1 | 1.7 |
|  | D | 3.5 | 1.2 |
| Road | A | 7.5 | 2.9 |
|  | B | 5.7 | 1.6 |
|  | C | 5.0 | 1.5 |
|  | D | 3.8 | 1.2 |
| Train Station | A | 8.0 | 3.3 |
|  | B | 6.5 | 1.6 |
|  | C | 5.4 | 1.5 |
|  | D | 4.2 | 1.2 |

Table 10: Deviations between two rooms, per setup

| Setup | Max. Deviation [dB] | Avg. Deviation [dB] |
| --- | --- | --- |
| A | 8.6 | 3.1 |
| B | 5.5 | 1.6 |
| C | 5.0 | 1.5 |
| D | 3.8 | 1.2 |

### 5.2.6	Discussion

For configuration A (simulation according to ES 202 396-1, 4.1 loudspeaker setup), a similarity between the sound fields is clearly present, but also larger deviations occur across the entire frequency range. Up to 7.5 dB difference between the two rooms can be overserved, the average deviation is 2.9 dB.

For configuration B (simulation according to TS 103 224, 4.0 loudspeaker setup), the two noise fields at the considered positions are more aligned than for configuration A, but still show some moderate deviations up to 5.7 dB. Compared to configuration A, the average deviation is reduced by almost the half (1.6 dB).

For configuration C (simulation according to TS 103 224, 4.1 loudspeaker setup), almost no differences to configuration B can be identified. The maximum deviation decreases slightly from 5.7 dB to 5.0 dB, while the average is almost the same (1.5 dB). The additionally used sub-woofer seems to have not much impact on the noise field simulation, since it mainly affects the low-frequencies.

For configuration D (simulation according to TS 103 224, 8.0 loudspeaker setup), the smallest deviation in maximum (3.8 dB) and average (1.2 dB) can be observed, i.e. the sound fields are better aligned across rooms than in all other configurations. The increased number of loudspeakers seem to improve the reproduction accuracy and provide the most consistent results across different measurement rooms.

Similar results can be observed for residual noise types, as shown in Table 9. For configuration A, the car noise which includes a substantial amount of low-frequency energy, shows the largest maximum and average deviations. These cannot be identified for the flexible configurations B, C and D. Here the deviation metrics are similar to the other noise types.

The averaged deviation metrics according to Table 10 show a similar trend as the per-noise results. The three flexible setups (configurations B, C and D) indicate an improved reproduction accuracy compared to configuration A. Here both deviation metrics are approximately halved (A: 8.6 dB Max. / 3.1 Avg.) compared to the most advanced configuration D (3.8 dB Max. / 1.2 dB Avg.).

In general, both deviation metrics decrease with an increasing number of loudspeakers. While the maximum deviation could be decreased from 5.5 dB for configuration B to 3.8 dB for configuration D, the improvement in average is rather small (from 1.6 dB for B to 1.2 dB for D).

### 5.2.7	Conclusions

This study presented analysis results of several variants of the flexible background noise simulation according to clause 7 of TS 103 224, which are compared to the default setup of ES 202 396-1. The four setups introduced allow the backward compatible usage of the binaural sound sources of ES 202 396-1. The analysis was carried out by recording the noise signals with microphones close to typical terminal microphone locations. These locations were not included in the equalization process, which was exclusively performed at the DRP of the HATS. The reproduction accuracy across different measurement rooms / labs at typical terminal microphone locations was investigated.

For the test conditions used in this study, the default setup according to ES 202 396-1 leads to largest deviations between the two measurement rooms. The reproduction methods using the TS 103 224 equalization method provide more consistent sound fields around the HATS across measurement rooms and deviations between the different test rooms can be reduced. The reproduction accuracy across the investigated rooms can be improved further by increasing the number of loudspeakers, as shown by comparing configuration B and D. However, this observation may need further verification by additional measurements / comparisons.

Still, it must be noted that the noise fields at the locations compared in this study are unknown and possibly incorrect since the recordings were made purely at the DRP of the artificial head. In the presented study, only the similarity across different rooms was considered, not the absolute reproduction accuracy. However, based on the obtained results, the flexible equalization setup may represent a promising alternative to the one according to ES 202 396-1, which is currently used in TS 26.132 for testing UEs in handset mode.

## 5.3		Noise Field Simulations in different Labs

### 5.3.1	Introduction

The study described in the following clauses investigate at the reproducibility of the background noise configurations among various lab by assessing the difference in the noise spectra between measurement at the HATS ear microphones and the reference noise.

### 5.3.2	Noise Field Simulations

#### 5.3.2.1	System according to ES 202 396-1 (binaural noise sources)

The noise field simulation system according to ETSI ES 202 396-1 [4] is currently used for handset UE testing in TS 26.132 [3]. It utilizes a 4.1 loudspeaker setup (four satellites and one sub-woofer) for reproduction and is equalized in order to play back binaurally recorded signals. The equalization procedure is conducted in several automated steps, but also manual adjustments for cross-talk cancellation are necessary.

However, recent investigations (see clause 4.2.3.5.2) indicated drawbacks regarding the reproducibility across different labs, presumably due to manual steps in the equalization process.

The setup is denoted as ES202 in the following sections.

#### 5.3.2.2	System according to TS 103 224 (8-channel sources)

The noise field simulation according to ETSI TS 103 224 [5] provides an improved reproduction accuracy by:

- An eight-channel microphone array (instead of two ear microphones) as shown in Figure 87, developed in particular for testing of handset and handheld hands-free devices,

- Eight loudspeakers (instead of four loudspeakers and a subwoofer),

- An automated procedure for room equalization and cross-talk cancellation (instead of manual adjustment).

Figure 87: Microphone array for handset testing according to ETSI TS 103 224

Due to the different simulation system, the types of noise sources are different as well – instead of binaural noise recordings, eight-channel signals recorded with the same microphone array have to be used. For this purpose, ETSI TS 103 224 provides a sound file database including similar scenarios as in ETSI ES 202 396-1.

The setup is denoted as TS103-HS in the following sections.

#### 5.3.2.3	System according to TS 103 224 (binaural noise sources)

Beside the default handset and hands-free setups, the recently updated version of ETSI TS 103 224 [5] provides an extension for so-called “flexible setups”. Here an arbitrary number of N microphones and M loudspeakers may be used for the equalization procedure, i.e. the automated equalization procedure is generalized to other configurations (with the restriction of N ≥ 2 and M ≥ N). As microphone inputs, the two ears of the HATS were used (N=2). This allows the usage of the same noise types as for ES202.

Figure 83 illustrates the principle of the different two-point equalizations more in detail. In general, all investigated measurement rooms are equipped with eight loudspeakers and a sub-woofer in order to realize the setups required for ES202 and TS103-HS. Depending on the selected configuration, either four (labelled as 1-4 in Figure 83) or eight (labelled as 1-8 in Figure 83) channels are used for equalization and playback. The subwoofer may also be used as a fifth output channel.

Figure 88: Possible flexible equalization configurations

For the present study, the following setups with M ≥ 4 loudspeakers were investigated:

- 4.1 (four loudspeaker, one sub-woofer – as specified in ETSI ES 202 396-1)

- 4.0 (without sub-woofer)

- 8.0 (same loudspeaker setup as specified in default setup of ETSI 103 224)

The flexible methodology requires the usage of a so-called “reference sound”, which should be a typical example of the microphone configuration. It seems trivial to simply use an arbitrary binaural recording for this purpose; however, several pitfalls were discovered during the equalization across several measurement rooms:

- Since the equalization procedure optimizes transfer functions and filter coefficients based on this sound, it is desirable to select a binaural noise which is not included in the later evaluation,

- The reference sound must provide sufficient “binaural content” in order to ensure adequate cross-talk cancellation between left and right ear,

- The reference sound must provide sufficient energy across all frequency bands in order to accurately determine a transfer function between source and recorded channel.

Based on these observations, the binaural noise TrainStation_bin was selected, which is provided in TS 103 224 as the binaural version of the eight-channel recording (see also Table 13).

The setups are denoted as TS103-4.0, TS103-4.1 and TS103-8.0 in the following sections.

### 5.3.3	Noise Types

According to the descriptions of clauses 7.12.1 (NB), 8.12.1 (WB) and 9.12.1 (SWB/FB) of [2], the binaural noise types as shown in Table 7 are used for the evaluation of noise field simulations ES202, TS103-4.0, TS103-4.1 and TS103-8.0. For the eight-channel system TS103-HS, these identical noise scenarios are not available in the background noise database of ETSI TS 103 224.

However, the noise database here provides similar / related scenarios, which may be used instead. Table 12 shows the eight noise types which are proposed and used for the present study. The database also provides binaural versions of the eight-channel recordings as shown in Table 13.

NOTE 1:	Since the noise recordings are in 8-channel-format, only level information for channel 1 (close to left ear) and 7 (close to right ear) are provided. More information on the specific recordings can be found in clause 8 of ETSI TS 103 224.

NOTE 2:	In the noise database of TS 103 224, there is no equivalent noise condition for “mensa”. The noise “sales counter” is used instead.

Table 11: Noise types acc. to ETSI ES 202 396-1 (binaural)

| Filename | Duration | Level (L & R) |
| --- | --- | --- |
| Pub_Noise_binaural_V2 | 30 s | L: 75,0 dB(A) R: 73,0 dB(A) |
| Outside_Traffic_Road_binaural | 30 s | L: 74,9 dB(A) R: 73,9 dB(A) |
| Outside_Traffic_Crossroads_binaural | 20 s | L: 69,1 dB(A) R: 69,6 dB(A) |
| Train_Station_binaural | 30 s | L: 68,2 dB(A) R: 69,8 dB(A) |
| Fullsize_Car1_130Kmh_binaural | 30 s | L: 69,1 dB(A) R: 68,1 dB(A) |
| Cafeteria_Noise_binaural | 30 s | L: 68,4 dB(A) R: 67,3 dB(A) |
| Mensa_binaural | 22 s | L: 63,4 dB(A) R: 61,9 dB(A) |
| Work_Noise_Office_Callcenter_binaural | 30 s | L: 56,6 dB(A) R: 57,8 dB(A) |

Table 12: Noise types acc. to ETSI TS 103 224 (8-channel versions)

| Filename | Duration | Level Channel 1 & 7 |
| --- | --- | --- |
| Pub_handset | 30 s | 1: 77,2 dB(A) 7: 76,0 dB(A) |
| Roadnoise_handset | 30 s | 1: 72,8 dB(A) 7: 73,0 dB(A) |
| Crossroadnoise_handset | 30 s | 1: 70,6 dB(A) 7: 71,2 dB(A) |
| TrainStation_handset | 30 s | 1: 78,9 dB(A) 7: 78,8 dB(A) |
| FullSizeCar_130_handset | 30 s | 1: 68,5 dB(A) 7: 70,8 dB(A) |
| SalesCounter_handset | 30 s | 1: 66,6 dB(A) 7: 66,6 dB(A) |
| Cafeteria_handset | 30 s | 1: 70,0 dB(A) 7: 70,6 dB(A) |
| Callcenter2_handset | 30 s | 1: 60,2 dB(A) 7: 60,2 dB(A) |

Table 13: Noise types acc. to ETSI TS 103 224 (binaural versions of Table 12)

| Filename | Duration | Level (L & R) |
| --- | --- | --- |
| Pub_bin | 30 s | L: 77,7 dB(A) R: 76,2 dB(A) |
| Roadnoise_bin | 30 s | L: 74,0 dB(A) R: 74,1 dB(A) |
| Crossroadnoise_bin | 30 s | L: 70,8 dB(A) R: 71,6 dB(A) |
| TrainStation_bin | 30 s | L: 78,1 dB(A) R: 78,8 dB(A) |
| FullSizeCar_130_bin | 30 s | L: 68,7 dB(A) R: 70,7 dB(A) |
| SalesCounter_bin | 30 s | L: 66,7 dB(A) R: 66,6 dB(A) |
| Cafeteria_bin | 30 s | L: 69,8 dB(A) R: 70,3 dB(A) |
| Callcenter2_bin | 30 s | L: 60,2 dB(A) R: 60,0 dB(A) |

### 5.3.4	Measurement rooms

The parameters of the measurement rooms evaluated in this study are described in Table 8. All rooms are equipped with sub-woofer setup for the playback according to ES 202 396-1 [4] and with eight satellite loudspeakers in overall for the playback according to TS 103 224 [5].

Table 14: Measurement rooms

| Company | Name | Length [m] | Width [m] | Height [m] | Comment |
| --- | --- | --- | --- | --- | --- |
| HEAD acoustics GmbH | Room#1 | 2.40 | 3.40 | 2.05 | Same chamber type and design as Room 2 |
|  | Room#2 | 2.40 | 3.40 | 2.05 | Same chamber type and design as Room 1 |
|  | Room#3 | 2.90 | 3.10 | 2.05 | Different manufacturer than Room 1,2,4 |
|  | Room#4 | 1.80 | 2.40 | 2.05 | Smallest chamber |
| Intel | Room#5 | 5.57 | 4.67 | 2.93 | Trapezoidal room |
|  | Room#6 | 3.60 | 3.00 | 2.05 |  |

The equalization procedures of the five noise field simulations as described in clause 2 were successfully applied for all rooms. The binaural as well as the eight-channel transfer functions passed the tolerance of +/- 3 dB across the frequency range 50 Hz to 10 kHz, which is specified in ES 202 396-1 as well as in TS 103 224.

Figure 89 to Figure 93 show the validation results of the equalization procedure for all investigated simulation systems and measurement rooms. Note that Figure 89 to Figure 92 show the left and right ear equalization curves, while Figure 93 shows the eight microphone equalization curves of the spatial array.

It can be seen that binaural equalization using TS 103 224-4.0, TS103-4.1 and TS103-8.0 system yield smoother equalization compared to ES202.

NOTE:	Due to time constraints, Room#6 could not perform equalization and measurements with TS103-4.0.

Figure 89: Validation results of equalization (ES202)

|  | Room #6 - N/A |
| --- | --- |

Figure 90: Validation results of equalization (TS103-4.0)

Figure 91: Validation results of equalization (TS103-4.1)

Figure 92: Validation results of equalization (TS103-8.0)

Figure 93: Validation results of equalization (TS103-HS)

### 5.3.5	Test Methodology

#### 5.3.5.1	Spectral Analysis

In order to investigate and quantify the reproduction accuracy of the noise field simulations, for each system, lab and noise type the background noise is recorded with the HATS ears. For each single measurement, the 1/3rd octave spectrum for left and right ear microphone is calculated from the recordings and is then referenced to the corresponding spectrum of the original noise source. If the noise field reproduction was perfect, the difference between the measured spectrum and the original reference spectrum should be zero.

For TS103-HS, the binaural versions of the eight-channel recordings are used as the reference spectrum. Table 8 summarizes the various configurations tested in the current investigation.

Table 15: Configurations of sources and references

| Noise field simulation | Equalization Points | Noise source from | Recording point | Reference noise |
| --- | --- | --- | --- | --- |
| ES202 | Ear Mics (2) | ES 202 396-1 (binaural, Table 7) | Ear Mics | ES 202 396-1 (binaural, Table 7) |
| TS103-4.0 | Ear Mics (2) | ES 202 396-1 (binaural, Table 7) | Ear Mics | ES 202 396-1 (binaural, Table 7) |
| TS103-4.1 | Ear Mics (2) | ES 202 396-1 (binaural, Table 7) | Ear Mics | ES 202 396-1 (binaural, Table 7) |
| TS103-8.0 | Ear Mics (2) | ES 202 396-1 (binaural, Table 7) | Ear Mics | ES 202 396-1 (binaural, Table 7) |
| TS103-HS | Mics of array (8) | TS 103 224 (8 Channels, Table 12) | Ear Mics | TS 103 224 (binaural, Table 13) |

The following definitions apply for the determined spectra and differences:

OX(fi)OX(fi)							Original 1/3rd octave spectrum for left (X=L) or right (X=R) ear at frequency index i.

MX(fi)MX(fi)						Measured 1/3rd octave spectrum for left (X=L) or right (X=R) ear at frequency index i.

HXfi=MX(fi)OX(fi)HXfi=MX(fi)OX(fi)	1/3rd octave spectral distance for left (X=L) or right (X=R) ear at frequency index i.

#### 5.3.5.2	Error Metrics

In order to quantify the accuracy of each system several error metrics are calculated on the difference between the measured and reference spectrum. Based on the definitions above, the metrics are specified in equations (1,2) for left (X=L) or right (X=R) ear (all on dB scale):

Standard deviation:

StdX=1Ni=1 N20⋅log10HXfi-AvgX2	(1)

Absolute maximum deviation:

AbsMaxX=maxi∈1,N20⋅log10HXfiAbsMaxX=maxi∈1,N20⋅log10HXfi	(2)

With:

AvgX=1Ni=1 N20⋅log10HXfiAvgX=1Ni=1 N20⋅log10HXfi		Average of spectral difference in dB

f1f1											Minimum frequency used for metric calculation (50 Hz)

fNfN											Maximum frequency used for metric calculation (10 kHz)

NN											Number of frequencies included in range

Figure 94 depicts two example results of two measured different test cases. The graph displays two spectral differences for the left and right ear. The graph on the left provides a “good” example (regarding metrics), where the transfer function for the noise “Train Station” is almost flat, and results for left and right ear are consistent, whereas the right one shows a “bad” example (regarding metrics), where the transfer function of the noise “Full-size car 130 km/h” shows quite substantial deviations. In addition, there are noticeable differences between left and right ear. Table 16 provides the corresponding metrics for left and right ears.

Figure 94: Example for “good” (left) and “bad” (right) binaural transfer function

Table 16: Deviation metrics for two examples

|  | AbsMax [dB] |  | Std [dB] |  |
| --- | --- | --- | --- | --- |
| Example | L | R | L | R |
| “bad” | 3.60 | 6.62 | 1.28 | 2.30 |
| “good” | 0.63 | 0.63 | 0.29 | 0.27 |

### 5.3.6	Measurement Results

#### 5.3.6.1	Overall Results

Table 17 shows the overall results for all the rooms, noise and both ears combined for the five different background noise (BGN) configurations. For both metrics AbsMax and Std, the maximum and average across all included cases. The lower the number the better.

Table 17: Overall results for the five BGN configuration

| System | Max (AbsMax) [dB] | Avg (AbsMax) [dB] | Max (Std) [dB] | Avg (Std) [dB] |
| --- | --- | --- | --- | --- |
| ES202 | 12.03 | 4.38 | 3.72 | 1.57 |
| TS103-4.0 | 6.31 | 2.25 | 1.43 | 0.66 |
| TS103-4.1 | 6.45 | 2.04 | 1.77 | 0.59 |
| TS103-8.0 | 4.05 | 1.70 | 1.48 | 0.46 |
| TS103-HS | 9.67 | 3.74 | 2.21 | 1.06 |

NOTE 1:	For Room#6, there is no data for TS103-4.0 system available.

NOTE 2:	For Room#5, the noise Train_Station_binaural is missing in TS103-4.1.

NOTE 3:	Due to the different equalization points and different noise types used in TS103-HS, this configuration may not be included in all analyses shown in the following clauses.

The ES202 configuration provides the worst metrics, while TS103-8.0 has the best ones (by a factor of about 3 compared to ES202). TS103-4.0 and TS103-4.1 have quite similar results, but still not as accurate as TS103-8.0.

The results of the TS103-HS seem, at first glance, not as good as the binaural configurations of the TS103. But since the ear microphones were not used as equalization points in TS103-HS, higher deviations than for binaural equalizations are expected. Nevertheless, results are still better than for ES202.

Table 18: Details for AbsMax metric for the five BGN configuration

| System | Max (AbsMax) [dB] | Room | Noise | Frequency [Hz] | Ear |
| --- | --- | --- | --- | --- | --- |
| ES202 | 12.03 | #3 | Pub | 50 | R |
| TS103-4.0 | 6.31 | #2 | FullSizeCar130 | 5000 | R |
| TS103-4.1 | 6.45 | #2 | FullSizeCar130 | 5000 | R |
| TS103-8.0 | 4.05 | #4 | FullSizeCar130 | 5000 | R |
| TS103-HS | 9.67 | #6 | FullSizeCar130 | 10000 | R |

Table 18 provides additional information about the maximum AbsMax (first column of Table 17). This aggregated metric can be tracked to a certain measurement room, noise type, frequency band and ear channel. For ES202, the pub noise causes the maximum value at 50 Hz and right ear. For TS103-4.0/-4.1/8.0, the right ear is affected as well, but here the car noise shows the largest deviations, commonly at 5 kHz. The same observation can be made for TS103-HS, but at 10 kHz.

#### 5.3.6.2	Results per Room

Figure 95 and Figure 96 show maximum AbsMax and average Std, respectively, for all noise types and ears as a function of the BGN system and room.

Figure 95: Maximum (top) and average (bottom) AbsMax deviation for all noises and ear depending on the room and BGN system

Figure 96: Maximum (top) and average (bottom) Std for all noises and ear depending on the room and BGN system

It can be seen that the behaviour between the rooms are consistent, with the ES202 giving the worst results in terms of AbsMax and Std. Both decrease consistently from TS103-4.0 to TS103-8.0, whereas the latter one provides the best results.

Due to the different equalization points, increased error metrics TS103-HS are expected. However, across most rooms, the metrics are better than for ES202.

#### 5.3.6.3	Results per Background Noise

Figure 97 and Figure 98 show maximum AbsMax and average Std, respectively, for all rooms and ear in function of the BGN systems and noise types.

NOTE:	Since TS103-HS is used with different noises than the binaural systems, it is excluded in this analysis.

Figure 97: Maximum (top) and average (bottom) AbsMax for all rooms and ear depending on the noise type and BGN system

Figure 98: Maximum (top) and average (bottom) Std for all rooms and ear depending on the noise type and BGN system

Similar trends as before can be observed: ES202 provides the worst results in all four metrics across all noise types. In particular, the error metrics here vary a lot across noises; the Pub noise always obtains much worse results here.

On the other hand, the method TS103-8.0 again obtains the best results, TS103-4.0 and TS103-4.1 are slightly worse. For these configurations a variance across noises is visible as well, but less pronounced.

#### 5.3.6.4	Impact of Bandwidth

In this chapter the effect of the bandwidth (BW) is examined, which is used to calculate the error metrics. The limit fN = 10 kHz was chosen according to the validation requirement described in [5]. However, most of the energy in the considered noise signals are located in lower frequencies. In order to investigate the dependency of the error metrics on the frequency, all of them are recalculated with fN = 4 kHz.

Figure 99 and Figure 100 show AbsMax and average Std, respectively, for all rooms and ears as a function of the BGN system and BW limit.

Figure 99: Effect of bandwidth on Maximum (top) and average (bottom) AbsMax for all rooms, noise and ear depending on the BGN system

Figure 100: Effect of bandwidth on Maximum (top) and average (bottom) Std for all rooms, noise and ear depending on the BGN system

It can be noticed that the largest differences between the limits 10 kHz and 4 kHz occur for TS103-HS. Results are much closer to the binaural equalizations TS103-4.0/4.1/8.0. This indicates that most spectral differences of TS103-HS occur in the frequency range 4-10 kHz, respectively that the system reproduces the noise field quite adequate in the low frequency domain. This could be explained by the difference in the calibration point (ears vs. microphone array), which will be more noticeable in high frequency.

For the systems TS103-4.0/4.1/8.0, error metrics also decrease from 10 kHz to 4 kHz. Since these configurations are based on the same equalization methodology as TS103-HS, also these results seem plausible.

On the other hand, error metrics for ES202 do not improve between the two different limits. Only for Avg(AbsMax), there is a very slight improvement. For Std metrics, results even get slightly worse. This indicates that most of the spectral differences occur in the low frequency domain between 50 Hz and 4 kHz.

### 5.3.7	Discussion

The present study investigated six labs/measurement rooms regarding the reproduction accuracy of noise field simulations. Four binaural (ES202, TS103-4.0, TS103-4.1 and TS103-8.0) and one microphone-array-based (TS103-HS) equalization methods were used. Eight noises each were evaluated and analysed based on spectral difference metrics between recording and original sound.

In overall, the configurations TS103-4.0, TS103-4.1 and TS103-8.0 provided consistently better results than ES202 with respect to the error metrics defined in 5.3.5.2. A subsequent analysis across labs and noise types confirmed this trend.

In several cases, ES202 was even outperformed by TS103-HS with respect to the error metrics defined in 5.3.5.2, in particular for the lower frequency domain below 4 kHz - even though this method was not equalized at the ear microphones. Since the asymmetric microphone array used for TS103-HS provides more equalization points on the right side of HATS (used for testing handset devices), error metrics as defined in 5.3.5.2 may also be even better when leaving out the left ear.

However, the results of this investigation may only give a note/hint about the variance of different rooms and noise field simulations on acoustic performance evaluation of handset terminals according to TS 26.131/132. Only a subsequent measurement series with multiple devices may confirm the impact on handset terminals.

# 6	Conclusions

## 6.1	Conclusions on handheld and desktop hands-free mode

The following conclusions can be drawn from the Round Robin Experiment:

	In general, the inter-lab consistency of the results is good for both noise reproduction systems.

	When comparing the background noises from ES 202 396-1 [4] and TS 103 224 [5], it can be observed that the corresponding noises from both data sets lead to similar measurement results – apart from two cases where the nature of the background noises and in particular the signal levels differ: Cafeteria noise and train station noise. In those cases, the absolute measurement values changed but their rank order did not.

	When using the binaural background noises from TS 103 224 [5] for the testing, the observed differences between the reproduction methods according to ES 202 396-1 and TS 103 224 are rather small. The offset is not constant, it depends on the type of terminal. The performance scores for all terminals are slightly higher when using ES 202 396-1 in hand-held hands-free mode. For desktop hands-free, the variations observed are slightly larger, but not systematic. The inter-lab correlation obtained by TS 103 224 is higher than with ES 202 396-1. For ES 202 396-1, RMSE between 0.13 and 0.25 MOS are observed, while for TS 103 224 the RMSE decreases to a range of 0.07 to 0.2 MOS.

	The differences between the two systems observed in hand-held hands-free mode when testing according to the current version of TS 26.131 [2] and TS 26.132 [3] are quite small (when using the background noises from TS 103 224). In this case, S-, N- and G-MOS are averaged across all evaluated background noises for a per-device comparison. Differences within a range of 0.1 – 0.2 MOS were found, whereas most of the deviation seem to depend on the terminal. The performance for all terminals is slightly improved when using the simulation method according to ES 202 396-1 compared to the one according to TS 103 224 [3].

	When testing in Desktop hands-free mode, the differences according to TS 26.131 and TS 26.132 (averaging S-, N- and G-MOS across all background noises per device) slightly increase compared to hand-held hands-free mode. However, it should be noted that only two labs with three different rooms were participating in this experiment with a rather limited amount of data.

	The accuracy and consistency of the background noise reproduction at the terminal is improved especially in the low frequency domain below 2 kHz when using TS 103 224. Spectral differences measured at the reference microphone drop from 5-15 dB per 1/3rd octave for the ES 202 396-1 simulation method to 1-5 dB when using the TS 103 224 simulation method.

	It was discussed during the meetings and telephone conferences, if it might be possible to reduce this variability of the ES 202 396-1 sound field reproduction further by defining and validating the delays between the loudspeakers in a different way than it is currently described in ES 202 396-1. Whether this is possible or not and if this leads to less variability was unknown at the time of this study. Any requests regarding these two specifications for noise field simulation needs to be addressed by the responsible committee ETSI TC STQ.

## 6.2	Conclusions on handset mode

TBD

Annex A:
Collection of reports of round robin test in HHHF mode

# A.1	Introduction

The results of the round robin test described in clause 4.2 were presented more in detail across several temporary meeting documents. This annex lists these reports as well as the corresponding test plan and work item description. The documents are provided as an electronic attachment.

# A.2	List of documents

SP-140740: "New work item on Acoustic Test Methods and Performance Objectives for Speakerphone Performance in Noisy Environments (ATeMPO_SPINE)".

S4-AHQ099: "Proposed test plan for a Round Robin Test for comparison of background noise simulations – Rev. 1".

S4-151040: "ATeMPO_SPINE round-robin tests conducted at Sony".

S4-151363: "Report on ATeMPO_SPINE round-robin tests conducted at Orange".

S4-151354: "ATeMPO-SPINE round-robin tests conducted at Knowles".

S4-151343: "Report on ATeMPO_SPINE round-robin tests conducted at HEAD acoustics".

S4-151365: "Results of the Round Robin Test on Different Background Noise Simulation Techniques for Hand-Held Hands-Free Terminals".

Annex B:
Detailed results for modified ambient noise playback

# B.1	Introduction

As described in clause 5.2.5, all result curves from all combinations of noise field configurations and background noises are provided in this Annex.

# B.2	Configuration A

Figure B.1 shows the referenced spectra according to clause 5.2.5 for configuration A (ES 202 396-1 with 4.1 setup).

| Figure B.1: Referenced spectra for configuration A, all noises |  |
| --- | --- |

# B.3	Configuration B

Figure B.2 shows the referenced spectra according to clause 5.2.5 for configuration B (TS 103 224 with 4.0 setup).

| Figure B.2: Referenced spectra for configuration B, all noises |  |
| --- | --- |

# B.4	Configuration C

Figure B.3 shows the referenced spectra according to clause 5.2.5 for configuration C (TS 103 224 with 4.1 setup).

| Figure B.3: Referenced spectra for configuration C, all noises |  |
| --- | --- |

# B.5	Configuration D

Figure B.4 shows the referenced spectra according to clause 5.2.5 for configuration D (TS 103 224 with 8.0 setup).

| Figure B.4: Referenced spectra for configuration D, all noises |  |
| --- | --- |

Annex <X> (informative):
Change history

| Change history |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Date | Meeting | TDoc | CR | Rev | Cat | Subject/Comment | New version |
| 2019-01 | SA4#102 | S4-190045 |  |  |  | Initial draft / skeleton | 0.0.1 |
| 2019-01 | SA4#102 | S4-190150 |  |  |  | Online edits during meeting | 0.1.0 |
| 2019-04 | SA4#103 | S4-190545 |  |  |  | pCR S4-190446, Additional investigations on noise field simulation | 0.2.0 |
| 2019-07 | SA4#104 | S4-190829 |  |  |  | pCR S4-190823, Noise field simulation across multiple labs added | 0.3.0 |