3GPP TSG RAN WG4 Meeting #107 	R4-2309289

Incheon, KR, May 22nd – 26th, 2023

Source:				Rohde & Schwarz

Title:					TP to TR 38.870 on contents for Annex B

Agenda Item:		8.16.2.3

Document for:	Approval

# Introduction

This contribution provides the Text Proposal [1] endorsed by RAN5 during RAN5#98 on MU tables to be included in Annex B of TR 38.870 [2].

Proposal 1: Approve the text proposal below.

# Conclusion

The following proposal is made in this contribution:

Proposal 1: Approve the text proposal below.

# References

- R5-232589, “Text proposal for TR 38.870 Annex B on MU for BHH”, Rohde & Schwarz, 3GPP TSG RAN5 Meeting #99, May 2023.

- 3GPP TR 38.870, v0.3.0 (2023-04)

# Appendix – Text Proposal to TR 38.870 v0.3.0

< Unchanged Text Deleted >

< Beginning of Changes >

# B.4	MU Assessment for TRP

## B.4.1	MU Assessment for TRP in Anechoic Chamber

The uncertainty contributions related to TRP are listed in Table B.4.1-1. A preliminary example uncertainty budgets are presented in Table B.4.1-2 and Table B.4.1-3.

Table B.4.1-1 Uncertainty contributions in TRP measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure 7.2-1, Figure 7.2-2) |  |  |
| 1 | Mismatch of receiver chain | B.2.1 |
| 2 | Insertion loss of receiver chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Measurement Receiver: uncertainty of the absolute level | B.2.4 |
| 5 | Measurement distance | B.2.7 |
| 6 | Quality of quiet zone | B.2.8 |
| 7 | DUT Tx-power drift | B.2.9 |
| 8 | Uncertainty related to the use of phantoms | B.2.11 |
| 9 | Coarse sampling grid | B.2.12 |
| 10 | Random uncertainty | B.2.13 |
| 11 | Frequency Response | B.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure 7.3-1) |  |  |
| 12 | Uncertainty of network analyzer | B.2.15 |
| 13 | Mismatch of receiver chain | B.2.1 |
| 14 | Insertion loss of receiver chain | B.2.2 |
| 15 | Mismatch in the connection of calibration antenna | B.2.1 |
| 16 | Influence of the calibration antenna feed cable | B.2.3 |
| 17 | Influence of the measurement antenna cable | B.2.3 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | B.2.16 |
| 19 | Measurement distance | B.2.7 |
| 20 | Quality of the Quiet Zone | B.2.8 |

Table B.4.1-2 Preliminary example of uncertainty budget for TRP hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гpower meter <0.05 Гmeasurement antenna <0.16 | 0.07 | U-shaped | 1.41 | 1 | 0.05 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Power Meter | 0.06 | Rectangular | 1.73 | 1 | 0.03 |
| 5 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | Rectangular | 1.73 | 1 | 0.16 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | Actual | 1 | 1 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | Rectangular | 1.73 | 1 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | 0.32 | Rectangular | 1.73 | 1 | 0.18 |
| 9 | Coarse sampling grid | Negligible 15° sampling grid | 0 | Actual | 1 | 1 | 0.00 |
| 10 | Random Uncertainty | Monoblock, clamshell and PDA design used for testing | 0.81 | Rectangular | 1.73 | 1 | 0.47 |
| 11 | Frequency Response | Average path loss corrected | 0 | Rectangular | 1.73 | 1 | 0.00 |
|  | Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | Manufacturer’s uncertainty | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | Rectangular | 1.73 | 1 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| 19 | Measurement distance | Dipole: aligned with phase center | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| Combined standard uncertainty |  |  |  |  |  |  | 0.91 |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | 1.78 |

Table B.4.1-3 Preliminary example of uncertainty budget for TRP Beside Head and Hand (speech mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гpower meter <0.05 Гmeasurement antenna <0.16 | 0.07 | U-shaped | 1.41 | 1 | 0.05 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Power Meter | 0.06 | Rectangular | 1.73 | 1 | 0.03 |
| 5 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | Rectangular | 1.73 | 1 | 0.16 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | Actual | 1 | 1 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | Rectangular | 1.73 | 1 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | 0.32 | Rectangular | 1.73 | 1 | 0.18 |
| 9 | Coarse sampling grid | Negligible 15° sampling grid | 0 | Actual | 1 | 1 | 0.00 |
| 10 | Random Uncertainty | Monoblock, clamshell and PDA design used for testing | 1.04 | Rectangular | 1.73 | 1 | 0.60 |
| 11 | Frequency Response | Average path loss corrected | 0 | Rectangular | 1.73 | 1 | 0.00 |
|  | Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | Manufacturer’s uncertainty | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | Rectangular | 1.73 | 1 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| 19 | Measurement distance | Dipole: aligned with phase center | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| Combined standard uncertainty |  |  |  |  |  |  | 0.98 |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | 1.93 |

# B.5	MU Assessment for TRS

## B.5.1	MU Assessment for TRS in Anechoic Chamber

The uncertainty contributions related to TRS are listed in Table B.5.1-1. A preliminary example uncertainty budgets are presented in Table B.5.1-2 and Table B.5.1-3.

Table B.5.1-1 Uncertainty contributions in TRS measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure 7.2-1, Figure 7.2-2) |  |  |
| 1 | Mismatch of transmitter chain | B.2.1 |
| 2 | Insertion loss of transmitter chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Communication Tester: uncertainty of the absolute output level | B.2.5 |
| 5 | Sensitivity measurement: output level step resolution | B.2.6 |
| 6 | Measurement distance | B.2.7 |
| 7 | Quality of quiet zone | B.2.8 |
| 8 | DUT sensitivity drift | B.2.10 |
| 9 | Uncertainty related to the use of phantoms | B.2.11 |
| 10 | Coarse sampling grid | B.2.12 |
| 11 | Random uncertainty | B.2.13 |
| 12 | Frequency Response | B.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure 7.3-1) |  |  |
| 13 | Uncertainty of network analyzer | B.2.15 |
| 14 | Mismatch of transmitter chain | B.2.1 |
| 15 | Insertion loss of transmitter chain | B.2.2 |
| 16 | Mismatch in the connection of calibration antenna | B.2.1 |
| 17 | Influence of the calibration antenna feed cable | B.2.3 |
| 18 | Influence of the measurement antenna cable | B.2.3 |
| 19 | Uncertainty of the absolute gain/radiation efficiency of the calibration antenna | B.2.16 |
| 20 | Measurement distance | B.2.7 |
| 21 | Quality of quiet zone | B.2.8 |

Table B.5.1-2: Preliminary example of uncertainty budget for TRS hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | ГCommTester <0.13 Г antenna connection <0.03 | 0.07 | U-shaped | 1.41 | 1 | 0.05 |  |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 4 | Communication Tester: uncertainty of the absolute output level | Manufacturer’s data sheet | 1 | Rectangular | 1.73 | 1 | 0.58 |  |
| 5 | Sensitivity measurement: output level step resolution | Step of 0.5 dB | 0.25 | Rectangular | 1.73 | 1 | 0.14 |  |
| 6 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | Rectangular | 1.73 | 1 | 0.16 |  |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | Actual | 1 | 1 | 0.5 |  |
| 8 | DUT sensitivity drift | Drift measurement | 0.2 | Rectangular | 1.73 | 1 | 0.12 |  |
| 9 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | 0.32 | Rectangular | 1.73 | 1 | 0.18 |  |
| 10 | Coarse sampling grid | 30° sampling grid | 0.15 | Actual | 1 | 1 | 0.15 |  |
| 11 | Random uncertainty | Monoblock, clamshell and PDA used for testing | 0.91 | Rectangular | 1.73 | 1 | 0.53 |  |
| 12 | Frequency Response | Included in the output level step resolution | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers NA setup | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| 14 | Mismatch of transmitter chain | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |  |
| 15 | Insertion loss of transmitter chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |  |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with dipole | 0.3 | Rectangular | 1.73 | 1 | 0.17 |  |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| 20 | Measurement distance | Dipole: aligned with phase center | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 21 | Quality of quiet zone | Peak-to-null ripple | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| Combined standard uncertainty |  |  |  |  |  |  | 1.12 |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | 2.20 |

Table B.5.1-3: Preliminary example of uncertainty budget for TRS Beside Head and Hand (speech mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | ГCommTester <0.13 Г antenna connection <0.03 | 0.07 | U-shaped | 1.41 | 1 | 0.05 |  |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 4 | Communication Tester: uncertainty of the absolute output level | Manufacturer’s data sheet | 1 | Rectangular | 1.73 | 1 | 0.58 |  |
| 5 | Sensitivity measurement: output level step resolution | Step of 0.5 dB | 0.25 | Rectangular | 1.73 | 1 | 0.14 |  |
| 6 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | Rectangular | 1.73 | 1 | 0.16 |  |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | Actual | 1 | 1 | 0.50 |  |
| 8 | DUT sensitivity drift | Drift measurement | 0.2 | Rectangular | 1.73 | 1 | 0.12 |  |
| 9 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | 0.32 | Rectangular | 1.73 | 1 | 0.18 |  |
| 10 | Coarse sampling grid | 30° sampling grid | 0.15 | Actual | 1 | 1 | 0.15 |  |
| 11 | Random uncertainty | Monoblock, clamshell and PDA used for testing | 0.91 | Rectangular | 1.73 | 1 | 0.81 |  |
| 12 | Frequency Response | Included in the output level step resolution | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers NA setup | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| 14 | Mismatch of transmitter chain | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |  |
| 15 | Insertion loss of transmitter chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |  |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with dipole | 0.3 | Rectangular | 1.73 | 1 | 0.17 |  |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| 20 | Measurement distance | Dipole: aligned with phase center | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 21 | Quality of quiet zone | Peak-to-null ripple | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| Combined standard uncertainty |  |  |  |  |  |  | 1.28 |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | 2.51 |

< End of Changes >