## Summary
| Expanded uncertainty (dB) | [11 empty] | A method exceeds agreed value |  |  |  | Comments
| Indoor Anechoic Chamber |  |  | Compact Antenna Test Range |  |  | Reverberation Chamber |  |  | Agreed value
| 24.25<f ≤29.5 GHz | 37<f ≤40 GHz |  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6<f ≤71 GHz | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6<f ≤71 GHz | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6<f ≤71 GHz | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 4.2<f≤6 GHz
EIRP - Normal conditions |  |  |  | 1.7367 | 2.0697 | 2.978 |  |  |  | 1.7 | 2 | 3 |  | x
EIRP - Extreme conditions |  |  |  | 3.0523 | 3.2533 | 3.8947 |  |  |  | 3.1 | 3.3 | 3.9
Power dynamics |  |  |  | 0.4 | 0.4 | 0.4 |  |  |  | 0.4 | 0.4 | 0.4
EVM (%) |  |  |  | 1 | 1 | 1 |  |  |  | 1 | 1 | 1
TX OFF | [5 empty] | 5.6321 | [5 empty] | 5.6
In-band TRP |  |  |  | 2.1079 | 2.3898 | 3.2107 | 1.8482 | 2.0827 | 2.9282 | 2.1 | 2.4 | 3.2
ACLR- absolute |  |  |  | 2.6922 | 2.7064 | 5.328 | 2.3594 | 2.3594 | 5.1265 | 2.7 | 2.7 | 5.3
ACLR - relative |  |  |  | 2.2821 | 2.5352 | 5.2352 | 2.1485 | 2.3594 | 5.1265 | 2.3 | 2.6 | 5.2
OBUE |  |  |  | 2.7011 | 2.7153 | 5.3325 | 2.3594 | 2.3594 | 5.1265 | 2.7 | 2.7 | 5.3
Co-existance EM
Co-location EM
TX IMD
| Indoor Anechoic Chamber |  |  | Compact Antenna Test Range |  |  | Reverberation Chamber |  |  |  |  | Agreed value
| 18<f≤26,5 GHz | 26,5<f ≤40 GHz | 40<f ≤60 GHz | 18<f≤26,5 GHz | 26,5<f ≤40 GHz | 40<f ≤60 GHz | 18<f≤26,5 GHz | 26,5<f ≤40 GHz | 40<f ≤60 GHz | 71GHz<f ≤110 GHz | 110<f ≤142 GHz | 18<f≤26,5 GHz | 26,5<f ≤40 GHz | 40<f ≤60 GHz | 71GHz<f ≤110 GHz | 110<f ≤142 GHz
OOB EM | x | x | 4.9386 | x | x | 4.957 | x | x | 3.5282 | 5.3427 | 5.9062 | x | x | 5 | 5.3 | 5.9
RX EM | x | x | 4.9386 | x | x | 4.957 | x | x | 3.5282 | 5.3427 | 5.9062 | x | x | 5 | 5.3 | 5.9

## Change history
Date | Meeting | TDoc | CR | Rev | Cat | Subject/Comment | New version
2020-01-01 | [5 empty] | checked 28/1/20: EIRP tab: same values as were in TR 38.817-02. EIRP-Ex tab: same values as were in TR 38.817-02 TX OFF: -pointing error 0.2 in EIRP and 0.3 in TX OFF - use 0.2 for now. -NA is 0.3 in EIRP and 0.2 here, use 0.3 cancels out pointing error -Receiver missmatch uncertainty different to EIRP: 0.72 rather than 0.43/0.57. Could be becuase of limiter? New error description. -Final values matched those as were in TR 38.817-02 In-band TRP: CATR -pointing error different: 0.2 for EIRP, 0.35 here -RX chain missmatch calc is worng (corrected) -Misalignment of positioning system, has calc error (std uncertainty is zero though which matches EIRP) -Misalignment of SGH and pointing system is zero in TRP/TX OFF but is 0.5 here - why? -Switching uncertainty different (uses EIRP uncertainty as std uncertainty) -With all above corrected get 1.74/2.07 compared to 1.78/2.06 - so total is ok Reverb: NA uncertainty shoudl be 0.3 not 0.2, small incraes in total but under agreed value ACLR-abs: CATR -pointing error different: 0.2 for EIRP, 0.3 here - missalignment values different to EIRP (where they are 0) - with EIRP values used the final value is 2.41/2.43 rather than 2.37/2.37 - propose using EIRP values Reverb - NA is wrong (0.2 not 0.3dB) - total very similar as was in TR38.817-02 (2.32) smaller than agreed value so OK ACLR-rel: Checked 28/1/20 CATR -pointing error is different again 0.2 for EIRP, 0.3 here - reciever missmatch different to EIRP -reciever insertion loss variation different (EIRP=0, here, 0.18/0.2) - Using EIRP value, final value is close 1.94/2.23 vs. 1.96/2.31 propose using EIRP values Reverb - NA uncertainty shoudl be 0.3 not 0.2, small incraes in total but under agreed value OBUE: Checked 29/1/20 CATR - NA is wrong -Influence of calibration feed cable is different to ACS abs/EIRP -numbetr of other minor differences - corrected total 2.53 vs 2.37 below agreed value so ok Reverb - NA is wrong (0.2 not 0,.3) total very similar to TR38.817-02 (2.32) smaller than agreed value so OK OOB EM: Checked 29/1/20 Below 40GHz some discrepancies biut suggest using in-band TRP limiits (this is whjat is done in 38.817-02) IAB: power measurement different for 3 methods CATR: power measurement different for 3 methods Reverb: power measurement equipment much larger than other methods - should synchronise RX EM: same as TX emssisons
2020-04-01 | RAN4#95-e |  |  |  |  | 7/4/20 - PWS added | 16.0.0
2020-08-01 | RAN4#96-e | R4-2011260 | 10 |  | A | MU Excel spreadsheets added to the TR 37.941 | 16.1.0
2020-10-01 | RAN4#97-e | R4-2015965 | 18 |  | A | PWS values updates; additional PWS tests added | 16.2.0
2023-05-01 | RAN4#107 | R4-230xxxx |  |  | F | Updated with FR2-2 values. | 17.1.0

## TE
Test Equipment errors | [13 empty] | Comments
UID | Uncertainty source | Uncertainty value (dB) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz |  |  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 0.5 | 0.7 |  | Gaussian | 1 | 0.5 | 0.7 | [5 empty] | RAN4#107: C1-1 is used up to 52.6 GHz. For 52.6-71GHz, the C1-10 is used instead for Power Meter.
C1-2 | RF signal generator | 0.9 | 0.9 |  | Gaussian | 1 | 0.9 | 0.9
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 0.3 | 0.3 | 0.85 |  |  |  |  | RAN4#107: Reuse of 0.85 dB value for FR2-2, based on RAN5 discussions.
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 | 0.51962 | 0.51962 | Rectangular | 1.7321 | 0.3 | 0.3 | 0.3
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 0.9 | 0.9 |  | Gaussian | 1 | 0.9 | 0.9
C1-7_mixer | RF power measurement equipment (e.g. spectrum analyzer, power meter) for IF used with mixer |  |  | 2.36 | Gaussian | 1 |  |  | 2.36
C1-8 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - relative (ACLR) | 0.75 | 0.9 | 2.36 | Gaussian | 1 | 0.75 | 0.9 | 2.36
C1-9 | RF power measurement equipment standard uncertainty σ (dB) of the absolute level for a time domain wideband measurement for FR2 | 1.25 | 1.45 | 2.61 | Gaussian | 1 | 1.25 | 1.45 | 2.61
C1-10 | Uncertainty of the RF power measurement equipment (power meter, power sensor) - high power (EIRP) |  |  | 0.98 | Gaussian | 1 |  |  | 0.98 |  |  |  |  | RAN4#107: used based on the assumption to use Power Meter
Test Equipment errors
UID | Uncertainty source | Uncertainty value (dB) |  |  |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 18<f≤26.5 GHz | 26.5<f ≤40 GHz | 40<f ≤60 GHz | 71<f ≤110 GHz | 110<f ≤142 GHz |  |  | 18<f≤26.5 GHz | 26.5<f ≤40 GHz | 40<f ≤60 GHz | 71<f ≤110 GHz | 110<f ≤142 GHz
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.3 | 0.8 | 1 | Gaussian | 1 | 0.3 | 0.3 | 0.3 | 0.8 | 1
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 | 0.52 | 0.52 | 0.52 | 0.52 | Rectangular | 1.7321 | 0.30022 | 0.30022 | 0.30022 | 0.30022 | 0.30022
C1-7 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 0.52 | 0.84 | 0.6 | 2 | 2.3 | Gaussian | 1 | 0.52 | 0.84 | 0.6 | 2 | 2.3 | RAN4#94: RK: 40-60GHz MU is different for different methods: IAC=0.6 CATR=0.7 Reverb=1.15
Conducted MU (minus missmatch) - up to 6 GHz
UID | Uncertainty source | Uncertainty value (dB) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | f<3 GHz | 3<f<4.2 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz
23 | DL-RS MU derived from conducted spec | 0.41 | 0.56 |  | Gaussian | 1 | 0.41 | 0.56
23 | Total power dynamic range conducted unbcertainty | 0.2 | 0.2 |  | Gaussian | 1 | 0.2 | 0.2
23 | ACLR relative - MU of TE from conducted | 0.41 | 0.56 |  | Gaussian | 1 | 0.41 | 0.56
| ACLR abb q(conducted minus missmatch) | 1.0108 | 1.2321 |  | Gaussian | 1 | 1.0108 | 1.2321
| OBUE (conducted minus missmatch) | 0.75239 | 0.85703 |  | Gaussian | 1 | 0.75239 | 0.85703
27 | Additional (co-location) emissions - condcuted uncertainty (minus missmatch) | 1.02 | 1.28 |  | Gaussian | 1 | 1.02 | 1.28
| TX IMD - conducted measurement uncertainty | 1 | 1.1 |  | Gaussian | 1 | 1 | 1.1
Conducted MU (minus missmatch) - whole spectrum
UID | Uncertainty source | Uncertainty value (dB) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz |  |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz
27 | OOB emsions - Conducted Uncertainty (minus missmatch) | 1.02 | 2.04 |  | Gaussian | 1 | 1.02 | 2.04
27 | Rx emisions - Conducted Uncertainty (minus missmatch) | 1 | 1.99 |  | Gaussian | 1 | 1 | 1.99
Co-location errors
UID | Co-locoation error source | Uncertainty value (dB) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz |  |  |  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz
| Uncertainty related to the selection of the CLTA (Note) | 1.5 | 1.5 |  | Rectangular | 1.7321 | 0.86603 | 0.86603
| Uncertainty related to the placement of the CLTA (Note) | 1.7 | 1.7 |  | Rectangular | 1.7321 | 0.9815 | 0.9815
| Uncertainty related to measuring close to noise floor - Emisions | 0.68 | 0.68 |  | Gaussian | 1 | 0.68 | 0.68
| Uncertainty related to measuring close to noise floor Tx OFF | 1 | 1 |  | Gaussian | 1 | 1 | 1
| Impedance mismatch between feeder cable and CLTA | 0.14 | 0.23 |  | U-Shaped | 1.4142 | 0.098995 | 0.16263
| Gain variations in LNA | 0.1 | 0.1 |  | Gaussian | 1 | 0.1 | 0.1
| Gain variations in measurement amplifier | 0.1 | 0.1 |  | Gaussian | 1 | 0.1 | 0.1

## CATR-Er
CATR | [12 empty] | Comments
UID | Uncertainty source | Uncertainty value (dB) |  |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 40GHz<f ≤60GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 40GHz<f ≤60GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 | 0.2 |  | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0 | 0.1
A2-18a | Misalignment and pointing error of BS (for TRP) | 0.3 | 0.3 | 0.3 | 0.3 | Exp. normal | 2 | 1 | 0.15 | 0.15 | 0.15 | 0.15
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 0.5 | 0.7 |  |  | Gaussian | 1 | 1 | 0.5 | 0.7 | 0
C1-10 | Uncertainty of the RF power measurement equipment (power meter, power sensor) - high power (EIRP) |  |  |  | 0.98 | Gaussian | 1 | 1 |  |  |  | 0.98
A2-2a | Standing wave between BS and test range antenna | 0.03 | 0.03 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.14849 | 0.14849 | RK: uncertainty changed to 0.03 dB
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.001 | 0.001 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.001 | 0.001
A2-4a | QZ ripple experienced by BS | 0.4 | 0.4 | 0.0928 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.0928 | 0.4 | RK: uncertainty value scorrect to match std uncertainty
A2-12 | Frequency flatness of test system | 0.25 | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.25
A2-13 | QZ ripple with BS (extreme test conditions) | 0.7 | 0.7 |  | 0.7 | Gaussian | 1 | 1 | 0.7 | 0.7 | 0 | 0.7
A2-15 | radome loss variation | 0.5 | 0.5 |  | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0 | 0.5
A2-14 | wet radome loss variation | 0.9 | 0.9 |  | 0.9 | Gaussian | 1 | 1 | 0.9 | 0.9 | 0 | 0.9
A2-16 | Change in absorber behaviour | 0.5 | 0.5 |  | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0 | 0.5
C1-9 | RF power measurement equipment standard uncertainty σ (dB) of the absolute level for a time domain wideband measurement for FR2 | 1.25 | 1.45 |  | 2.61 | Gaussian | 1 | 1 | 1.25 | 1.45 | 0 | 2.61
A2-19 | Uncertainty of the LNA (FR2 only) | x | x | 0 |  | Gaussian | 1 | 1 | x | x | 0
A2-20 | Uncertainty of the mixer (FR2 only) | x | x | 2.25 | 2.25 | Gaussian | 1 | 1 | x | x | 2.25 | 2.25
A2-10 | Miscellaneous uncertainty | x | x | 0 | 0 | Rectangular | 1.7321 | 1 | x | x | 0 | 0
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 |  | 0.85
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 0.43 | 0.57 | 0.57 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 |  | 0.40305 | RK: calcualtion corrected - same unceratinty value
A2-5b | Mismatch of receiver chain for low power receiver | 0.72 | 0.72 |  | 0.72 | U-shaped | 1.4142 | 1 | 0.50912 | 0.50912 |  | 0.50912 | Nok: these are non-zero in 37.843. can set the same as in 37.843?
A2-6 | Insertion loss of receiver chain | 0 | 0 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0 | 0 |  | 0.10392 | RK: uncertainty changed to 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.001 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 |  | 0.01
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | 0.29 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 |  | 0.20506
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 | 0.51962 | 0.52 | 0.51962 | Rectangular | 1.7321 | 1 | 0.3 | 0.3 |  | 0.3 | RK: calcualtion corrected - same unceratinty value
A2-8 | Misalignment positioning system | 0 | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  | 0 | RK: uncertainty changed to 0 Nok: these were non-zero in TRE37.843. Can set the same as in TR37.843?
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  | 0 | RK: uncertainty changed to 0 Nok: these were non-zero in TRE37.843. Can set the same as in TR37.843?
A2-9 | Rotary joints | 0 | 0 | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 |  | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 |  | 0.06364
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 |  | 0.009
A2-11 | Switching uncertainty | 0.1 | 0.1 | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 |  | 0.24826 | RK: uncertainty changed to 0.1
Note on C11: Richard Kybett: This is most liely a typo, the std uncertainty is correct just the uncertainty value has FR1 value (0.928) in it.

## Reverb-Er
Reverberation Chamber | [14 empty] | Comments
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 0.5 | 0.7 |  | Gaussian | 1 | 1 | 0.5 | 0.7
C1-10 | Uncertainty of the RF power measurement equipment (power meter, power sensor) - high power (EIRP) |  |  | 0.98 | Gaussian | 1 | 1 |  |  | 0.98
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.14142
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.3 | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.3
A6-4 | Mean value estimation of reference antenna mismatch efficiency | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.85
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5
UID | Uncertainty source | Uncertainty value (dB) |  |  |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 18<f≤26.5 GHz | 26.5<f ≤40 GHz | 40<f ≤60 GHz | 71<f ≤110 GHz | 110<f ≤142 GHz |  |  |  | 18<f≤26.5 GHz | 26.5<f ≤40 GHz | 40<f ≤60 GHz | 71<f ≤110 GHz | 110<f ≤142 GHz
Stage 2: BS measurement
C1-7 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 0.52 | 0.84 | 0.6 | 2 | 2.3 | Gaussian | 1 | 1 | 0.52 | 0.84 | 0.6 | 2 | 2.3 | Values below 40GHz overlap with the table above and are different. Use above values.
A6-1 | Impedance mismatch in the receiving chain | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | U-shaped | 1.4142 | 1 | 0.3182 | 0.3182 | 0.3182 | 0.3182 | 0.3182 | Values below 40GHz overlap with the table above and are different. Use above values.
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.057735 | 0.057735 | Values below 40GHz overlap with the table above and are different. Use above values.
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | Normal | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | Values below 40GHz overlap with the table above and are different. Use above values.
A6-4 | Mean value estimation of reference antenna mismatch efficiency | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | Normal | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | Values below 40GHz overlap with the table above and are different. Use above values.
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.3 | 0.8 | 1 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.3 | 0.8 | 1 | Values below 40GHz overlap with the table above and are different. Use above values.
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | Normal | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | Values below 40GHz overlap with the table above and are different. Use above values.
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | Normal | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | Values below 40GHz overlap with the table above and are different. Use above values.
A6-7 | Uniformity of transfer function | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | Normal | 1 | 1 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | Values below 40GHz overlap with the table above and are different. Use above values.

## IAC-Er
Indoor anechoic Chamber | [10 empty] | Comments
UID | Uncertainty source | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz | Distribution of the probability | Divisor based on distribution shape | ci | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: BS measurement
A1-1 | Positioning misalignment between the BS and the reference antenna | x | x | 0.018 | Rectangular | 1.7321 | 1 | x | x | 0.010392
A1-2 | Pointing misalignment between the BS and the receiving antenna | x | x | 0 | Rectangular | 1.7321 | 1 | x | x | 0
A1-3 | Quality of quiet zone | x | x | 0.1 | Gaussian | 1 | 1 | x | x | 0.1
A1-4a | Polarization mismatch between the BS and the receiving antenna | x | x | 0.018 | Rectangular | 1.7321 | 1 | x | x | 0.010392
A1-5a | Mutual coupling between the BS and the receiving antenna | x | x | 0 | Rectangular | 1.7321 | 1 | x | x | 0
A1-6 | Phase curvature | x | x | 0.05 | Gaussian | 1 | 1 | x | x | 0.05
C1-7 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | x | x | 0.6 | Gaussian | 1 | 1 | x | x | 0.6
A1-16 | Frequency flatness of test system | x | x | 0.25 | Gaussian | 1 | 1 | x | x | 0.25
A1-21 | Uncertainty of the LNA (FR2 only) | x | x | 0 | Gaussian | 1 | 1 | x | x | 0
A1-22 | Uncertainty of the mixer (FR2 only) | x | x | 2.25 | Gaussian | 1 | 1 | x | x | 2.25
A1-7 | Impedance mismatch in the receiving chain | x | x | 0.42 | U-shaped | 1.4142 | 1 | x | x | 0.29698
A1-8 | Random uncertainty | x | x | 0.1 | Rectangular | 1.7321 | 1 | x | x | 0.057735
Stage 1: Calibration measurement
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | x | x | 0.57 | U-shaped | 1.4142 | 1 | x | x | 0.40305
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | x | x | 0.43 | Rectangular | 1.7321 | 1 | x | x | 0.24826
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | x | x | 0.57 | U-shaped | 1.4142 | 1 | x | x | 0.40305
A1-3 | Quality of quiet zone | x | x | 0.1 | Gaussian | 1 | 1 | x | x | 0.1
A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | x | x | 0.018 | Rectangular | 1.7321 | 1 | x | x | 0.010392
A1-5b | Mutual coupling between the reference antenna and the receiving antenna | x | x | 0 | Rectangular | 1.7321 | 1 | x | x | 0
A1-6 | Phase curvature | x | x | 0.07 | Gaussian | 1 | 1 | x | x | 0.07
C1-3 | Uncertainty of the network analyzer | x | x | 0.3 | Gaussian | 1 | 1 | x | x | 0.3
A1-12 | Influence of the reference antenna feed cable | x | x | 0.18 | Rectangular | 1.7321 | 1 | x | x | 0.10392
A1-13 | Reference antenna feed cable loss measurement uncertainty | x | x | 0.1 | Gaussian | 1 | 1 | x | x | 0.1
A1-14 | Influence of the receiving antenna feed cable | x | x | 0.18 | Rectangular | 1.7321 | 1 | x | x | 0.10392
C1-4 | Uncertainty of the absolute gain of the reference antenna | x | x | 0.52 | Rectangular | 1.7321 | 1 | x | x | 0.30022
A1-15 | Uncertainty of the absolute gain of the receiving antenna | x | x | 0 | Rectangular | 1.7321 | 1 | x | x | 0

## EIRP
| Table 9.2.7-2: OTA test system specific measurement uncertainty values for the EIRP accuracy, FR2, Normal test conditions | [12 empty] | Comments
Expanded uncertainty (dB)
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
Indoor Anechoic Chamber
| Compact Antenna Test Range | 1.7367 | 2.0697 | 2.978
One Dimensional Compact Range Chamber
Near Field Test Range
PWS
| Agreed value | 1.7 | 2 | 3 | [9 empty] | Checked 28/1/20: same values as were in TR 38.817-02.
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Table 9.2.3.4-2: CATR MU value derivation for EIRP accuracy measurements, Normal test conditions, FR2 | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 0.5 | 0.7 |  | Gaussian | 1 | 1 | 0.5 | 0.7 |  | 0.25 | 0.49 | 0
C1-10 | Uncertainty of the RF power measurement equipment (power meter, power sensor) - high power (EIRP) |  |  | 0.98 | Gaussian | 1 | 1 |  |  | 0.98 | 0 | 0 | 0.9604
A2-2a | Standing wave between BS and test range antenna | 0.03 | 0.03 | 0.21 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.14849 | 0.00045 | 0.00045 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.001 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.001 | 0.0001 | 0.0001 | 1e-06
A2-4a | QZ ripple experienced by BS | 0.4 | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.4 | 0.16 | 0.16 | 0.16
A2-12 | Frequency flatness of test system | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.85 | 0.09 | 0.09 | 0.7225
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 0.43 | 0.57 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | 0.40305 | 0.09245 | 0.16245 | 0.16245
A2-6 | Insertion loss of receiver chain | 0 | 0 | 0.18 | Rectangular | 1.7321 | 1 | 0 | 0 | 0.10392 | 0 | 0 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01 | 0.0001 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.20506 | 0.02205 | 0.04205 | 0.04205
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 | 0.51962 | 0.51962 | Rectangular | 1.7321 | 1 | 0.3 | 0.3 | 0.3 | 0.09 | 0.09 | 0.09
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.1 | 0.1 | 0.43 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.24826 | 0.0033333 | 0.0033333 | 0.061633
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0.88607 | 1.056 | 1.5194 | 0.88607 | 1.056 | 1.5194
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 1.7367 | 2.0697 | 2.978 | 1.7367 | 2.0697 | 2.978
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0

## EIRP extreme
| Table 9.3.4-2: Test system specific measurement uncertainty values for the EIRP accuracy in Extreme test conditions, FR2 | [12 empty] | Comments
Expanded uncertainty (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz
Indoor anechoic
| CATR | 3.0523 | 3.2533 | 3.8947
Near field
1D CATR
Plane wave sythesiser
| Agreed value | 3.1 | 3.3 | 3.9 | [9 empty] | Checked 28/1/20: same values as were in TR 38.817-02.
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Table 9.3.3.4-1: CATR MU value derivation for EIRP accuracy measurements in Extreme test conditions, FR2 | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 0.5 | 0.7 |  | Gaussian | 1 | 1 | 0.5 | 0.7 |  | 0.25 | 0.49 | 0
C1-10 | Uncertainty of the RF power measurement equipment (power meter, power sensor) - high power (EIRP) |  |  | 0.98 | Gaussian | 1 | 1 | 0 | 0 | 0.98 | 0 | 0 | 0.9604
A2-2a | Standing wave between BS and test range antenna | 0.03 | 0.03 | 0.21 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.14849 | 0.00045 | 0.00045 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.001 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.001 | 0.0001 | 0.0001 | 1e-06
A2-13 | QZ ripple with BS (extreme test conditions) | 0.7 | 0.7 | 0.7 | Gaussian | 1 | 1 | 0.7 | 0.7 | 0.7 | 0.49 | 0.49 | 0.49
A2-12 | Frequency flatness of test system | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
A2-15 | radome loss variation | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25
A2-14 | wet radome loss variation | 0.9 | 0.9 | 0.9 | Gaussian | 1 | 1 | 0.9 | 0.9 | 0.9 | 0.81 | 0.81 | 0.81
A2-16 | Change in absorber behaviour | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.85 | 0.09 | 0.09 | 0.7225
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 0.43 | 0.57 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | 0.40305 | 0.09245 | 0.16245 | 0.16245
A2-6 | Insertion loss of receiver chain | 0 | 0 | 0.18 | Rectangular | 1.7321 | 1 | 0 | 0 | 0.10392 | 0 | 0 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01 | 0.0001 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.20506 | 0.02205 | 0.04205 | 0.04205
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 | 0.51962 | 0.51962 | Rectangular | 1.7321 | 1 | 0.3 | 0.3 | 0.3 | 0.09 | 0.09 | 0.09
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.1 | 0.1 | 0.43 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.24826 | 0.0033333 | 0.0033333 | 0.061633
Combined standard uncertainty (1σ) (dB) | [7 empty] | 1.5573 | 1.6599 | 1.9871 | 1.5573 | 1.6599 | 1.9871
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 3.0523 | 3.2533 | 3.8947 | 3.0523 | 3.2533 | 3.8947
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0

## power dynamics
|  | Expanded uncertainty (dB) | [9 empty] | Comments
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
Indoor anechoic
CATR
Near field
1D CATR
Plane wave sythesiser
| Agreed value | 0.4 | 0.4 | 0.4 | [7 empty] | Same MU values as for FR1.
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!

## EVM
|  | Expanded uncertainty (dB) | [9 empty] | Comments
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
Indoor anechoic
CATR
Near field
1D CATR
Plane wave sythesiser
| Agreed value |  |  | 1 | 1 | 1 | [5 empty] | Values for test equipment only (1% EVM, 12Hz Freq error and 25ns TAE).
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!

## TX OFF
| Table 9.10.3-1: Maximum accepted test system uncertainty values for the EIRP accuracy, FR2 | [12 empty] | Comments
Expanded uncertainty (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz
| Indoor anechoic | [12 empty] | Checked 28/1/20 pointing error 0.2 in EIRP and 0.3 in TX OFF - use 0.2 for now. NA is 0.3 in EIRP and 0.2 here, use 0.3 cancels out pointing error Receiver missmatch uncertainty different to EIRP: 0.72 rather than 0.43/0.57. Could be becuase of limiter? New error description. Final values matched those as were in TR 38.817-02
| CATR | 2.9494 | 3.2939 | 5.6321
Near field
1D CATR
Plane wave sythesiser
| Agreed value | 2.9 | 3.3 | 5.6
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Table 9.10.2.3-1: CATR uncertainty assessment for EIRP measurements for transmitter OFF power and transmitter transient period | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
C1-9 | RF power measurement equipment standard uncertainty σ (dB) of the absolute level for a time domain wideband measurement for FR2 | 1.25 | 1.45 | 2.61 | Gaussian | 1 | 1 | 1.25 | 1.45 | 2.61 | 1.5625 | 2.1025 | 6.8121
A2-2a | Standing wave between BS and test range antenna | 0.03 | 0.03 | 0.21 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.14849 | 0.00045 | 0.00045 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.001 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.001 | 0.0001 | 0.0001 | 1e-06
A2-4a | QZ ripple experienced by BS | 0.4 | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.4 | 0.16 | 0.16 | 0.16
A2-12 | Frequency flatness of test system | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.85 | 0.09 | 0.09 | 0.7225
A2-5b | Mismatch of receiver chain for low power receiver | 0.72 | 0.72 | 0.72 | U-shaped | 1.4142 | 1 | 0.50912 | 0.50912 | 0.50912 | 0.2592 | 0.2592 | 0.2592
A2-6 | Insertion loss of receiver chain | 0 | 0 | 0.18 | Rectangular | 1.7321 | 1 | 0 | 0 | 0.10392 | 0 | 0 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01 | 0.0001 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.20506 | 0.02205 | 0.04205 | 0.04205
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 | 0.51962 | 0.51962 | Rectangular | 1.7321 | 1 | 0.3 | 0.3 | 0.3 | 0.09 | 0.09 | 0.09
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.1 | 0.1 | 0.43 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.24826 | 0.0033333 | 0.0033333 | 0.061633
Combined standard uncertainty (1σ) (dB) | [7 empty] | 1.5048 | 1.6806 | 2.8735 | 1.5048 | 1.6806 | 2.8735
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 2.9494 | 3.2939 | 5.6321 | 2.9494 | 3.2939 | 5.6321
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0

## In-band TRP
| Table 11.2.7-2: Test system specific MU values for the OTA BS output power test, Normal test conditions, FR2 | [12 empty] | Comments
Expanded uncertainty (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz
| Indoor anechoic | [12 empty] | Checked 28/1/20: CATR -pointing error different: 0.2 for EIRP, 0.35 here -RX chain missmatch calc is worng (corrected) -Misalignment of positioning system, has calc error (std uncertainty is zero though which matches EIRP) -Misalignment of SGH and pointing system is zero in TRP/TX OFF but is 0.5 here - why? -Switching uncertainty different (uses EIRP uncertainty as std uncertainty) -With all above corrected get 1.74/2.07 compared to 1.78/2.06 - so total is ok Reverb: NA uncertainty shoudl be 0.3 not 0.2, small incraes in total but under agreed value
| CATR | 2.1079 | 2.3898 | 3.2107
Near field
1D CATR
Plane wave sythesiser
| Reverberation Chamber | 1.8482 | 2.0827 | 2.9282
| Agreed value | 2.1 | 2.4 | 3.2
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Table 11.2.3.4-1: CATR MU value derivation for OTA BS output power measurement, FR2 | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 0.5 | 0.7 |  | Gaussian | 1 | 1 | 0.5 | 0.7 |  | 0.25 | 0.49 | 0
C1-10 | Uncertainty of the RF power measurement equipment (power meter, power sensor) - high power (EIRP) |  |  | 0.98 | Gaussian | 1 | 1 |  |  | 0.98 | 0 | 0 | 0.9604
A2-2a | Standing wave between BS and test range antenna | 0.03 | 0.03 | 0.21 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.14849 | 0.00045 | 0.00045 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.001 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.001 | 0.0001 | 0.0001 | 1e-06
A2-4a | QZ ripple experienced by BS | 0.4 | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.4 | 0.16 | 0.16 | 0.16
A2-12 | Frequency flatness of test system | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.85 | 0.09 | 0.09 | 0.7225
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 0.43 | 0.57 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | 0.40305 | 0.09245 | 0.16245 | 0.16245
A2-6 | Insertion loss of receiver chain | 0 | 0 | 0.18 | Rectangular | 1.7321 | 1 | 0 | 0 | 0.10392 | 0 | 0 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01 | 0.0001 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.20506 | 0.02205 | 0.04205 | 0.04205
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 | 0.51962 | 0.51962 | Rectangular | 1.7321 | 1 | 0.3 | 0.3 | 0.3 | 0.09 | 0.09 | 0.09
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty |  |  | 0.43 | Rectangular | 1.7321 | 1 |  |  | 0.24826 | 0 | 0 | 0.061633 | RAN4#107: switching uncertainty added, motivated by the PM added to the test setup. Resulting total MU unchanged.
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0.88418 | 1.0544 | 1.5194 | 0.88418 | 1.0544 | 1.5194
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 1.733 | 2.0666 | 2.978 | 1.733 | 2.0666 | 2.978
TRP Summation error | [7 empty] | 1.2 | 1.2 | 1.2
Total MU | [7 empty] | 2.1079 | 2.3898 | 3.2107
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Table 11.2.5.4-1: Reverberation chamber MU value derivation for OTA BS output power, FR2 | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 0.5 | 0.7 |  | Gaussian | 1 | 1 | 0.5 | 0.7 |  | 0.25 | 0.49 | 0
C1-10 | Uncertainty of the RF power measurement equipment (power meter, power sensor) - high power (EIRP) |  |  | 0.98 | Gaussian | 1 | 1 |  |  | 0.98 | 0 | 0 | 0.9604
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.14142 | 0.02 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.3 | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.3 | 0.09 | 0.09 | 0.09
A6-4 | Mean value estimation of reference antenna mismatch efficiency | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.85 | 0.09 | 0.09 | 0.7225
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0.94294 | 1.0626 | 1.494 | 0.94294 | 1.0626 | 1.494
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 1.8482 | 2.0827 | 2.9282 | 1.8482 | 2.0827 | 2.9282

## ACLR-abs
| Table 11.3.7-3: Test system specific MU values for the absolute OTA ACLR, Normal test conditions, FR2 | [12 empty] | Comments
Expanded uncertainty (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz
| Indoor anechoic | [12 empty] | Checked 28/1/20 CATR -pointing error different: 0.2 for EIRP, 0.3 here - missalignment values different to EIRP (where they are 0) - with EIRP values used the final value is 2.41/2.43 rather than 2.37/2.37 - propose using EIRP values Reverb - NA is wrong (0.2 not 0.3dB) - total very similar as was in TR38.817-02 (2.32) smaller than agreed value so OK
| CATR | 2.6922 | 2.7064 | 5.328
Near field
1D CATR
Plane wave sythesiser
| Reverberation Chamber | 2.3594 | 2.3594 | 5.1265
| Agreed value | 2.7 | 2.7 | 5.3
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Table 11.3.3.4-1: CATR MU value derivation for the EIRP measurement of the absolute OTA ACLR, FR2 | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 0.9 | 0.9 | 2.36 | Gaussian | 1 | 1 | 0.9 | 0.9 | 2.36 | 0.81 | 0.81 | 5.5696
A2-2a | Standing wave between BS and test range antenna | 0.03 | 0.03 | 0.21 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.14849 | 0.00045 | 0.00045 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.001 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.001 | 0.0001 | 0.0001 | 1e-06
A2-4a | QZ ripple experienced by BS | 0.4 | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.4 | 0.16 | 0.16 | 0.16
A2-12 | Frequency flatness of test system | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.85 | 0.09 | 0.09 | 0.7225
A2-5b | Mismatch of receiver chain for low power receiver | 0.72 | 0.72 | 0.72 | U-shaped | 1.4142 | 1 | 0.50912 | 0.50912 | 0.50912 | 0.2592 | 0.2592 | 0.2592
A2-6 | Insertion loss of receiver chain | 0 | 0 | 0.18 | Rectangular | 1.7321 | 1 | 0 | 0 | 0.10392 | 0 | 0 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01 | 0.0001 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.20506 | 0.02205 | 0.04205 | 0.04205
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 | 0.51962 | 0.51962 | Rectangular | 1.7321 | 1 | 0.3 | 0.3 | 0.3 | 0.09 | 0.09 | 0.09
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.1 | 0.1 | 0.43 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.24826 | 0.0033333 | 0.0033333 | 0.061633
Combined standard uncertainty (1σ) (dB) | [7 empty] | 1.2296 | 1.2377 | 2.6485 | 1.2296 | 1.2377 | 2.6485
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 2.41 | 2.4259 | 5.1911 | 2.41 | 2.4259 | 5.1911
TRP Summation error | [7 empty] | 1.2 | 1.2 | 1.2
Total MU | [7 empty] | 2.6922 | 2.7064 | 5.328
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Table 11.3.5.4-1: Reverberation chamber MU value derivation for absolute ACLR measurement, FR2 | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 0.9 | 0.9 | 2.36 | Gaussian | 1 | 1 | 0.9 | 0.9 | 2.36 | 0.81 | 0.81 | 5.5696
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.14142 | 0.02 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.3 | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.3 | 0.09 | 0.09 | 0.09
A6-4 | Mean value estimation of reference antenna mismatch efficiency | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.85 | 0.09 | 0.09 | 0.7225
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25
Combined standard uncertainty (1σ) (dB) | [7 empty] | 1.2038 | 1.2038 | 2.6156 | 1.2038 | 1.2038 | 2.6156
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 2.3594 | 2.3594 | 5.1265 | 2.3594 | 2.3594 | 5.1265

## ACLR-rel
| Table 11.3.7-4: Test system specific MU values for relative OTA ACLR, Normal conditions, FR2 | [12 empty] | Comments
Expanded uncertainty (dB)
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
Indoor anechoic
| CATR | 2.2821 | 2.5352 | 5.2352
Near field
1D CATR
Plane wave sythesiser
| Reverberation Chamber | 2.1485 | 2.3594 | 5.1265
| Agreed value | 2.3 | 2.6 | 5.2
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 0 | 0 |  | 0 | 0
TRP Summation error | [7 empty] | 0.75 | 0.75
Total MU | [7 empty] | 0.75 | 0.75
Table 11.3.3.4-2: CATR MU value derivation for the EIRP measurement of the relative OTA ACLR, FR2 | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
C1-8 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - relative (ACLR) | 0.75 | 0.9 | 2.36 | Gaussian | 1 | 1 | 0.75 | 0.9 | 2.36 | 0.5625 | 0.81 | 5.5696
A2-4a | QZ ripple experienced by BS | 0.4 | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.4 | 0.16 | 0.16 | 0.16
A2-12 | Frequency flatness of test system | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.85 | 0.09 | 0.09 | 0.7225
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 0.43 | 0.57 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | 0.40305 | 0.09245 | 0.16245 | 0.16245
A2-6 | Insertion loss of receiver chain | 0 | 0 | 0.18 | Rectangular | 1.7321 | 1 | 0 | 0 | 0.10392 | 0 | 0 | 0.0108
A2-11 | Switching uncertainty | 0.1 | 0.1 | 0.43 | Rectangular | 1.7321 | 0 | 0.057735 | 0.057735 | 0.24826 | 0.0033333 | 0.0033333 | 0.061633
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0.99035 | 1.1394 | 2.5999 | 0.99035 | 1.1394 | 2.5999
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 1.9411 | 2.2333 | 5.0958 | 1.9411 | 2.2333 | 5.0958
TRP Summation error | [7 empty] | 1.2 | 1.2 | 1.2
Total MU | [7 empty] | 2.2821 | 2.5352 | 5.2352
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Table 11.3.5.4-2: Reverberation chamber MU value derivation for relative ACLR measurement, FR2 | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
C1-8 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - relative (ACLR) | 0.75 | 0.9 | 2.36 | Gaussian | 1 | 1 | 0.75 | 0.9 | 2.36 | 0.5625 | 0.81 | 5.5696
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.14142 | 0.02 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.3 | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.3 | 0.09 | 0.09 | 0.09
A6-4 | Mean value estimation of reference antenna mismatch efficiency | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.85 | 0.09 | 0.09 | 0.7225
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25
Combined standard uncertainty (1σ) (dB) | [7 empty] | 1.0962 | 1.2038 | 2.6156 | 1.0962 | 1.2038 | 2.6156
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 2.1485 | 2.3594 | 5.1265 | 2.1485 | 2.3594 | 5.1265

## OBUE
| Table 11.4.7-2: Test system specific MU values for the OTA OBUE measurement, FR2 | [12 empty] | Comments
Expanded uncertainty (dB)
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
| Indoor anechoic | [12 empty] | Checked 29/1/20 CATR - NA is wrong -Influence of calibration feed cable is different to ACS abs/EIRP -numbetr of other minor differences - corrected total 2.53 vs 2.37 below agreed value so ok Reverb - NA is wrong (0.2 not 0,.3) total very similar to TR38.817-02 (2.32) smaller than agreed value so OK
| CATR | 2.7011 | 2.7153 | 5.3325
Near field
1D CATR
Plane wave sythesiser
| Reverberation Chamber | 2.3594 | 2.3594 | 5.1265
| Agreed value | 2.7 | 2.7 | 5.3
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Table 11.4.3.4-1: CATR MU value derivation for OTA OBUE measurement, FR2 | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
A2-18a | Misalignment and pointing error of BS (for TRP) | 0.3 | 0.3 | 0.3 | Exp. normal | 2 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 0.9 | 0.9 | 2.36 | Gaussian | 1 | 1 | 0.9 | 0.9 | 2.36 | 0.81 | 0.81 | 5.5696
A2-2a | Standing wave between BS and test range antenna | 0.03 | 0.03 | 0.21 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.14849 | 0.00045 | 0.00045 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.001 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.001 | 0.0001 | 0.0001 | 1e-06
A2-4a | QZ ripple experienced by BS | 0.4 | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.4 | 0.16 | 0.16 | 0.16
A2-12 | Frequency flatness of test system | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.85 | 0.09 | 0.09 | 0.7225
A2-5b | Mismatch of receiver chain for low power receiver | 0.72 | 0.72 | 0.72 | U-shaped | 1.4142 | 1 | 0.50912 | 0.50912 | 0.50912 | 0.2592 | 0.2592 | 0.2592
A2-6 | Insertion loss of receiver chain | 0 | 0 | 0.18 | Rectangular | 1.7321 | 1 | 0 | 0 | 0.10392 | 0 | 0 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01 | 0.0001 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.20506 | 0.02205 | 0.04205 | 0.04205
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 | 0.51962 | 0.51962 | Rectangular | 1.7321 | 1 | 0.3 | 0.3 | 0.3 | 0.09 | 0.09 | 0.09
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.1 | 0.1 | 0.43 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.24826 | 0.0033333 | 0.0033333 | 0.061633
Combined standard uncertainty (1σ) (dB) | [7 empty] | 1.2347 | 1.2427 | 2.6509 | 1.2347 | 1.2427 | 2.6509
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 2.4199 | 2.4357 | 5.1957 | 2.4199 | 2.4357 | 5.1957
TRP summation error | [7 empty] | 1.2 | 1.2 | 1.2
Total MU | [7 empty] | 2.7011 | 2.7153 | 5.3325
Table 11.4.5.4-1: Reverberation chamber MU value derivation for OTA OBUE measurement, FR2 | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz |  |  |  | 24.25<f ≤29.5 GHz | 37<f ≤40 GHz | 52.6 < f ≤ 71 GHz
Stage 2: BS measurement
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 0.9 | 0.9 | 2.36 | Gaussian | 1 | 1 | 0.9 | 0.9 | 2.36 | 0.81 | 0.81 | 5.5696
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.14142 | 0.02 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.3 | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.3 | 0.09 | 0.09 | 0.09
A6-4 | Mean value estimation of reference antenna mismatch efficiency | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | 0.85 | Gaussian | 1 | 0.3 | 0.3 | 0.3 | 0.85 | 0.09 | 0.09 | 0.7225
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25
Combined standard uncertainty (1σ) (dB) | [7 empty] | 1.2038 | 1.2038 | 2.6156 | 1.2038 | 1.2038 | 2.6156
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 2.3594 | 2.3594 | 5.1265 | 2.3594 | 2.3594 | 5.1265
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Combined standard uncertainty (1σ) (dB) | [7 empty] | 0 | 0
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0

## OOB EM
| Table 12.2.5-2: Spurious emissions MU values – FR2 | [12 empty] | Comments
Expanded uncertainty (dB)
|  | 30MHz < f ≤ 6GHz | 6<f≤40 GHz | 40<f ≤60 GHz | 71 < f ≤ 110 GHz | 110 < f ≤ 142 GHz
| Indoor Anechoic Chamber | x | x | 4.9386 | [9 empty] | Checked 29/1/20 Below 40GHz some discrepancies biut suggest using in-band TRP limiits (this is whjat is done in 38.817-02) IAB: power measurement different for 3 methods CATR: power measurement different for 3 methods Reverb: power measurement equipment much larger than other methods - should synchronise
| CATR | x | x | 4.957
| Reverberation Chamber | x | x | 3.5282 | 5.3427 | 5.9062
| Agreed value | 2.5 | 2.7 | 5 | 5.3 | 5.9 | [7 empty] | There was an error in TR 38.817-02: value for 40-60 GHz was recorded as 4.9 dB
Table 12.2.4.3-2: Reverberation Chamber value derivation for TX spurious emissions, 43 GHz – 142 GHz | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 40GHz<f ≤60GHz | 71 < f ≤ 110 GHz | 110 < f ≤ 142 GHz |  |  |  | 40<f ≤60 GHz | 71 < f ≤ 110 GHz | 110 < f ≤ 142 GHz
Stage 2: BS measurement
C1-7 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 0.6 | 2 | 2.3 | Gaussian | 1 | 1 | 0.6 | 2 | 2.3 | 0.36 | 4 | 5.29
A6-1 | Impedance mismatch in the receiving chain | 0.45 | 0.45 | 0.45 | U-shaped | 1.4142 | 1 | 0.3182 | 0.3182 | 0.3182 | 0.10125 | 0.10125 | 0.10125
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.5 | 0.5 | 0.5 | Normal | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25
A6-4 | Mean value estimation of reference antenna mismatch efficiency | 0.27 | 0.27 | 0.27 | Normal | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.8 | 1 | Gaussian | 1 | 1 | 0.3 | 0.8 | 1 | 0.09 | 0.64 | 1
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Normal | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Normal | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 1.5 | 1.5 | 1.5 | Normal | 1 | 1 | 1.5 | 1.5 | 1.5 | 2.25 | 2.25 | 2.25
Combined standard uncertainty (1σ) (dB) | [7 empty] | 1.8001 | 2.7259 | 3.0134 | 1.8001 | 2.7259 | 3.0134
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [7 empty] | 3.5282 | 5.3427 | 5.9062 | 3.5282 | 5.3427 | 5.9062
Table 12.2.2.4-1: IAC value derivation for TX spurious emissions, FR2 | [10 empty] | rms calculations
UID | Uncertainty source |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 40GHz<f ≤60GHz |  |  |  | 40<f ≤60 GHz
Stage 2: BS measurement
A1-1 | Positioning misalignment between the BS and the reference antenna | 0.018 | Rectangular | 1.7321 | 1 | 0.010392 |  |  |  |  | 0.000108
A1-2 | Pointing misalignment between the BS and the receiving antenna | 0 | Rectangular | 1.7321 | 1 | 0 |  |  |  |  | 0
A1-3 | Quality of quiet zone | 0.1 | Gaussian | 1 | 1 | 0.1 |  |  |  |  | 0.01
A1-4a | Polarization mismatch between the BS and the receiving antenna | 0.018 | Rectangular | 1.7321 | 1 | 0.010392 |  |  |  |  | 0.000108
A1-5a | Mutual coupling between the BS and the receiving antenna | 0 | Rectangular | 1.7321 | 1 | 0 |  |  |  |  | 0
A1-6 | Phase curvature | 0.05 | Gaussian | 1 | 1 | 0.05 |  |  |  |  | 0.0025
C1-7 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 0.6 | Gaussian | 1 | 1 | 0.6 |  |  |  |  | 0.36
A1-16 | Frequency flatness of test system | 0.25 | Gaussian | 1 | 1 | 0.25 |  |  |  |  | 0.0625
A1-21 | Uncertainty of the LNA (FR2 only) | 0 | Gaussian | 1 | 1 | 0 |  |  |  |  | 0
A1-22 | Uncertainty of the mixer (FR2 only) | 2.25 | Gaussian | 1 | 1 | 2.25 |  |  |  |  | 5.0625
A1-7 | Impedance mismatch in the receiving chain | 0.42 | U-shaped | 1.4142 | 1 | 0.29698 |  |  |  |  | 0.0882
A1-8 | Random uncertainty | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 |  |  |  |  | 0.0033333
Stage 1: Calibration measurement | [10 empty] | 0
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.57 | U-shaped | 1.4142 | 1 | 0.40305 |  |  |  |  | 0.16245
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.43 | Rectangular | 1.7321 | 1 | 0.24826 |  |  |  |  | 0.061633
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.57 | U-shaped | 1.4142 | 1 | 0.40305 |  |  |  |  | 0.16245
A1-3 | Quality of quiet zone | 0.1 | Gaussian | 1 | 1 | 0.1 |  |  |  |  | 0.01
A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.018 | Rectangular | 1.7321 | 1 | 0.010392 |  |  |  |  | 0.000108
A1-5b | Mutual coupling between the reference antenna and the receiving antenna | 0 | Rectangular | 1.7321 | 1 | 0 |  |  |  |  | 0
A1-6 | Phase curvature | 0.07 | Gaussian | 1 | 1 | 0.07 |  |  |  |  | 0.0049
C1-3 | Uncertainty of the network analyzer | 0.3 | Gaussian | 1 | 1 | 0.3 |  |  |  |  | 0.09
A1-12 | Influence of the reference antenna feed cable | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 |  |  |  |  | 0.0108
A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.1 | Gaussian | 1 | 1 | 0.1 |  |  |  |  | 0.01
A1-14 | Influence of the receiving antenna feed cable | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 |  |  |  |  | 0.0108
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 |  |  |  |  | 0.090133
A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | Rectangular | 1.7321 | 1 | 0 |  |  |  |  | 0
Combined standard uncertainty (1σ) (dB) | [5 empty] | 2.4905 |  |  |  |  | 2.4905
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [5 empty] | 4.8814 |  |  |  |  | 4.8814
TRP Summation error | [5 empty] | 0.75
Total MU | [5 empty] | 4.9386
Table 12.2.3.3-1: CATR value derivation for TX spurious emissions, FR2 | [10 empty] | rms calculations
UID | Uncertainty source |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 40GHz<f ≤60GHz |  |  |  | 40<f ≤60 GHz
Stage 2: BS measurement
A2-18a | Misalignment and pointing error of BS (for TRP) | 0.3 | Exp. normal | 2 | 1 | 0.15 |  |  |  |  | 0.0225
C1-7 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 0.6 | Gaussian | 1 | 1 | 0.6 |  |  |  |  | 0.36
A2-19 | Uncertainty of the LNA (FR2 only) | 0 | Gaussian | 1 | 1 | 0 |  |  |  |  | 0
A2-20 | Uncertainty of the mixer (FR2 only) | 2.25 | Gaussian | 1 | 1 | 2.25 |  |  |  |  | 5.0625
A2-2a | Standing wave between BS and test range antenna | 0.21 | Gaussian | 1 | 1 | 0.21 |  |  |  |  | 0.0441
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.001 | Gaussian | 1 | 1 | 0.001 |  |  |  |  | 1e-06
A2-4a | QZ ripple experienced by BS | 0.0928 | Gaussian | 1 | 1 | 0.0928 |  |  |  |  | 0.0086118
A2-12 | Frequency flatness of test system | 0.25 | Gaussian | 1 | 1 | 0.25 |  |  |  |  | 0.0625
A2-10 | Miscellaneous uncertainty | 0 | Rectangular | 1.7321 | 1 | 0 |  |  |  |  | 0
Stage 1: Calibration measurement | [10 empty] | 0
C1-3 | Uncertainty of the network analyzer | 0.3 | Gaussian | 1 | 1 | 0.3 |  |  |  |  | 0.09
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 0.57 | U-shaped | 1.4142 | 1 | 0.40305 |  |  |  |  | 0.16245
A2-5b | Mismatch of receiver chain for low power receiver | 0 | U-shaped | 1.4142 | 1 | 0 |  |  |  |  | 0
A2-6 | Insertion loss of receiver chain | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 |  |  |  |  | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.001 | Gaussian | 1 | 1 | 0.001 |  |  |  |  | 1e-06
A2-7 | Influence of the calibration antenna feed cable | 0.29 | U-shaped | 1.4142 | 1 | 0.20506 |  |  |  |  | 0.04205
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 |  |  |  |  | 0.090133
A2-8 | Misalignment positioning system | 0 | Exp. normal | 2 | 1 | 0 |  |  |  |  | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 | Exp. normal | 2 | 1 | 0 |  |  |  |  | 0
A2-9 | Rotary joints | 0 | U-shaped | 1.4142 | 1 | 0 |  |  |  |  | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 |  |  |  |  | 0.00405
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 | Gaussian | 1 | 1 | 0.009 |  |  |  |  | 8.1e-05
A2-11 | Switching uncertainty | 0.43 | Rectangular | 1.7321 | 1 | 0.24826 |  |  |  |  | 0.061633
Combined standard uncertainty (1σ) (dB) | [5 empty] | 2.4539 |  |  |  |  | 2.4539
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [5 empty] | 4.8096 |  |  |  |  | 4.8096
TRP Summation error | [5 empty] | 1.2
Total MU | [5 empty] | 4.957

## RX EM
|  | Expanded uncertainty (dB) | [11 empty] | Comments
|  | 18<f≤26.5 GHz | 26.5<f ≤40 GHz | 40<f ≤60 GHz | 71 < f ≤ 110 GHz | 110 < f ≤ 142 GHz
| Indoor anechoic chamber | x | x | x | x | x
| CATR | x | x | x | x | x
| Reverberation Chamber | x | x | x | 5.3427 | 5.9062
| Agreed value | x | x | 5 | 5.3 | 5.9 | [7 empty] | Checked 29/1/20: same as TX emssisons
Indoor anechoic chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz |  |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz |  |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Reverberation Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz |  |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0