3GPP TSG-RAN WG4 Meeting #110	R4-2400033
Athens, GR, 26 Feb – 01 Mar, 2024

Agenda item:	8.12.3

Source:	CAICT

Title:	Analysis of 3GPP Rel-18 FR1 MIMO OTA lab alignment results

Document for:	Discussion

1	Introduction

Based on the Framework for FR1 MIMO OTA [1], the FR1 MIMO OTA lab alignment should be concluded at this meeting. All 6 volunteer labs have submitted their measurement results of the performance alignment devices (PADs) in the 1st round [2]-[7].

At this meeting, a CR to the test plan TS 38.151 is agreed in [8], which removed the statement that power validation reuslts can be considered a systematic offset to correct the final UE sensitivity vlaues. Accordingly, it was agreed in the WF [9] that the lab alignment outcome should be based on the measurement results without power validation systematic offsets.

Besides, it was agreed at this meeting [9] that the pass/fail limit for Rel-18 lab alignment is updated based on the lated MU agreed in RAN5 (specifiedy in Table B.1-1 of TS 38.551). For bands≤3GHz, the pass/fail limit is updated to 0.75*2.61=1.96 dB.

This contribution presents an analysis of the Rel-18 FR1 MIMO OTA lab alignment result submitted by all the 6 labs in the 1st round, both the original measurement results (with systematic offsets) and the valid measurement results (without systematic offsets) are presented.  According to the Framework [1] and the WF of this meeting [9] and the last meeting [10], the reference value of each PAD is derived by the linear average (in dB) of the measurement results submitted by all 6 labs in the 1st round. The pass/fail limit is specified as +/- 0.75* MU (+/- 1.96 dB). The detailed analysis data is attached in the Excel worksheet of this contribution.

2	Discussion

2.1	Original lab alignment results (with systematic offsets)

The original FR1 MIMO OTA lab alignment results submitted by the 6 labs in the 1st round are summarized and analysed as shown in Table 1.

Table 1. Summary of the original FR1 MIMO OTA lab alignment results (with systematic offsets)

| Device | Band | TRMS measurement result [dBm/15kHz] |  |  |  |  |  | Average approach | Average value | Max-Min deviation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 |  |  |  |
| PAD_1 | n28 | -89.34 | -90.58 | -89.95 | -88.26 | -89.22 | -91.39 | Linear average | -89.47 | 2.32 |
| PAD_2 | n28 | -86.46 | -88.39 | -88.61 | -85.50 | -85.86 | -87.38 |  | -87.03 | 3.12 |
| PAD_3 | n28 | -86.26 | -89.80 | -88.75 | -86.19 | -86.49 | -88.00 |  | -87.50 | 3.61 |
| Device | Band | TRMS offset [dBm/15kHz] |  |  |  |  |  | Pass/fail limit |  |  |
|  |  | Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 |  |  |  |
| PAD_1 | n28 | 0.13 | -1.11 | -0.48 | 1.21 | 0.25 | -1.60 | +/- 0.75 MU, i.e., +/- 1.96 dB |  |  |
| PAD_2 | n28 | 0.57 | -1.35 | -1.58 | 1.54 | 1.17 | -0.35 |  |  |  |
| PAD_3 | n28 | 1.23 | -2.30 | -1.25 | 1.31 | 1.01 | -0.42 |  |  |  |

Figure 1 presents the original TRMS measurement results submitted by all labs in the 1st round, and Figure 2 visualizes the deviations between TRMS measurement results from each lab and reference values.

Figure 1. Original TRMS measurement results submitted by 6 labs (with systematic offsets)

Figure 2. Deviations between original TRMS measurement results from each lab and reference values (with systematic offsets)

2.2	Valid lab alignment results (without systematic offsets)

During this meeting [11], volnteer labs reported the systematic offsets that applied to their submitted resutls as listed in Table 2.

Table 2. Systematic offsets applied to the original lab alignment results

|  | Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 |
| --- | --- | --- | --- | --- | --- | --- |
| Systematic offset (dBm/15kHz) | 0 | -0.46 | 0.48 | 0 | 0 | 0.36 |

After removing the systematic offsets from the original lab alignment results, the valid updated lab alignment results in the 1st round are summarised in Table 3.

Table 3. Summary of the valid FR1 MIMO OTA lab alignment results (without systematic offsets)

| Device | Band | TRMS measurement result [dBm/15kHz] |  |  |  |  |  | Average approach | Average value | Max-Min deviation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 |  |  |  |
| PAD_1 | n28 | -89.34 | -90.12 | -90.43 | -88.26 | -89.22 | -91.75 | Linear average | -89.85 | 3.49 |
| PAD_2 | n28 | -86.46 | -87.93 | -89.09 | -85.50 | -85.86 | -87.74 |  | -87.10 | 3.60 |
| PAD_3 | n28 | -86.26 | -89.34 | -89.23 | -86.19 | -86.49 | -88.36 |  | -87.65 | 3.15 |
| Device | Band | TRMS offset [dBm/15kHz] |  |  |  |  |  | Pass/fail limit |  |  |
|  |  | Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 |  |  |  |
| PAD_1 | n28 | 0.52 | -0.27 | -0.58 | 1.59 | 0.63 | -1.90 | +/- 0.75 MU, i.e., +/- 1.96 dB |  |  |
| PAD_2 | n28 | 0.64 | -0.83 | -2.00 | 1.60 | 1.24 | -0.64 |  |  |  |
| PAD_3 | n28 | 1.38 | -1.69 | -1.59 | 1.46 | 1.15 | -0.71 |  |  |  |

Figure 3 presents the valid TRMS measurement results from all labs, and Figure 4 visualizes the deviations between valid TRMS measurement results from each lab and reference values.

Figure 2. Valid TRMS measurement results from 6 labs (without systematic offsets)

Figure 3. Deviations between valid TRMS measurement results from each lab and reference values (without systematic offsets)

Observation 1: All 6 volunteer labs have submitted measurement results for the Rel-18 FR1 MIMO OTA lab alignment in the 1st round, no outlier value is identified.

Observation 2: As agreed, the reference value is derived by the linear average (in dB) of the measurement results submitted by all labs in the 1st round. The maximum deviation of one lab between the measurement results and the reference values is beyond 0.75*MU. Other 5 labs are well aligned.

Observation 3: The 1st round of 3GPP Rel-18 FR1 MIMO OTA lab alignment activity can be concluded with 5 volunteer labs well aligned at band n28.

3	Conclusion

This contribution presents an analysis of the Rel-18 FR1 MIMO OTA lab alignment result submitted by all the 6 labs in the 1st round.

Observation 1: All 6 volunteer labs have submitted measurement results for the Rel-18 FR1 MIMO OTA lab alignment in the 1st round, no outlier value is identified.

Observation 2: As agreed, the reference value is derived by the linear average (in dB) of the measurement results submitted by all labs in the 1st round. The maximum deviation of one lab between the measurement results and the reference values is beyond 0.75*MU. Other 5 labs are well aligned.

Observation 3: The 1st round of 3GPP Rel-18 FR1 MIMO OTA lab alignment activity can be concluded with 5 volunteer labs well aligned at band n28.

References

- R4-2321137, Updated Framework and time plan for FR1 MIMO OTA performance requirements development (Nov 2023), CAICT, 3GPP RAN4#109, Nov. 2023.

- R4-2320067, CAICT Rel-18 FR1 MIMO OTA Lab Alignment PAD Results, CAICT, 3GPP RAN4#109, Nov. 2023.

- R4-2318927, PADs measurement results for n28, CMCC, BUPT, 3GPP RAN4#109, Nov. 2023.

- R4-2318269, Measurement results of Mediatek lab for 3GPP Rel-18 FR1 MIMO OTA lab alignment activity, Mediatek, 3GPP RAN4#109, Nov. 2023.

- R4-2318896, Measure results for 3GPP Rel-18 FR1 MIMO OTA Lab Alignment, Xiaomi, 3GPP RAN4#109, Nov. 2023.

- R4-2319919, 3GPP Rel-18 FR1 MIMO OTA PAD measurement, OPPO, 3GPP RAN4#109, Nov. 2023.

- R4-2403017, On FR1 MIMO OTA Lab Alignment PADs measurement results, Apple, 3GPP RAN4#110, Feb. 2024.

- R4-2400534, Update to FR1 Channel model validation, MVG, Spirent, Apple, ETS-Lindgren, Telecom Italia, Sony, 3GPP RAN4#110, Feb. 2024.

- R4-2402876, Way forward on [110][332] NR_MIMO_OTA_enh, CAICT, 3GPP RAN4#110, Feb. 2024.

- R4-2321085, WF on [109][336] NR_MIMO_OTA_enh, CAICT, 3GPP RAN4#109, Nov. 2023.

- R4-2403011, Ad-hoc minutes for [110][332] NR_MIMO_OTA_enh, CAICT, 3GPP RAN4#110, Feb. 2024.