3GPP TSG-RAN WG4 Meeting #79	R4-163897

Nanjing, China, 23 - 27 May, 2016

Agenda Item:	6.2.2	Conformance requirements

Source:	Kathrein

Title:	Text proposal for chapter 10.3.2.1 (Test methods)

Document for:	Approval

# Introduction

The TR 37.842 [1] has to describe in sub-clause 10 the potential test methodologies for the AAS testing in detail. In this TP, we provide the required uncertainty budget for the one dimensional compact range chamber.

It is proposed to approve this TP and include it in the TR.

# References

[1]	TR 37.842,  Radio Frequency (RF) requirement background for Active Antenna System (AAS) Base Station (BS)

# Text proposal TR 37.842

10.3.2 OTA sensitivity

10.3.2.1 Test methods

##### 10.3.2.1.3 One Dimensional Compact Range Chamber

###### 10.3.2.1.3.3 Uncertainty budget format

Editor: It is proposed to change chapter C1 “Indoor Anechoic Chamber” to “General Chamber Uncertainties”.

| UID | Description of uncertainty contribution | Details in paragraph | Uncertainty [dB] |
| --- | --- | --- | --- |
| Stage 1, DUT measurement |  |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | C1-1 | 0 (0.01°) |
| 2 | Pointing misalignment between the AAS BS and the reference antenna. | C1-2 | 0 (0.01°) |
| 3 | Quality of quiet zone | C1-3 | 0.0325 |
| 4 | Polarization mismatch between the AAS BS and the transmitting antenna | C1-4 | 0.05 (2°) |
| 5 | Mutual coupling between the AAS BS and the transmitting antenna | C1-5 | 0 |
| 6 | Phase curvature | C1-6 | 0 |
| 7 | Uncertainty of the signal generator | C1-11 | 0.1 |
| Stage 2, Calibration measurement |  |  |  |
| 8 | Impedance mismatch in the transmitting chain | C1-8 | 0.0056 |
| 9 | Positioning and pointing misalignment of the reference antenna and the transmitting antenna | C1-9 | 0 |
| 10 | Impedance mismatch at the reference antenna and measurement equipment. | C1-10 | 0.0029 |
| 11 | Quality of quiet zone | C1-3 | 0.178 |
| 12 | Polarization mismatch for reference antenna | C1-4 | 0.05 (2°) |
| 13 | Mutual coupling between the reference antenna and the transmitting antenna | C1-5 | 0 |
| 14 | Phase curvature | C1-6 | 0 |
| 15 | Uncertainty of the measurement equipment | C1-7 | 0.1 |
| 16 | Uncertainty of the RF signal source generator | C1-11 | 0.25 |
| 17 | Influence of the reference antenna feed cable - Flexing cables, adapters, attenuators, connector repeatability | C1-12 | 0.082 |
| 18 | Influence of the transmitting antenna feed cable - Flexing cables, adapters, attenuators, connector repeatability | C1-13 | 0.082 |
| 19 | Uncertainty of the absolute gain of the reference antenna | C1-14 | 0.5 |