## Slide 1: Measurement uncertainty for EIRP
3GPP TSG-RAN WG4 Meeting #84-bis
Dubrovnik, Croatia, 9 – 13 October, 2017
Agenda item: 9.8.2.2
Document for: Discussion
R4-1710386
Intel

## Slide 2: Measurement Setup EIRP using a Power Meter @ 60GHz
DUT
Array
Power Sensor
Antenna array position within UE is known
Open Ended Waveguide Probe (OEWP) is used as Measurement Antenna
OEWP is aligned with the direction of the main beam.
OEWP is directly connected to a Power Sensor

## Slide 3: Measurement Uncertainty EIRP using a Power Meter @ 60GHz
Source of Uncertainty | Value [dB] | Distribution | Divisor | Std Uncert [dB] | Comments
Measurement Antenna Gain | 0.93 | Normal | 2 | 0.47 | 3 Antenna Gain method
Mismatch VSWR Antenna – Power Meter | 0.26 | U-shaped | √2 | 0.18 | Measurement Antenna VSWR=1.5 / Power Meter VSWR=1.35
Power Meter Accuracy | 0.25 | Rectangular | √3 | 0.14 | Power meter specifications. Absolute level accuracy
Range Length (Near/Far Field condition) | 0.00 | Rectangular | √3 | 0.00 | Measurement is done on Far field conditions
Position of the phase center (Measurement Antenna) | 0.13 | Rectangular | √3 | 0.08 | Introduced as a result of the inaccurate positioning of the phase center of the receiving antenna
Ambient temperature impact | 0.10 | Actual | 1 | 0.10 | Influence of the ambient temperature on the output power of the carrier
Repeatability | 0.50 | Actual | 1 | 0.50 | Repeatability of Measurement based on statistics (Class A)
Zero Offset | 0.14 | Rectangular | √3 | 0.08 | Power meter specifications. Zero offset impact on the measurement
Measurement Noise | 0.02 | Rectangular | √3 | 0.01 | Power meter specifications. Sensitivity, sampling rate, aperture time, etc.
Combined Standard Uncertainty |  |  |  | 0.74 | 
Expanded Uncertainty (k=2)* |  |  |  | 1.47 | 
* The expanded Measurement Uncertainty with coverage factor (k=2) corresponds to a confidence level of 95%.

## Slide 4: Black Box vs White Box approach 1
1 From R4-1708200

## Slide 5: Black Box vs White Box approach 1 Path loss difference antenna on-centre vs. off-centre
The Black Box approach yields a measurement uncertainty related to the offset of the centre of active antenna array from the centre of the quiet zone, specifically for TX powers or RX sensitivities that are defined as single-directional requirements (EIRP, EIS).
 | Path Loss Difference ≈ 20 log[r/(r-Dd)] [dB] |  |  |  |  |  |  |  |  | 
r [m] Dd [cm] | 0.5 | 0.6 | 0.7 | 0.8 | 0.9 | 1 | 1.5 | 2 | 3 | 4
0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0
5 | 0.9 | 0.8 | 0.6 | 0.6 | 0.5 | 0.4 | 0.3 | 0.2 | 0.1 | 0.1
7.5 | 1.4 | 1.2 | 1.0 | 0.9 | 0.8 | 0.7 | 0.4 | 0.3 | 0.2 | 0.2
10 | 1.9 | 1.6 | 1.3 | 1.2 | 1.0 | 0.9 | 0.6 | 0.4 | 0.3 | 0.2
15 | 3.1 | 2.5 | 2.1 | 1.8 | 1.6 | 1.4 | 0.9 | 0.7 | 0.4 | 0.3
20 | 4.4 | 3.5 | 2.9 | 2.5 | 2.2 | 1.9 | 1.2 | 0.9 | 0.6 | 0.4
25 | 6.0 | 4.7 | 3.8 | 3.3 | 2.8 | 2.5 | 1.6 | 1.2 | 0.8 | 0.6
30 | 8.0 | 6.0 | 4.9 | 4.1 | 3.5 | 3.1 | 1.9 | 1.4 | 0.9 | 0.7
1 From R4-1708200

## Slide 6: Conclusions
Observation 1: Uncertainty due to the offset of the center of the antenna array from the center of the quiet zone is a significant contributor to the overall budget; for some DUT size/measurement distance this element becomes dominant.
Observation 2: In general, large uncertainty budgets may lead to large test tolerances for the corresponding UE RF requirements; this can create discrepancies between product design requirements (for UEs and network equipment) and the specification.
Proposal 1: It is proposed to define two options of the MU budget; baseline MU assuming black box DUT positioning and optimized MU assuming white box DUT positioning.
Proposal 2: Optimized MU is applicable to DUTs where the centers of the antenna arrays are known to the testing lab.
Proposal 3: The test tolerances are derived assuming optimized MU; if a DUT is tested in a setup using black box DUT positioning, then the impact of additional MU is taken on as additional risk.