TSG-RAN Working Group 4 (Radio) meeting #85	R4-1713564

Reno, USA, 27th November- 1st December 2017

Source:	Ericsson

Title:	Measurement uncertainty budget for ACLR for CATR

Agenda item:	8.27.4.1

Document for:	Discussion

# Introduction

This contribution presents an uncertainty budget for ACLR measurement in a Compact Antenna Test Range which maybe is a method for conformance testing in [2].

# Discussion

Description of uncertainty contribution can be found in [3].

From framework on uncertainty budget for ACLR in [4] following measurement uncertainty budget can be estimated.

| ACLR uncertainty budget for Compact Antenna Test Range |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3 GHz < f ≦ 4.2 GHz | Uncertainty value 4.2 GHz < f ≦ 6 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3 GHz < f ≦ 4.2 GHz | Standard uncertainty ui [dB] 4.2 GHz < f ≦ 6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 |
| 5 | QZ ripple with DUT | 0.0928 | 0.0928 | 0.0928 | Normal | 1 | 1 | 0.0928 | 0.0928 | 0.0928 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| 6 | Network Analyzer | 0.13 | 0.20 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 | 0.20 |
| 7 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.127 | 0.325 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 | 0.23 |
| 8 | Insertion loss variation in receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 | 0.10 |
| [TBD] | TRP Estimation Accuracy |  |  |  | Normal | 1 | 1 | [TBD] | [TBD] | [TBD] |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  |  |  |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  |  |  |

# Conclusion

This contribution presents measurement uncertainty budget for ACLR measurement in a Compact Antenna Test Range.

# References

[1] 	TS 37.105, Active Antenna System (AAS) Base Station (BS) transmission and reception

[2]	 TS 37.145-2, Active Antenna System (AAS) Base Station (BS) conformance testing; Part 2: radiated conformance testing

[3] 	TS 37.842, Evolved Universal Terrestrial Radio Access (E-UTRA) and Universal Terrestrial Radio Access (UTRA; Radio Frequency (RF) requirement background for Active Antenna System (AAS) Base Station (BS)

[4]	R4-1705386, “Framework on Uncertainty Budget for ACLR”, Ericsson