## Slide 1: 3GPP TSG-RAN4 Meeting RAN4#86 Tdoc: R4-1802988 Athens, Greece, 26 February – 2 March 2018
Source: Keysight Technologies
Title: Proposals for Characterization of Quiet Zone at mm-wave NR frequencies for Combined axis Positioner
Agenda item: 7.10.2
Document for: Approval

## Slide 2: Objective of this Contribution
For below 6GHz, characterization of the quiet zone test procedures, also commonly referred to as ‘ripple test’ have been defined by CTIA (normative) [1] and 3GPP (informative) [2, 3].
  The outcome of the ripple test is a surface standard deviation that effectively predicts the variation of the TRP/TRS measurements, spherical surface integrals of EIRP/EIS, when the DUT is placed anywhere within the quiet zone.
In R4-1711278, QZ quality characterization procedure to capture the impact of amplitude variation in the test zone was described
This procedure is most suited for positioners using distributed Azimuth – Elevation axis as shown in R4-1713682 of partial test results using the described procedure
The objective of this contribution is to introduce a ‘Characterization of the quality of Quiet Zone’ test procedure for TRP/TRS (3D surface integrals) and for EIRP/EIS (single point power measurements) for combined-axis Azimuth – Roll type of positioner for mmwave frequencies combined with the MU of offset between UE antenna array and center of quiet zone [4,5]. This procedure is applicable to the any FF measurement setup.
Once agreement is reached, a TP will be created

## Slide 3: CTIA Ripple Test Procedure for combined axis system
A characterization of the quiet zone test is generally only performed once when the chamber is commissioned
The surface standard deviation of the pattern variations corresponds to the MU element of the ‘Quality of the Quiet Zone’

## Slide 4: Use of Directional Antennas
In R4-1800849, a mask was proposed for a directional reference antenna which was agreed.

## Slide 5: MU for TRP/TIS based Efficiency Measurements
As described in TP R4-1711958, the quality of the QZ measurements for integrated RF parameters should be based on 3D pattern measurements
  A signal source is connected to the reference antenna via a coaxial cable
  A measurement receiver is connected to the measurement antenna (ideally, a multi-port VNA is used to simultaneously measure the q and f polarizations)
Suitable TRP/TRS grids have been proposed and studied in [R4-1800952] but more work is planned, e.g., adaptive grids, before proposing a grid
Proposal 1: The quality of the QZ measurements shall be based on efficiency measurements, i.e., full 3D pattern measurements as in R4-1711958
Proposal 2: the 3D pattern measurement grid is FFS

## Slide 6: QZ Quality for combined axis positioner
P3
P1
P2
P5
P6
P7
To Measurement Antenna
Phi Axis
Theta Axis
P4
Number of positions to be measured
Proposal 3: Use 7 reference antenna positions to determine the quality of the quiet zone and the effect of offsetting the reference antenna from the centre of the quiet zone to the dimension of the QZ
Observation 1: For a complete measurement a total of 49 different reference antenna orientations, i.e., 7 different Theta axis orientations, qref = 0o, 45o, 90o, 135o, 225o, 270o, 315o ,180o and 7 different Phi axis orientations fref = 0o, 45o, 90o, 135o, 225o, 270o, and 315o
z
y
x

## Slide 7: Positions
qref = 90o
fref = 0o
qref = 45o
fref = 45o
fref = 90o
fref = 270o
qref = 0o
fref = 135o
fref = 315o
fref = 225o
qref = 315o
qref = 135o

## Slide 8: Test point reduction
qref | 135o | 90o | 45o | 0o | 315o | 270o | 225o
135o |  |  | X | X | X |  | 
90o |  |  | X | X | X |  | 
45o |  |  | X | X | X |  | 
0o |  | X | X | X | X | X | 
225o |  |  | X | X | X |  | 
270o |  |  | X | X | X |  | 
315o |  |  | X | X | X |  | 
180o | X | X | X |  | X | X | X
fref
To keep the number of test points limited and still be able to capture the worst case scenario, the following table recommends reduction of the number of test point
Proposal 4: Use the 29 orientations mentioned in Table 1 for the antenna positions as in Proposal 3.
Table 1

## Slide 9: Reference Antenna Polarization
As agreed in R4-1711958, QZ procedure is to be done separately for both the polarizations.
Proposal 5: Characterize the QZ quality for q and f-polarized field components individually
q-Pol
f-Pol
q-Pol
f-Pol
Requires two separate antenna orientations  Double the test time!
Requires a single antenna orientation
q-Pol
f-Pol

## Slide 10: Calculation of Combined MU for TRP/TRS
The combined MU element related to the quality of the quiet zone for TRP/TRS and offset between UE antenna array and centre of quiet zone is the standard deviation of the all the positions and orientations.
Proposal 6: The MU element quantifying the quality of the quiet zone for TRP/TRS combined with the MU element of offset between UE antenna array and center of quiet zone is the standard deviation of the all the efficiency measurement results as described in the procedure (Proposal 3,4,5)

## Slide 11: Calculation of Combined MU for EIRP/EIS
The combined MU of quality of the QZ and offset between UE antenna array and centre of QZ for EIRP and EIS can be determined from a subset of the measurements taken to determine the combined MU of TRP/TRS, specifically the respective peak gain measurements
  The combined MU element corresponds to the standard deviation of the peak gain radiated by the reference antenna radiated in the respective reference orientations
Proposal 7: The MU element quantifying the quality of the quiet zone for EIRP/EIS combined with the MU element of offset between UE antenna array and centre of quiet zone is the standard deviation of the all the peak gain measurements as described in the procedure (Proposal 3,4,5)

## Slide 12: Proposals
Proposal 1: The quality of the QZ measurements shall be based on efficiency measurements, i.e., full 3D pattern measurements as in R4-1711958
Proposal 2: the 3D pattern measurement grid is FFS
Proposal 3: Use 7 reference antenna positions to determine the quality of the quiet zone and the effect of offsetting the reference antenna from the centre of the quiet zone to the dimension of the QZ
Proposal 4: Use the 29 orientations mentioned in Table 1 for the antenna positions as in Proposal 3.
Proposal 5: Characterize the QZ quality for q and f-polarized field components individually
Proposal 6: The MU element quantifying the quality of the quiet zone for TRP/TRS combined with the MU element of offset between UE antenna array and center of quiet zone is the standard deviation of the all the efficiency measurement results as described in the procedure (Proposal 3,4,5)
Proposal 7: The MU element quantifying the quality of the quiet zone for EIRP/EIS combined with the MU element of offset between UE antenna array and centre of quiet zone is the standard deviation of the all the peak gain measurements as described in the procedure (Proposal 3,4,5)

## Slide 13: References
[1] CTIA, “Test Plan for Wireless Device Over-the-Air Performance” Version 3.7, June 2017
[2] 3GPP TR 25.914 “Measurements of radio performances for UMTS terminals in speech mode”
[3] 3GPP TS 34.114, “User Equipment (UE) / Mobile Station (MS) Over The Air (OTA) antenna performance; Conformance testing”
[4] R4-1708200, Black Box vs White Box Testing for NR, Rohde & Schwarz, TSG RAN WG4 Meeting #84, Berlin, China, August 2017
[5] R4-1711958, TP (TR38.810) on Quality of Quiet Zone charecterization for FR2, Rohde & Schwarz, September, 2017
[6] R4-1707756, WF on NR MU and test tolerance, CATR, Intel Corporation, Rohde & Schwarz, Anritsu, Fraunhofer HHI, 3GPP TSG-RAN WG4 Meeting #84, Berlin, Germany, August, 2017
[7] R4-1709769, On Measurement Grids for mm-wave NR, Rohde & Schwarz, 3GPP TSG RAN WG4 NR#3, Nagoya, Japan, September 2017