3GPP TSG-RAN WG4 Meeting #90bis	R4-1904341

Xi’an, Shaanxi, China, 8 – 12 April 2019

Source:	Nokia, Nokia Shanghai Bell

Title:	Uncertainty budget values for the FR2 transmitter ON/OFF power requirement

Agenda item:	6.8.5.3.3.1

Document for:	Discussion

# 1	Introduction

During the RAN4#90 meeting, the way forward on FR2 transient period [1] was approved, which captured agreements and open issues to be addressed at forthcoming meetings.

This document assesses the uncertainty budget for OTA measurements of the FR2 transmitter ON/OFF power test requirement in a CATR chamber.

# 2	Discussion

From [1], the following agreements are captured:

- Measuring TX OFF power level with assumption OFF ant gain = ON antenna gain (= full array gain) is feasible

- The conformance EIRP OFF level for the transient test will be Core TRP level plus the ON antenna directivity.

Based on the above agreements, the TRP off level can be mathematically expressed as

TRPOFF=EIRPOFFDON			(in linear units)									(1)

where EIRPOFF is the EIRP off level measured during the transmitter OFF period

TRPOFF is the TRP off level measured during the transmitter OFF period

DON is the antenna directivity measured during the transmitter ON period

DON can be expressed as

DON=EIRPONTRPON  																	(2)

where EIRPON is the EIRP on level measured during the transmitter ON period

TRPON is the TRP level measured during the transmitter ON period, which could be the TRP power measured for the FR2 OTA BS output power requirement or it is measured using the same TRP method that is for the FR2 OTA BS output power measurement (see 3GPP TS 38.141-2, Annex I for more details)

Substituting Equation (2) into Equation (1) and rearranging, we get

TRPOFF=EIRPOFFEIRPON.TRPON													(3)

As can be observed from Equation (3), the MU budget for TRPOFF consists of the following:

- MUEIRPOFFEIRPON – as this is a relative term, most of the measurement uncertainty values cancel out, but there are other uncertainty contributors due to the transmitter transient period measurement, which include the limiter uncertainty because a limiter is used to protect test equipment from high transmitter ON power levels

- MU(TRPON) – if the same TRP method is used to measure TRPON, then the measurement uncertainty value of FR2 OTA BS output power can be adopted

From 3GPP TR 38.817-02, the measurement uncertainty of FR2 OTA BS output power is shown in Table 1. An estimated limiter uncertainty of 0.5 dB [2] is added to the table. Further, the spectrum analyser uncertainty is revised for transmitter transient period measurements.

From 3GPP TR 38.817-02, 12.10, the FR2 TRP SE is 1.2 dB. The MU for FR2 transmitter ON/OFF power is

- MU = 2.502+1.22 = 2.8 dB, 	24.25 GHz < f ≤ 29.5 GHz

- MU = 2.522+1.22 = 2.8 dB, 	37 GHz < f ≤ 40 GHz

Since the transmitter OFF power is not a regulatory requirement, it is recommended to set TT = MU.

Table 1: Compact antenna test range uncertainty assessment for EIRP measurements for FR2 BS output power

| UID (Note 1) | Uncertainty Source | Uncertainty value 24.25<f <29.5GHz | Uncertainty value 37<f <40GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)(dB) 24.25<f <29.5GHz | Standard uncertainty ui (dB)(dB) 37<f <40GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| E2-1 | Misalignment DUT & pointing error | 0.35 | 0.35 | Exp. normal | 2 | 1 | 0.174 | 0.174 |
| E2-2 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.9 | 0.9 | Gaussian | 1 | 1 | [0.9] | [0.9] |
| E2-3 | Standing wave between DUT and test range antenna | 0.03 | 0.03 | U-shaped | √2 | 1 | 0.02 | 0.02 |
| E2-4 | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | Normal | 1 | 1 | 0.01 | 0.01 |
| E2-5 | QZ ripple with DUT | 0.0928 | 0.0928 | Normal | 1 | 1 | 0.4 | 0.4 |
| E2-16 | Frequency flatness | 0.25 | 0.25 | Normal | 1 | 1 | 0.25 | 0.25 |
| E2-YY | Limiter uncertainty | 0.5 | 0.5 | normal | 1 | 1 | [0.5] | [0.5] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| E2-6 | Network Analyzer | 0.3 | 0.3 | Normal | 1 | 1 | 0.3 | 0.3 |
| E2-7 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.43 | 0.57 | U-shaped | √2 | 1 | 0.21 | 0.26 |
| E2-8 | Insertion loss variation in receiver chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0 | 0 |
| E2-4 | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | Normal | 1 | 1 | 0.01 | 0.01 |
| E2-9 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | √2 | 1 | 0.15 | 0.2 |
| E2-10 | SGH Calibration uncertainty | 0.52 | 0.52 | Rectangular | √3 | 1 | 0.3 | 0.3 |
| E2-11 | Misalignment positioning system | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0 | 0 |
| E2-1 | Misalignment SGH and pointing error | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 |
| E2-12 | Rotary joints | 0 | 0 | U-shaped | √2 | 1 | 0 | 0 |
| E2-3 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.06 | 0.06 |
| E2-5 | QZ ripple with SGH | 0.009 | 0.009 | Normal | 1 | 1 | 0.009 | 0.009 |
| E2-15 | Switching uncertainty | 0.43 | 0.43 | Rectangular | √3 | 1 | 0.1 | 0.1 |
| Combined standard uncertainty (1σ) (dB)(dB) |  |  |  |  |  |  | 1.27 | 1.29 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB)(dB) |  |  |  |  |  |  | [2.50] | [2.52] |
| Note 1: UID are referenced to TR 37.843 [9]. |  |  |  |  |  |  |  |  |

# 3	Conclusions

The document has outlined an analysis for the calculation of the measurement uncertainty budget values for transmitter OFF power in a CATR chamber. Based on our analysis, the MU budget is as follows:

- MU = [2.8] dB, 	24.25 GHz < f ≤ 29.5 GHz

- MU = [2.8] dB, 	37 GHz < f ≤ 40 GHz

Since the FR2 transmitter OFF power is not a regulatory requirement, it is recommended to set TT = MU, i.e.,

- TT = [2.8] dB, 	24.25 GHz < f ≤ 29.5 GHz

- TT = [2.8] dB, 	37 GHz < f ≤ 40 GHz

# References

- R4-1902286, WF on FR2 transient period, Keysight

- R4-1901549, Discuss FR2 TX OFF MU, Huawei