3GPP TSG-RAN4 Meeting #92	R4-1907993

Ljubljana, Slovenia, 26th Aug 2019 - 30th Aug 2019

Agenda Item:		6.1.3.5

Source: 		Rohde & Schwarz

Title: 		2D Compact Range for BS Testing – MU budget analysis

Document for:		Approval

# Introduction

This contribution provides a thorough analysis of the Measurement Uncertainty budget applicable to the 2D Compact Range system presented in [1] and further defined in [2].

- Using the CATR uncertainty budget for EIRP and EIS measurements currently available in [4] and [5] as a reference, the following sections in this contribution present a study of the uncertainty value for each of the terms together with the corresponding description of those additional terms applicable to the plane wave converter methodology.

# MU budgets

The tables 2-1 and 2-2 below present the summary list of MU terms for both EIRP and EIS measurements.

| UID | Uncertainty Source | Details in section |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | 3.1 |
| 2 | Measurement Receiver | 3.2 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 3.3 |
| 4 | RF leakage (calibration antenna connector terminated) | 3.4 |
| 5 | Normalized QZ ripple with DUT | 3.5 |
| 6 | Miscellaneous Uncertainty | 3.6 |
| 7 | Frequency flatness | 3.7 |
| 8 | Field imbalance | 3.8 |
| Stage 1: Calibration measurement |  |  |
| 9 | Calibration equipment (i.e. reference receiver, measurement receiver and generator) | 3.9 |
| 10 | Mismatch (i.e. reference receiver, reference antenna, generator and reference cable) | 3.10 |
| 11 | Insertion loss variation in receiver chain | 3.11 |
| 4 | RF leakage (calibration antenna connector terminated) | 3.4 |
| 12 | Influence of the calibration antenna feed cable | 3.12 |
| 13 | Uncertainty of the absolute gain of the calibration antenna | 3.13 |
| 14 | Misalignment of positioning system | 3.14 |
| 15 | Misalignment of calibration antenna & pointing error | 3.1 |
| 16 | Rotary joints | 3.15 |
| 17 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 3.3 |
| 18 | Normalized QZ ripple with calibration antenna | 3.5 |
| 19 | Switching uncertainty | 3.16 |
| 20 | Field calibration repeatability | 3.17 |

Table 2-1: 2D Compact Range uncertainty budget format for EIRP measurements

| UID | Uncertainty Source | Details in section |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | 3.1 |
| 21 | RF signal generator | 3.18 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 3.3 |
| 4 | RF leakage (calibration antenna connector terminated) | 3.4 |
| 5 | Normalized QZ ripple with DUT | 3.5 |
| 6 | Miscellaneous Uncertainty | 3.6 |
| 7 | Frequency flatness | 3.7 |
| 8 | Field imbalance | 3.8 |
| Stage 1: Calibration measurement |  |  |
| 9 | Calibration equipment (i.e. reference receiver and generator) | 3.9 |
| 22 | Mismatch (i.e. reference receiver, reference antenna, generator and reference cable) | 3.10 |
| 11 | Insertion loss variation in receiver chain | 3.11 |
| 4 | RF leakage (calibration antenna connector terminated) | 3.4 |
| 12 | Influence of the calibration antenna feed cable | 3.12 |
| 13 | Uncertainty of the absolute gain of the calibration antenna | 3.13 |
| 14 | Misalignment of positioning system | 3.14 |
| 15 | Misalignment of calibration antenna & pointing error | 3.1 |
| 16 | Rotary joints | 3.15 |
| 17 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 3.3 |
| 18 | Normalized QZ ripple with calibration antenna | 3.5 |
| 19 | Switching uncertainty | 3.16 |
| 20 | Field calibration repeatability | 3.17 |

Table 2-2: 2D Compact Range uncertainty budget format for EIS measurements

# Error contribution descriptions

The following sections describe in detail the evaluation for each of the MU terms in Tables 2-1 and 2-2 applicable to R&S implementation of PWC. The following assumptions are common to all of them:

- RF equipment

  - Reference receiver: R&S NRQ6

  - Measurement receiver: R&S FSW8 or R&S NRQ6

  - Signal generator: R&S SMW200A

  - Vector network analyzer: R&S ZVA8

- Antennas

  - Calibration antenna: R&S PWC-CAA1 (active Vivaldi antenna)

  - DUT Antenna:

    - MARS MA-WA36-DP25 (3.2 to 3.8 GHz)

    - MARS MA-WA25-DP23 (2.3 to 2.7 GHz)

- Test settings

  - Minimum set of frequencies: 2350, 2600, 3500 and 3750 MHz

  - Signal bandwidth <= 100 MHz

  - Test distance = 1500 mm

## Misalignment DUT/calibration antenna & pointing error

Following section E2-1 in [4], this error is evaluated using the DUT or calibration antenna by taking the maximal difference between PDUT or PCAL measured for a roll angle of +2° or -2° and PDUT or PCAL measured for a roll angle of 0° for all considered frequencies. This evaluation is a worst case compared to the positioner repeatability, equal 0.05º.

A rectangular distribution is assumed for this uncertainty.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Misalignment DUT & pointing error | 0.100 | 0.100 | dB | Rectangular | √3 | 1 | 0.058 | 0.058 |
| 15 | Misalignment of calibration antenna & pointing error | 0.050 | 0.050 | dB | Rectangular | √3 | 1 | 0.029 | 0.029 |

## Measurement Receiver

As described in [2], a relative measurement approach is used following the description in section G.4 in [6] where the same receiver is used for both calibration and DUT measurements steps. Therefore, the final DUT measurement becomes relative and the absolute accuracy of the measurement receiver is cancelled out.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2 | Measurement Receiver | 0.000 | 0.000 | dB | Gaussian | 1 | 1 | 0.000 | 0.000 |

## Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT/calibration antenna

This MU term covers both the standing wave between DUT or calibration antenna and the test range antenna as per section E2-3 in [4], but it also counts for the PWC imperfect field synthesis over distance.

It is evaluated by means of an S21 measurement between the PWC and the DUT or calibration antenna, where the measurement distance is varied ±35mm from the reference at 1500mm. This variation corresponds to ~λ02 for the lowest frequency range (i.e. 2300 MHz).

The maximum coupling deviation for each center frequency is calculated as follows, assuming a rectangular distribution:

ufc,m=maxdi20logS21di,fc,mS211500,fc,m

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | dB | Rectangular | √3 | 1 | 0.029 | 0.081 |
| 17 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | dB | Rectangular | √3 | 1 | 0.069 | 0.069 |

## RF leakage

This contribution denotes interference and noise SI+Nleaking in to connector and cable(s) between test range antenna and receiving/transmitting equipment. Assuming a worst case SINR ≥ 40dB, the error is calculated as follows:

uSINR=20log1+10-SINR20.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4 | RF leakage | 0.086 | 0.086 | dB | Normal | 1 | 1 | 0.086 | 0.086 |

Figure 3.4-1 shows the setup to verify this maximum SINR in case of Tx measurements, where the reference antenna connector cable is loaded with 50 ohm instead of connecting the calibration antenna (point F). PRX is the power received during range path loss calibration.

Figure 3.4-1: SINR check setup for Tx measurements

Figure 3.4-2 shows the corresponding setup to verify the maximum SINR in case of Rx measurements, where the PWC connector cable is loaded with 50 ohm instead of connecting the PWC (point C). PRX is the power received during range path loss calibration.

Figure 3.4-2: SINR check setup for Rx measurements

## Normalized QZ ripple with DUT/calibration antenna

This term describe the maximum antenna to antenna coupling error due to the synthesized field distribution E over the QZ surface different in relation to the plane wave distribution EREF=1 taken as reference. Only dominant field component is taken as E.

This term depends on:

- Field distribution in the QZ surface:

  - imperfect field synthesis by PWC200

  - room reflections

  - signal bandwidth

  - misalignment of PWC and DUT coordinate systems

  - accuracy of near-field scanner (negligible)

- aperture size of CAL antenna

- aperture efficiency of CAL antenna

As described in [7], the metric to assess the quality of the Quiet Zone for PWC differs from classical systems. In this case:

- Selection of metric depends on the Antenna Under Test aperture efficiency ηCAL

ηCAL=maximum effective areaphysical area=Aeff,calAp,cal=λ2D0,cal4π1Ap,cal

- Electric field deviation δE does not depend on mean magnitude or phase difference between the evaluation field E and the reference field EREF, therefore normalization factor c is introduced

δE=(E*c)-EREF

where electric fields E,EREF are taken only over MN sampling points included in the physical area of the AUT and c is chosen such that, δE2 is minimum.

- Normalized QZ ripple for antennas with high aperture efficiency ηCAL≅1 at discrete frequency fn is

δ2=1MNδE2E

        - Normalized antenna coupling error is ∆VV≤δ2ηCAL

ufn≤20log10V-∆VV=20log101-δ2ηCAL

- Normalized QZ ripple for antennas with low aperture efficiency ηCAL≪1 at discrete frequency fn is

δ∞=δE∞E

        - Antenna coupling error is ∆VV≤δ∞

ufn≤20log10V-∆VV=20log101-δ∞

- Finally average coupling error u is calculated for all uniformly distributed frequencies fn within signal bandwidth

u=20log1Nn=1N10ufn20

Using the calibration antenna described at the beginning of section 3, an aperture size of 40 mm x 170 mm is assumed which provides very high aperture efficiency ηCAL≅1.

In the case of the DUT antennas described at the beginning of section 3, a 60 cm x 60 cm aperture is assumed, providing an aperture efficiency close to 1 for both DUT antennas.

Therefore, in all cases the normalized QZ ripple formula with δ2 is used:

ufn=20log101-δ2

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 5 | Normalized QZ ripple with DUT | 0.66 | 0.92 | dB | Rectangular | √3 | 1 | 0.381 | 0.531 |
| 18 | Normalized QZ ripple with calibration antenna | 0.22 | 0.24 | dB | Rectangular | √3 | 1 | 0.127 | 0.139 |

The field distribution measurements (magnitude and phase) are performed using an Open Ended Waveguide Probe (OEWP) and a Vector Network Analyzer (e.g. R&S ZVA8) in order to get results like the one shown in figure 3.5-1 and 3.5-2, where the evaluation areas for calibration and DUT antenna are marked with black and red rectangles, respectively.

Figure 3.5-1: QZ field distribution; f = 2600 MHz.

Figure 3.5-2: QZ field distribution; f = 3750 MHz.

## Miscellaneous Uncertainty

The term 'miscellaneous uncertainty' is used to define all the unknown, unquantifiable, etc. uncertainties associated with EIRP/EIS measurements.

According to Appendix G.13 in [6], a fixed value of 0.2 dB with a normal distribution shall be incorporated in the budget for the miscellaneous uncertainty.

Measurements of the repeatability of the DUT measurement were carried out to confirm the value of 0.2 dB given in [6]. Before each measurement, the DUT antenna is removed from the positioner. After mounting the DUT on the positioner again and aligning it with the measurement antenna, the power received by the spectrum analyser is measured.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 6 | Miscellaneous Uncertainty | 0.200 | 0.200 | dB | Normal | 1.000 | 1 | 0.200 | 0.200 |

## Frequency flatness

The error contribution due to frequency flatness and interpolation error is calculated using equation G.33 in [6]:

εj=10logk=j-N2j+N2PLk(N+1)PLj

with εj the expected relative error in the average power result for a given channel in dB, PLj the linear path loss at the center frequency of the given channel, PLk the linear path loss at each frequency point across the corresponding channel and N the number of frequency steps across a given channel bandwidth.

The data in dB are linearized with the following equation.

PLk=10-PLk, dB10

As worst case, εj is calculated for a center frequency of 3700 MHz with a bandwidth of 100 MHz. A rectangular distribution is assumed for this uncertainty.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 6 | Frequency flatness | 0.130 | 0.130 | dB | Rectangular | √3 | 1 | 0.075 | 0.075 |

## Field imbalance

This term describes offset of DUT measurement result due to different aperture size between calibration and DUT antennas. Difference in average of incident field level over aperture sizes creates this offset. Only dominant field component is considered. This uncertainty is not included in the normalized QZ ripple term.

ufn=20log10EAUTECAL

EAUT=1MAUTNAUTm,n=1MAUT,NAUTEAUTm,n

ECAL=1MCALNCALm,n=1MCAL,NCALECALm,n

Finally the imbalance error u is calculated for all uniformly distributed frequencies fn within signal bandwidth

u=20log1Nn=1N10ufn20

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 6 | Field imbalance | 0.170 | 0.100 | dB | Rectangular | √3 | 1 | 0.098 | 0.058 |

## Calibration equipment

As described in [2], a relative measurement approach is used following the description in section G.4 and G.5 in [6] where the same receiver/generator is used for both calibration and DUT measurements steps. Therefore, the final DUT measurement becomes relative and the absolute accuracy of the measurement receiver/generator is cancelled out. In order to achieve this a high accuracy reference receiver (e.g. R&S NRQ6) is used to minimize the measurement uncertainty.

This uncertainty term is calculated as RSS of the following terms:

- Signal generator: absolute output level. The absolute output level of the signal generator stays constant during both measurement stages. This uncertainty is set to zero.

- Signal generator: output level stability. The output level stability of the signal generator is measured connecting a power sensor at the output port of the generator. The power in dB is measured at five different time steps. The uncertainty is determined by taking the maximal difference in the measurements results for each considered frequency. A normal distribution is assumed for this uncertainty.

- Measurement receiver: absolute level. The error due to the absolute level of the spectrum analyzer is cancelled due to its presence in both path loss calibration and DUT measurement. This uncertainty is set to zero. Not applicable to sensitivity measurements.

- Measurement receiver: linearity. The error due to the linearity of the spectrum analyzer is also set to zero assuming that the power received by the spectrum analyzer in both measurement stages is assumed constant for all considered frequencies. Not applicable to sensitivity measurements.

- Reference receiver: absolute level. The error due to the absolute level of the power sensor is taken from the manufacturer’s datasheet [11]. Assuming an absolute value of 0 dBm, the error is equal to 0.11 dB. An expanded normal distribution is assumed for this uncertainty.

- Reference receiver: mismatch loss. The VSWR of the power sensor is equal to 1.2 according to the manufacturer’s datasheet [11]. The absolute value of the complex reflection coefficient of the port of the power sensor ΓPS is calculated as per section 3.10. The following equation is used to estimate the mismatch loss of the power sensor, and a U-shaped distribution is assumed for εVSWR,PS.

εVSWR, PS=20log1+ΓPS

- Reference receiver: linearity. The error due to the linearity of the power sensor is set to zero: the power sensor is used for single power level measurement only for each considered frequency.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 9 | Calibration equipment (i.e. reference receiver, measurement receiver and generator) | 0.059 | 0.059 | dB | Normal | 1 | 1 | 0.059 | 0.059 |

## Mismatch

In general, the absolute value of the complex reflection coefficient Γi can be calculated from the VSWR, or the Return Loss, according to:

Γi=VSWR-1VSWR+1=10-RL20

Equation G.1 from [6] is used to calculate the error due to mismatch the 2 ends of a RF path:

εVSWR=20log1+Γ1∙Γ2∙S21∙S12

As described for the previous term, the mismatch error is also minimized because of the relative measurement approach. Therefore, this term is calculated as RSS of the following terms:

- Mismatch between the reference receiver and the reference cable. The reflection coefficient of the reference receiver and the generator, together with the path loss between these two elements is used to calculate this term.

- Mismatch between the reference cable and the reference antenna. The reflection coefficient of the reference antenna and the generator, together with the path loss between these two elements is used to calculate this term.

- Mismatch between the measurement receiver/generator and the measurement cable. This mismatch is present during path loss calibration and the DUT measurement as there is no movement of the cable connecting both devices and no connector change. Thus the uncertainty is set to zero.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 10 | Mismatch (i.e. reference receiver, reference antenna, generator and reference cable) | 0.06 | 0.05 | dB | U-shaped | √2 | 1 | 0.041 | 0.034 |
| 22 | Mismatch (i.e. reference receiver, reference antenna, generator and reference cable) | 0.11 | 0.11 | dB | U-shaped | √2 | 1 | 0.078 | 0.078 |

## Insertion loss variation in receiver chain

The insertion loss of the measurement antenna cable stays constant during both measurement stages: there is no movement of the cable and no connector change during measurements. This uncertainty is set to zero.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 11 | Insertion loss variation in receiver chain | 0.00 | 0.00 | dB | Rectangular | √3 | 1 | 0.00 | 0.00 |

## Influence of the calibration antenna feed cable

This uncertainty term is calculated as RSS of the following items, assuming a :

- Insertion loss variation due to adaptor. The reference antenna input connector is different from the reference receiver, so an adaptor is needed only for the first step of the range path calibration.

- Insertion loss variation due to the reference cable. This is assessed by repeated measurements while flexing the cable reference cable. The largest difference between the results is recorded as the uncertainty.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 12 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | dB | Rectangular | √3 | 1 | 0.060 | 0.060 |

## Uncertainty of the absolute gain of the calibration antenna

The calibration antenna used for the PWC is a dedicated active Vivaldi antenna as described at the beginning of section 3. Therefore, the reference antenna calibration is performed under very well controlled conditions where the MU is reduced. An expanded normal distribution is assumed for this uncertainty.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 13 | Uncertainty of the absolute gain of the calibration antenna. | 0.140 | 0.140 | dB | Exp. Normal | 2 | 1 | 0.070 | 0.070 |

## Misalignment of positioning system

The misalignment between the measurement antenna and calibration antenna is negligible due to the low directivity of the reference antenna. An expanded normal distribution is assumed for this uncertainty.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 14 | Misalignment of positioning system | 0.00 | 0.00 | dB | Exp. Normal | 2 | 1 | 0.00 | 0.00 |

## Rotary joints

This contribution is not applicable since the PWC antenna is single polarized and only one path loss calibration is performed.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 14 | Rotary joints | 0.00 | 0.00 | dB | U-shaped | √2 | 1 | 0.00 | 0.00 |

## Switching uncertainty

This term is evaluated using the Tx range calibration setup described in [2] by measuring the value of the power PDUT received by the spectrum analyzer five times after switching several times the switching unit (i.e. R&S OSP).

The uncertainty is determined by taking the maximal difference between the five different measurements. A rectangular distribution is assumed for this uncertainty.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 19 | Switching uncertainty | 0.020 | 0.020 | dB | Rectangular | √3 | 1 | 0.012 | 0.012 |

## Field calibration repeatability

Each execution of field calibration of the measurement antenna array to find the PWC settings provides a slightly different set of settings for the RF components for each antenna path. This results in variation of the synthesized plane wave in the QZ and variation of PWC antenna to reference antenna coupling. This variation u is described by field calibration repeatability term.

In other to determine this term, the following step by step process is followed:

- Perform field calibration

- Enable PWC setting

- Measure S21 between the PWC and calibration reference antenna at the PWC setting frequency

- Repeat field calibration and S21 measurement for all PWC setting frequencies at least I=5 times

- Maximum error is determined as follows.

u=maxi20logS21i-mini20logS21i

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 20 | Field calibration repeatability | 0.13 | 0.31 | dB | Rectangular | √3 | 1 | 0.075 | 0.179 |

## RF signal generator

As described in [2], a relative measurement approach is used following the description in section G.5 in [6] where the same generator is used for both calibration and DUT measurements steps. Therefore, the final DUT measurement becomes relative and the absolute accuracy of the RF signal generator receiver is cancelled out.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 21 | RF signal generator | 0.000 | 0.000 | dB | Gaussian | 1 | 1 | 0.000 | 0.000 |

# MU values

As an outcome of previous sections, tables 4-1 and 4-2 show the final MU value tables applicable to the PWC methodology.

| UID | Uncertainty Source | Uncertainty Value |  | Distribution | Divisor | ci | Std. uncertainty ui [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |  |  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 2 | Measurement Receiver | 0.00 | 0.00 | Gaussian | 1 | 1 | 0.00 | 0.00 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | Rectangular | √3 | 1 | 0.03 | 0.08 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 5 | Normalized QZ ripple with DUT | 0.66 | 0.92 | Rectangular | √3 | 1 | 0.38 | 0.53 |
| 6 | Miscellaneous Uncertainty | 0.20 | 0.20 | Normal | 1 | 1 | 0.20 | 0.20 |
| 7 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| 8 | Field imbalance | 0.17 | 0.10 | Rectangular | √3 | 1 | 0.10 | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 9 | Calibration equipment (i.e. reference receiver, measurement receiver and generator) | 0.059 | 0.059 | Normal | 1 | 1 | 0.06 | 0.06 |
| 10 | Mismatch (i.e. reference receiver, reference antenna, generator and reference cable) | 0.06 | 0.05 | U-shaped | √2 | 1 | 0.04 | 0.03 |
| 11 | Insertion loss variation in receiver chain | 0.00 | 0.00 | Rectangular | √3 | 1 | 0.00 | 0.00 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 12 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 13 | Uncertainty of the absolute gain of the calibration antenna | 0.14 | 0.14 | Exp. normal | 2 | 1 | 0.07 | 0.07 |
| 14 | Misalignment of positioning system | 0.00 | 0.00 | Exp. normal | 2 | 1 | 0.00 | 0.00 |
| 15 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 16 | Rotary joints | 0.00 | 0.00 | U-shaped | √2 | 1 | 0.00 | 0.00 |
| 17 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | Rectangular | √3 | 1 | 0.07 | 0.07 |
| 18 | Normalized QZ ripple with calibration antenna | 0.22 | 0.24 | Rectangular | √3 | 1 | 0.13 | 0.14 |
| 19 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 20 | Field calibration repeatability | 0.13 | 0.31 | Rectangular | √3 | 1 | 0.08 | 0.18 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.52 | 0.66 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.01 | 1.29 |

Table 4-1: 2D Compact Range uncertainty values for EIRP measurements

| UID | Uncertainty Source | Uncertainty Value |  | Distribution | Divisor | ci | Std. uncertainty ui [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |  |  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 21 | RF signal generator | 0.00 | 0.00 | Gaussian | 1 | 1 | 0.00 | 0.00 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | Rectangular | √3 | 1 | 0.03 | 0.08 |
| 4 | RF leakage (measurement antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 5 | Normalized QZ ripple with DUT | 0.66 | 0.92 | Rectangular | √3 | 1 | 0.38 | 0.53 |
| 6 | Miscellaneous Uncertainty | 0.20 | 0.20 | Normal | 1 | 1 | 0.20 | 0.20 |
| 7 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| 8 | Field imbalance | 0.17 | 0.10 | Rectangular | √3 | 1 | 0.10 | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 9 | Calibration equipment (i.e. reference receiver and generator) | 0.059 | 0.059 | Normal | 1 | 1 | 0.06 | 0.06 |
| 22 | Mismatch (i.e. reference receiver, reference antenna, generator and reference cable) | 0.11 | 0.11 | U-shaped | √2 | 1 | 0.08 | 0.08 |
| 11 | Insertion loss variation in receiver chain | 0.00 | 0.00 | Rectangular | √3 | 1 | 0.00 | 0.00 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 12 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 13 | Uncertainty of the absolute gain of the calibration antenna | 0.14 | 0.14 | Exp. normal | 2 | 1 | 0.07 | 0.07 |
| 14 | Misalignment of positioning system | 0.00 | 0.00 | Exp. normal | 2 | 1 | 0.00 | 0.00 |
| 15 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 16 | Rotary joints | 0.00 | 0.00 | U-shaped | √2 | 1 | 0.00 | 0.00 |
| 17 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | Rectangular | √3 | 1 | 0.07 | 0.07 |
| 18 | Normalized QZ ripple with calibration antenna | 0.22 | 0.24 | Rectangular | √3 | 1 | 0.13 | 0.14 |
| 19 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 20 | Field calibration repeatability | 0.13 | 0.31 | Rectangular | √3 | 1 | 0.08 | 0.18 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.52 | 0.66 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.02 | 1.30 |

Table 4-2: 2D Compact Range uncertainty values for EIS measurements

# Conclusion

In this contribution, in addition to [1] and [2], an extensive description of the MU for the 2D Compact Range has been provided. It is proposed to incorporate the MU budgets, as detailed in Tables 2-1 and 2-2, and the MU values, as detailed in Tables 4-1 and 4-2, into TR 37.843.

Proposal: incorporate Tables 2-1, 2-2, 4-1 and 4-2 into TR 37.843.

# References

- R4-1806605, “2D Compact Range for Testing of AAS Basestations”, Rohde & Schwarz, RAN4#87, May 2018

- R4-1907992, “2D Compact Range for BS Testing – Calibration and Measurement procedure”,  Rohde & Schwarz, RAN4#92, August 2019

- R4-1908987, “CR to TR 37.843: Addition of 2D Compact Range. MU, calibration and test procedure for EIRP and EIS”,  Rohde & Schwarz, RAN4#92, August 2019

- 3GPP TR 37.843: “Radio Frequency (RF) requirement background for Active Antenna System (AAS) Base Station (BS) radiated requirements”.

- 3GPP TR 37.842: “Radio Frequency (RF) requirement background for Active Antenna System (AAS) Base Station (BS)”.

- CTIA Test Plan for Wireless Device Over-the-Air Performance, v3.8.2

- O. M. Bucci, M. D. Migliore, G. Panariello and D. Pinchera; “Plane-Wave Generators: Design Guidelines,  Achievable Performances and Effective Synthesis”; IEEE Transactions on Antennas and Propagation, April 2013

- R.  Haupt,  “Generating  a  Plane  Wave  in  the  Near  Field  with  a  Planar Array Antenna”, Microwave Journal, September 2003

- C. Rowell, A. Tankielun, “Plane Wave Converter for 5G Massive MIMO Basestation Measurements”, the 12th European Conference on Antennas and Propagation (EuCAP2018), London, UK, 9 - 13 April, 2018

- B. Derat, “Creating Far-field Condition at a Tenth of Far-field Distance: An Innovative Technique for 5G OTA measurement”, Electronic Design Innovation Conference (EDICON), Beijing, China, 20 – 22 March, 2018

- R&S®NRQ6 Frequency Selective Power Sensor - Product Brochure (with complete specifications), V 03.00, Apr 04, 2019 - PD 3607.1888.12.