3GPP TSG-RAN4 Meeting #92	R4-1908987

# Ljubljana, Slovenia, 26th Aug 2019 - 30th Aug 2019

| CR-Form-v12.0 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 37.843 | CR | 0029 | rev | - | Current version: | 15.4.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | CR to TR 37.843: Addition of 2D Compact Range. MU, calibration and test procedure for EIRP and EIS. |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Rohde & Schwarz |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R4 |  |  |  |  |  |  |  |  |  |
| Work item code: | AASenh_BS_LTE_UTRA-Perf |  |  |  |  |  | Date: |  |  | 2019-08-16 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-15 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) Rel-12 (Release 12) Rel-13 (Release 13) Rel-14 (Release 14) Rel-15 (Release 15) Rel-16 (Release 16) |  |
| Reason for change: |  | The addition of the 2D Compact Range as accepted methodology for radiated testig of AAS BS requires modifications of some clauses. |  |  |  |  |  |  |  |  |
| Summary of change: |  | 2D Compact Range system description, together with the calibration and test procedure added to the corresponding clauses for EIRP and EIS test cases. The corresponding MU evaluation and MU terms description added. New reference added to clause 2. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | 2D Compact Range could not be used as methodology for AAS BS conformance testing. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | Clauses 2, 10.2.2, 10.3.2 and Annex E |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | - |  |  |  |  |  |  |  |  |

| This CR's revision history: | First revision. |
| --- | --- |

< Unchanged Text Deleted >

< Beginning of Changes >

# 2	References

The following documents contain provisions which, through reference in this text, constitute provisions of the present document.

-	References are either specific (identified by date of publication, edition number, version number, etc.) or non-specific.

-	For a specific reference, subsequent revisions do not apply.

-	For a non-specific reference, the latest version applies. In the case of a reference to a 3GPP document (including a GSM document), a non-specific reference implicitly refers to the latest version of that document in the same Release as the present document.

[1]	3GPP TR 21.905: "Vocabulary for 3GPP Specifications"

[2]	RP-170844: "Work Item Description for Enhancements of Base Station (BS) RF and EMC requirements for Active Antenna System (AAS)"

[3]	3GPP TS 37.105: "Active Antenna System (AAS) Base Station (BS) transmission and reception"

[4]	3GPP TR 37.842: "Frequency (RF) requirement background for Active Antenna System (AAS) Base Station (BS)"

[5]	3GPP TR 36.942: "Evolved Universal Terrestrial Radio Access (E-UTRA); Radio Frequency (RF) system scenarios"

[6]	3GPP TR 37.840: "Study of Radio Frequency (RF) and Electromagnetic Compatibility (EMC) requirements for Active Antenna Array System (AAS) base station"

[7]	3GPP TS 37.104: "E-UTRA, UTRA and GSM/EDGE; Multi-Standard Radio (MSR) Base Station (BS) radio transmission and reception"

[8]	3GPP TS 36.104: "Evolved Universal Terrestrial Radio Access (E-UTRA); Base Station (BS) radio transmission and reception"

[9]	Recommendation ITU-R SM.328-11: "Spectra and bandwidth of emissions", ITU/R

[10]	3GPP TR 25.951: "FDD Base Station (BS) classification"

[11]	3GPP TR 37.809: "Evolved Universal Terrestrial Radio Access (E-UTRA) medium range and Multi-Standard Radio (MSR) medium range / local area Base Station (BS) class requirements"

[12]	3GPP TR 36.931: "Evolved Universal Terrestrial Radio Access (E-UTRA); Radio Frequency (RF) requirements for LTE Pico Node B"

[13]	3GPP TS 37.114: "Active Antenna System (AAS) Base Station (BS) Electromagnetic Compatibility (EMC)"

[14]	3GPP TS 37.113: "E-UTRA, UTRA and GSM/EDGE; Multi-Standard Radio (MSR) Base Station (BS) Electromagnetic Compatibility (EMC)"

[15]	3GPP TS 36.113: "Evolved Universal Terrestrial Radio Access (E-UTRA); Base Station (BS) and repeater ElectroMagnetic Compatibility (EMC)"

[16]	3GPP TS 25.113: "Base Station (BS) and repeater ElectroMagnetic Compatibility (EMC)"

[17]	Conversion of power to field strength in the book "Automotive Antenna Design and Applications", by Victor Rabinovich, Nikolai Alexandrov, Basim Alkhateeb; 2010

[18]	Recommendation ITU-R SM.329-10: "Unwanted emissions in the spurious domain"

[19]	3GPP TS 25.104: "Universal Terrestrial Radio Access (UTRA); Base Station (BS) radio transmission and reception (FDD)"

[20]	3GPP TS 25.105: "Universal Terrestrial Radio Access (UTRA); Base Station (BS) radio transmission and reception (TDD)"

[21]	3GPP TS 25.141: "Base Station (BS) conformance testing (FDD)"

[22]	3GPP TS 36.141: "Evolved Universal Terrestrial Radio Access (E-UTRA); Base Station (BS) conformance testing"

[23]	3GPP TS 37.145-1: "Active Antenna System (AAS) Base Station (BS) conformance testing; Part 1: Conducted conformance testing"

[24]	3GPP TS 37.145-2: "Active Antenna System (AAS) Base Station (BS) conformance testing; Part 2: Radiated conformance testing"

[25]	IEEE Std 149: "IEEE Standard Test Procedures for Antennas", IEEE

[26]	TSGR4#8(99)631: "Antenna-to-Antenna Isolation Measurements", Allgon

[27]	Report ITU-R M.2244: "Isolation between antennas of IMT base stations in the land mobile service", ITU-R

[28]	IEC 61000-4-3: 2006+AMD1:2007+AMD2:2010: “Electromagnetic compatibility (EMC) - Part 4-3: Testing and measurement techniques - Radiated, radio-frequency, electromagnetic field immunity test”

[29]	Hald, J.; Hansen, J. E.; Jensen, F. & Holm Larsen, F. Hansen, J. (Ed.) Spherical Near-Field Antenna Measurements, Peter Peregrinus Ltd., 1998, vol 26.

[30]	"Sparse Sampling Analysis Tool" Matlab code: https://se.mathworks.com/matlabcentral/fileexchange/67143-sparse-sampling-analysis-tool?s_tid=srchtitle

[31]	J. Fridén, A. Razavi, and A. Stjernman, “Angular sampling, Test Signal, and Near-Field Aspects for Over-the-Air Total Radiated Power Assessment in Anechoic Chambers”, IEEE Access, 2018, https://ieeexplore.ieee.org/stamp/stamp.jsp?arnumber=8470084.

[32]	IEEE Std 145: “IEEE Standard Definitions of Terms for Antennas”, IEEE

[33]	IEC 61000-4-21. “Electromagnetic compatibility (EMC) - Part 4-21: Testing and measurement techniques - Reverberation chamber test methods” Edition 2.0 2011-01, The International Electrotechnical Commision (IEC), 2011

[34]	Hill, D.A., “Boundary fields in reverberation chambers,” IEEE Transactions on Electromagnetic Compatibility, vol. 47, no. 2, pp. 281-290, May 2005

[35]	Krauthäuser, H. G.; Winzerling, T.; J., N.; Eulig, N. & Enders, A. “Statistical interpretation of autocorrelation coefficients for fields in mode-stirred chambers” 2005 International Symposium on Electromagnetic Compatibility, 2005. EMC 2005., 2005, 2, 550-555

[36]	O. M. Bucci, M. D. Migliore, G. Panariello and D. Pinchera; “Plane-Wave Generators:  Design  Guidelines,  Achievable  Performances  and  Effective Synthesis”; IEEE Transactions on Antennas and Propagation,  April 2013

< End of Changes >

< Unchanged Text Deleted >

< Beginning of Changes >

#### 10.2.2.3	2D Compact Range

##### 10.2.2.3.1	Description

This method is a natural extension of a compact antenna test range (CATR) where the reflector is replaced by a phased antenna array, or Plane Wave Converter (PWC), consisting of phase shifters, attenuators, and an antenna array capable of forming planar waves within a specified quiet zone in front of the antenna in a similar manner as the reflector in a CATR system. The plane wave converter converts field modes of a test signal from a plane wave in free-space into a TEM mode in a coaxial cable or vice versa.

The principle of the 2D Compact Range is shown in figure 10.2.2.3.1-1:

Figure 10.2.2.3.1-1: 2D Compact Range test system diagram

A phased antenna array is comprised of a set of antenna elements with periodic separations where the phase and amplitude of each signal is controlled by a signal distribution network (SDN) in order to create a high directivity radiation pattern in the far field. Using similar technique of setting the phase and amplitude of the phased array signals, a quiet zone containing planar waves as a linear superposition of the array elements' spherical radiation waves can be created within the near field of the phased array.

Considering a phased planar array of N antenna elements with phase and amplitude weights, a quiet zone (defined as plane waves with constant phase and amplitude) can be formed in front of the PWC at the distance R.

##### 10.2.2.3.2	Calibration

The calibration measurement is done by using a reference antenna with known efficiency or gain values. In the calibration measurement the reference antenna is measured in the same place as the DUT, and the attenuation of the complete transmission path (D↔A, see figure 10.2.2.3.1-1) from the DUT to the measurement receiver is calibrated out. In order to minimize the measurement uncertainty due to the receiver/transmitter absolute accuracy, a relative measurement approach is used.

The step-by-step procedure is as follows:

- Connect the reference receiver to the RF output port of the signal generator, as shown in figure 10.2.2.3.2-1, using a reference cable and a matching attenuator.

Figure 10.2.2.3.2-1: Absolute Power Calibration with Reference Receiver

- Set the frequency fn and the output power Pout of CW signal at the output E of the signal generator.

- Measure PRX'fn at the interface F which is the input of the reference receiver.

- Repeat step 2 and 3 for all N frequencies fn, n=1,2,…,N within signal bandwidth BW with sufficiently small step size, e.g. ∆f=1 MHz, such as N∆f=BW. N shall be an odd number.

- Replace the reference receiver with the reference antenna at the interface F and connect the measurement receiver to the RF output port C of the PWC as shown in figure 10.2.2.3.2-2. The phase center of the reference antenna A shall be aligned with the PWC coordinate system B. The interface A shall coincide with the origin of the DUT positioner coordinate system. Realized gain of the reference antenna GREFfn in the direction towards PWC antenna is known.

Figure 10.2.2.3.2-2: Tx range path loss Calibration with Reference Receiver

- Configure the PWC setting corresponding to the center frequency fPWC=fN+12 within the signal bandwidth and the prepared QZ target field. Keep the same setting of the center frequency and output power Pout of CW signal at the signal generator.

- Measure PRXfn at the interface D which is the input of the measurement receiver.

- Calculate path loss value PLEIRPfn between the interfaces A and D for EIRP measurement. Note that PLEIRPfn includes absolute level uncertainty of the measurement receiver which is cancelled out during EIRP measurement.
PLEIRPfn=10logPRX'fnGREFfnPRXfn [dB]

##### 10.2.2.3.3	Procedure

- Install the AAS BS as DUT with the manufacturer declared coordinate system reference point in the same place as the phase center of the reference antenna A as shown in figure 10.2.2.3.3-1. The manufacturer declared coordinate system orientation of the AAS BS is set to be aligned with the testing system.

Figure 10.2.2.3.3-1: 2D Compact Range setup for EIRP measurements

- Set the AAS BS to generate the tested beam with the appropriate test model with the beam peak direction intended to be the same as the testing direction.

- Set the AAS BS to generate the tested beam with the appropriate test model with the beam peak direction intended to be the same as the testing direction.

- Measure PDUT at the interface D which is the input of the measurement receiver.

- Calculate EIRP value at the beam peak direction of DUT:
EIRP=10logPDUTPLEIRP[dB]

- Repeat steps 2-5 for all conformance test beam direction pairs and test conditions.

##### 10.2.2.3.4	MU assessment

###### 10.2.2.3.4.1	MU Budget

Table 10.2.2.3.4.1-1: 2D Compact Range uncertainty contributions for EIRP measurements

| UID | Uncertainty Source | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | E6-1 |
| 2 | Measurement Receiver | E6-2 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | E6-3 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-4 |
| 5 | Normalized QZ ripple with DUT | E6-5 |
| 6 | Miscellaneous Uncertainty | E6-6 |
| 7 | Frequency flatness | E6-7 |
| 8 | Field imbalance | E6-8 |
| Stage 1: Calibration measurement |  |  |
| 9 | Calibration equipment (i.e. reference receiver, measurement receiver and generator) | E6-9 |
| 10 | Mismatch (i.e. reference receiver, reference antenna, generator and reference cable) | E6-10 |
| 11 | Insertion loss variation in receiver chain | E6-11 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-4 |
| 12 | Influence of the calibration antenna feed cable | E6-12 |
| 13 | Uncertainty of the absolute gain of the calibration antenna | E6-13 |
| 14 | Misalignment of positioning system | E6-14 |
| 15 | Misalignment of calibration antenna & pointing error | E6-1 |
| 16 | Rotary joints | E6-15 |
| 17 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | E6-3 |
| 18 | Normalized QZ ripple with calibration antenna | E6-5 |
| 19 | Switching uncertainty | E6-16 |
| 20 | Field calibration repeatability | E6-17 |

###### 10.2.2.3.4.2	MU Value

Table 10.2.2.3.4.2-1: 2D Compact Range uncertainty assessment for EIRP measurements

| UID | Uncertainty Source | Uncertainty Value |  | Distribution of the probability | Divisor based on distribution shapre | ci | Std. uncertainty ui [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |  |  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 2 | Measurement Receiver | 0.00 | 0.00 | Gaussian | 1 | 1 | 0.00 | 0.00 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | Rectangular | √3 | 1 | 0.03 | 0.08 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 5 | Normalized QZ ripple with DUT | 0.66 | 0.92 | Rectangular | √3 | 1 | 0.38 | 0.53 |
| 6 | Miscellaneous Uncertainty | 0.20 | 0.20 | Normal | 1 | 1 | 0.20 | 0.20 |
| 7 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| 8 | Field imbalance | 0.17 | 0.10 | Rectangular | √3 | 1 | 0.10 | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 9 | Calibration equipment (i.e. reference receiver, measurement receiver and generator) | 0.059 | 0.059 | Normal | 1 | 1 | 0.06 | 0.06 |
| 10 | Mismatch (i.e. reference receiver, reference antenna, generator and reference cable) | 0.06 | 0.05 | U-shaped | √2 | 1 | 0.04 | 0.03 |
| 11 | Insertion loss variation in receiver chain | 0.00 | 0.00 | Rectangular | √3 | 1 | 0.00 | 0.00 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 12 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 13 | Uncertainty of the absolute gain of the calibration antenna | 0.14 | 0.14 | Exp. normal | 2 | 1 | 0.07 | 0.07 |
| 14 | Misalignment of positioning system | 0.00 | 0.00 | Exp. normal | 2 | 1 | 0.00 | 0.00 |
| 15 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 16 | Rotary joints | 0.00 | 0.00 | U-shaped | √2 | 1 | 0.00 | 0.00 |
| 17 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | Rectangular | √3 | 1 | 0.07 | 0.07 |
| 18 | Normalized QZ ripple with calibration antenna | 0.22 | 0.24 | Rectangular | √3 | 1 | 0.13 | 0.14 |
| 19 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 20 | Field calibration repeatability | 0.13 | 0.31 | Rectangular | √3 | 1 | 0.08 | 0.18 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.52 | 0.66 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.01 | 1.29 |

< End of Changes >

< Unchanged Text Deleted >

< Beginning of Changes >

### 10.3.2	OTA sensitivity

#### 10.3.2.1	General

The OTA sensitivity MU was derived in TR 37.842 [4] as follows:

Table 10.3.2-1: Test system specific measurement uncertainty values for the EIS test

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 1.22 | 1.25 |
| Compact Antenna Test Range | 1.33 | 1.40 |
| One Dimensional Compact Range Chamber | 1.29 | 1.43 |
| Near Field Test Range | 1.24 | 1.24 |
| 2D Compact Range | 1.02 | 1.30 |
| Common maximum accepted test system uncertainty | 1.3 | 1.4 |

#### 10.3.2.2	2D Compact Range

##### 10.3.2.2.1	Calibration

The calibration measurement is done by using a reference antenna with known efficiency or gain values. In the calibration measurement the reference antenna is measured in the same place as the DUT, and the attenuation of the complete transmission path (D↔A, see figure 10.2.2.3.1-1) from the DUT to the measurement receiver is calibrated out. In order to minimize the measurement uncertainty due to the receiver/transmitter absolute accuracy, a relative measurement approach is used.

The step-by-step procedure is as follows:

- Connect the signal generator to the input C of PWC and the input D of the reference receiver to the RF output port F of the reference antenna, as shown in figure 10.3.2.2.1-1. The phase center of the reference antenna A shall be aligned with the PWC coordinate system B. The interface A shall coincide with the origin B of the DUT positioner coordinate system. Realized gain of the reference antenna GREFfn in the direction towards PWC antenna is known.

Figure 10.3.2.2.1-1: Rx range path loss Calibration with Reference Receiver

- Configure the PWC setting corresponding to the center frequency fPWC=fN+12 within the signal bandwidth and the prepared QZ target field.

- Set the frequency fn and the output power Pout1 of CW signal at the output E of the signal generator.

- Measure PRXfn at the interface F which is the input of the reference receiver.

- Repeat step 3 and 4 for all N frequencies fn, n=1,2,…,N within signal bandwidth BW with sufficiently small step size, e.g. ∆f=1 MHz, such as N∆f=BW. N shall be an odd number.

- Calculate path loss value PLEISfn between the interfaces A and E for EIS measurement. Note that PLEISfn includes absolute level uncertainty of the signal generator which will be cancelled out during EIS measurement.
PLEISfn=10logPout1fnGREFfnPRXfn [dB]

##### 10.3.2.2.2	Procedure

- Install AAS BS as DUT with the manufacturer declared coordinate system reference point in the same place as the phase center of the reference antenna A as shown in figure 10.3.2.2.2-1. The manufacturer declared coordinate system orientation of the AAS BS is set to be aligned with the testing system.

Figure 10.3.2.2.2-1: 2D Compact Range setup for EIS measurements

- Set the AAS BS to generate the tested beam with the appropriate test model with the beam peak direction intended to be the same as the testing direction.

- Reduce output power Pout2 at the interface E which is the output of the signal generator until sensitivity threshold will be reported by DUT.

- Calculate EIS value at the beam peak direction of DUT:
EIS=10logPout2PLEIS [dB]

- Repeat steps 2-4 for all conformance test beam direction pairs and test conditions.

##### 10.3.2.2.3	MU assessment

###### 10.3.2.2.3.1	MU Budget

Table 10.3.2.2.3.1-1: 2D Compact Range uncertainty contributions for EIS measurements

| UID | Uncertainty Source | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | E6-1 |
| 21 | RF signal generator | E6-18 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | E6-3 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-4 |
| 5 | Normalized QZ ripple with DUT | E6-5 |
| 6 | Miscellaneous Uncertainty | E6-6 |
| 7 | Frequency flatness | E6-7 |
| 8 | Field imbalance | E6-8 |
| Stage 1: Calibration measurement |  |  |
| 9 | Calibration equipment (i.e. reference receiver and generator) | E6-9 |
| 22 | Mismatch (i.e. reference receiver, reference antenna, generator and reference cable) | E6-10 |
| 11 | Insertion loss variation in receiver chain | E6-11 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-4 |
| 12 | Influence of the calibration antenna feed cable | E6-12 |
| 13 | Uncertainty of the absolute gain of the calibration antenna | E6-13 |
| 14 | Misalignment of positioning system | E6-14 |
| 15 | Misalignment of calibration antenna & pointing error | E6-1 |
| 16 | Rotary joints | E6-15 |
| 17 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | E6-3 |
| 18 | Normalized QZ ripple with calibration antenna | E6-5 |
| 19 | Switching uncertainty | E6-16 |
| 20 | Field calibration repeatability | E6-17 |

###### 10.3.2.2.3.2	MU Value

Table 10.3.2.2.3.2-1: 2D Compact Range uncertainty assessment for EIRP measurements

| UID | Uncertainty Source | Uncertainty Value |  | Distribution of the probability | Divisor based on distribution shapre | ci | Std. uncertainty ui [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |  |  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 21 | RF signal generator | 0.00 | 0.00 | Gaussian | 1 | 1 | 0.00 | 0.00 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | Rectangular | √3 | 1 | 0.03 | 0.08 |
| 4 | RF leakage (measurement antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 5 | Normalized QZ ripple with DUT | 0.66 | 0.92 | Rectangular | √3 | 1 | 0.38 | 0.53 |
| 6 | Miscellaneous Uncertainty | 0.20 | 0.20 | Normal | 1 | 1 | 0.20 | 0.20 |
| 7 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| 8 | Field imbalance | 0.17 | 0.10 | Rectangular | √3 | 1 | 0.10 | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 9 | Calibration equipment (i.e. reference receiver and generator) | 0.059 | 0.059 | Normal | 1 | 1 | 0.06 | 0.06 |
| 22 | Mismatch (i.e. reference receiver, reference antenna, generator and reference cable) | 0.11 | 0.11 | U-shaped | √2 | 1 | 0.08 | 0.08 |
| 11 | Insertion loss variation in receiver chain | 0.00 | 0.00 | Rectangular | √3 | 1 | 0.00 | 0.00 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 12 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 13 | Uncertainty of the absolute gain of the calibration antenna | 0.14 | 0.14 | Exp. normal | 2 | 1 | 0.07 | 0.07 |
| 14 | Misalignment of positioning system | 0.00 | 0.00 | Exp. normal | 2 | 1 | 0.00 | 0.00 |
| 15 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 16 | Rotary joints | 0.00 | 0.00 | U-shaped | √2 | 1 | 0.00 | 0.00 |
| 17 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | Rectangular | √3 | 1 | 0.07 | 0.07 |
| 18 | Normalized QZ ripple with calibration antenna | 0.22 | 0.24 | Rectangular | √3 | 1 | 0.13 | 0.14 |
| 19 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 20 | Field calibration repeatability | 0.13 | 0.31 | Rectangular | √3 | 1 | 0.08 | 0.18 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.52 | 0.66 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.02 | 1.30 |

< End of Changes >

< Unchanged Text Deleted >

< Beginning of Changes >

# E.6	2D Compact Range

E6-1 Misalignment DUT/calibration antenna & pointing error

This contribution denotes uncertainty in DUT/calibration antenna alignment and DUT/calibration antenna pointing error.  In this measurement the DUT/calibration antenna is aligned to maximum, also allowing for a zero contribution for polarization mismatch uncertainty. By adjusting for maximums to align, this contribution can be a small contribution. The calibration antenna's phase centre and polarization purity changes slightly according to the frequency. Therefore, there should be some uncertainty reserved for this. To ensure that the point error is at a minimal, this contribution should be captured using the antenna pattern cut which is broadest (in the case of the DUT this would most likely be in the azimuth domain).

E6-2 Measurement Receiver: uncertainty of the absolute level

The receiving device is used to measure the received signal level in the EIRP tests as a relative level, according to clauses 10.2.2.3.2 and 10.2.2.3.3. Therefore, the absolute accuracy of the measurement receiver is cancelled out.

E6-3 Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT/calibration antenna

This value covers the effect of standing wave between DUT or calibration antenna and the test range antenna, but also counts for the PWC imperfect field synthesis over distance. This value can be captured by moving the DUT or calibration antenna towards the test range antenna.

E6-4 RF leakage

This contribution denotes noise leaking in to connector and cable(s) between test range antenna and receiving equipment.  The contribution also includes the noise leakage between the connector and cable(s) between reference antenna and transmitting equipment.

E6-5 Normalized QZ ripple with DUT/calibration antenna

This term describes the maximum antenna to antenna coupling error due to the synthesized field distribution E over the QZ surface different in relation to the plane wave distribution EREF=1 taken as reference. Only dominant field component is taken as E.

As described in [36], the metric to assess the quality of the Quiet Zone for PWC differs from classical systems. In this case:

- Selection of metric depends on the Antenna Under Test aperture efficiency ηCAL

ηCAL=maximum effective areaphysical area=Aeff,calAp,cal=λ2D0,cal4π1Ap,cal

- Electric field deviation δE does not depend on mean magnitude or phase difference between the evaluation field E and the reference field EREF, therefore normalization factor c is introduced

δE=(E*c)-EREF

where electric fields E,EREF are taken only over MN sampling points included in the physical area of the AUT and c is chosen such that, δE2 is minimum.

- Normalized QZ ripple for antennas with high aperture efficiency ηCAL≅1 at discrete frequency fn is

δ2=1MNδE2E

- Normalized antenna coupling error is ∆VV≤δ2ηCAL

ufn≤20log10V-∆VV=20log101-δ2ηCAL

- Normalized QZ ripple for antennas with low aperture efficiency ηCAL≪1 at discrete frequency fn is

δ∞=δE∞E

- Antenna coupling error is ∆VV≤δ∞

ufn≤20log10V-∆VV=20log101-δ∞

- Finally average coupling error u is calculated for all uniformly distributed frequencies fn within signal bandwidth

u=20log1Nn=1N10ufn20

E6-6 Miscellaneous Uncertainty

The term 'miscellaneous uncertainty' is used to define all the unknown, unquantifiable, etc. uncertainties associated with EIRP measurements. This term should include truly random effects as well as systematic uncertainties, such as that arising from dissimilarity between the patterns of the reference antenna and the DUT.

E6-7 Frequency flatness

This uncertainty contribution to account for the frequency interpolation error caused by a finite frequency resolution during the calibration stage.

E6-8 Field imbalance

This term describes offset of DUT measurement result due to different aperture size between calibration and DUT antennas. Difference in average of incident field level over aperture sizes creates this offset. Only dominant field component is considered. This uncertainty is not included in the normalized QZ ripple term.

E6-9 Calibration equipment

According to clauses 10.2.2.3.2 and 10.2.2.3.3, a relative measurement approach is used where the same receiver/generator is used for both calibration and DUT measurements steps. Therefore, the final DUT measurement becomes relative and the absolute accuracy of the measurement receiver/generator is cancelled out. In order to achieve this, a high accuracy reference receiver is used to minimize the measurement uncertainty.

This uncertainty term is calculated as root sum squared (RSS) of the following terms:

- Signal generator: absolute output level. The absolute output level of the signal generator stays constant during both measurement stages. This uncertainty is set to zero.

- Signal generator: output level stability. The output level stability of the signal generator is measured connecting a power sensor at the output port of the generator. The power in dB is measured at five different time steps. The uncertainty is determined by taking the maximal difference in the measurements results for each considered frequency. A normal distribution is assumed for this uncertainty.

- Measurement receiver: absolute level. The error due to the absolute level of the measurement receiver is cancelled due to its presence in both path loss calibration and DUT measurement. This uncertainty is set to zero. Not applicable to sensitivity measurements.

- Measurement receiver: linearity. The error due to the linearity of the spectrum analyzer is also set to zero assuming that the power received by the spectrum analyzer in both measurement stages is assumed constant for all considered frequencies. Not applicable to sensitivity measurements.

- Reference receiver: absolute level. The error due to the absolute level of the power sensor is taken from the manufacturer’s datasheet. An expanded normal distribution is assumed for this uncertainty.

- Reference receiver: mismatch loss. The absolute value of the complex reflection coefficient of the port of the power sensor ΓPS is calculated as per clause E6-10. The following equation is used to estimate the mismatch loss of the power sensor, and a U-shaped distribution is assumed for εVSWR,PS.

εVSWR, PS=20log1+ΓPS

- Reference receiver: linearity. The error due to the linearity of the power sensor is set to zero: the power sensor is used for single power level measurement only for each considered frequency.

E6-10 	Mismatch

In general, the absolute value of the complex reflection coefficient Γi can be calculated from the VSWR, or the Return Loss, according to:

Γi=VSWR-1VSWR+1=10-RL20

The following equation is used to calculate the error due to mismatch at the two ends of an RF path:

εVSWR=20log1+Γ1∙Γ2∙S21∙S12

As described for E6-9, the mismatch error is also minimized because of the relative measurement approach. Therefore, this term is calculated as RSS of the following terms:

- Mismatch between the reference receiver and the reference cable. The reflection coefficient of the reference receiver and the generator, together with the path loss between these two elements is used to calculate this term.

- Mismatch between the reference cable and the reference antenna. The reflection coefficient of the reference antenna and the generator, together with the path loss between these two elements is used to calculate this term.

- Mismatch between the measurement receiver/generator and the measurement cable. This mismatch is present during path loss calibration and the DUT measurement as there is no movement of the cable connecting both devices and no connector change. Thus, the uncertainty is set to zero.

E6-11 	Insertion loss variation in receiver chain

This uncertainty is the residual uncertainty contribution coming from introducing an antenna at the end of the cable.  If this cable does not change/move between the calibration Stage 1 and the measurement Stage 2, the uncertainty is assumed to be systematic and negligible during the measurement stage. Alternatively, the insertion loss can also be calculated by taking the measurement of the cable where port 2 is the end of the cable connected to the AAS BS or calibration antenna.

IL = -20log10|S21| dB

E6-12 	Influence of the calibration antenna feed cable

This uncertainty term is calculated as RSS of the following items, assuming a rectangular distribution:

- Insertion loss variation due to adaptor. The reference antenna input connector is different from the reference receiver, so an adaptor is needed only for the first step of the range path calibration.

- Insertion loss variation due to the reference cable. This is assessed by repeated measurements while flexing the cable reference cable. The largest difference between the results is recorded as the uncertainty.

E6-13 	Uncertainty of the absolute gain of the calibration antenna

The calibration antenna used for the PWC is a dedicated antenna with a special calibration. This term correspond to the MU determined by the calibration lab, and assumes an expanded normal distribution.

E6-14 Misalignment of positioning system

This contribution originates from uncertainty in sliding position and turn table angle accuracy. If the calibration antenna is aligned to maximum this contribution can be considered negligible and therefore set to zero.

E6-15 Rotary Joints

If applicable, this uncertainty term corresponds to the accuracy in changing from azimuth to vertical measurements.

E6-16 Switching Uncertainty

The purpose of the switching unit is to switch electromechanically different RF path to different measurement instruments of different measurement modes. The electromechanical switching clearly reduces the errors arising from manual switching work. Switching is also used to measure the path loss values of each polarization component.  Even though the electromechanical switching is preferable during path loss and antenna performance measurements, some minor uncertainties can occur when the switch states are programmed to change their polarity.

E6-17 Field calibration repeatability

Each execution of field calibration of the measurement antenna array to find the PWC settings provides a slightly different set of settings for the RF components for each antenna path. This results in variation of the synthesized plane wave in the QZ and variation of PWC antenna to reference antenna coupling. This variation u is described by field calibration repeatability term.

E6-18 	RF signal generator

This uncertainty consists of the uncertainty of the signal generator used to illuminate the DUT when DUT is in receiving mode.

According to clauses 10.3.2.2.1 and 10.3.2.2.2, a relative measurement approach is used where the same generator is used for both calibration and DUT measurements steps. Therefore, the final DUT measurement becomes relative and the absolute accuracy of the RF signal generator receiver is cancelled out.

< End of Changes >