3GPP TSG-RAN WG4 Meeting #94bis-e 	R4-2004530

Online, 20 Apr - 01 May 2020

Source: 	Huawei

Title: 	TP to TR 37.941: Tx MU value derivation sub-clause 9 update (FR2)

Agenda Item:	6.19.5

Document for:	Approval

# Introduction

This TP updates the MU value derivation sections for the FR2 Tx directional requirements.

It uses the pre-approved tables agreed in R4-2004529.

# TP to TR 39.941 v1.0.0

--- Start of changes ---

#### 9.2.2.3	MU value derivation, FR1

Table 9.2.2.3-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in Indoor Anechoic Chamber (Normal test conditions, FR1).

Table 9.2.2.3-1: Indoor Anechoic Chamber measurement uncertainty value derivation for EIRP accuracy measurements, Normal test conditions, FR1

| UID |  | Uncertainty source |  | Uncertainty value |  |  |  | Distribution of the probability |  | Divisor based on distribution shape |  | ci |  | Standard uncertainty ui [dB] |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.02 | 0.02 | 0.02 |
| A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 | 0.17 | 0.17 |
| A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1.00 | 1 | 0.10 | 0.10 | 0.10 |
| A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 | 0.01 |
| A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 | 0.05 | 0.05 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 | 1 | 0.14 | 0.26 | 0.26 |
| A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.41 | 1 | 0.10 | 0.23 | 0.23 |
| A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.73 | 1 | 0.06 | 0.06 | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 | 0.04 | 0.04 |
| A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 | 0.01 |
| A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 | 0.04 | 0.04 |
| A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1.00 | 1 | 0.10 | 0.10 | 0.10 |
| A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 | 0.01 |
| A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 | 0.05 | 0.05 |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1.00 | 1 | 0.13 | 0.20 | 0.20 |
| A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 | 0.03 |
| A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1.00 | 1 | 0.06 | 0.06 | 0.06 |
| A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 | 0.03 |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.29 | 0.25 | 0.25 |
| A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  | 0.44 | 0.54 | 0.54 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | 0.87 | 1.06 | 1.06 |

--- Next change ---

#### 9.2.3.3	MU value derivation, FR1

Table 9.2.3.3-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in CATR (Normal test conditions, FR1).

Table 9.2.3.3-1: CATR MU value derivation for EIRP accuracy measurements, Normal test conditions, FR1

| UID | Uncertainty source | Uncertainty value |  |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f≤3 GHz |  | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |
| A2-1a | Misalignment DUT & pointing error for EIRP | 0.00 |  | 0.00 | 0.00 | Exp. normal | 2.00 | 1 | 0.00 | 0.00 | 0.00 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 |  | 0.26 | 0.26 | Gaussian | 1.00 | 1 | 0.14 | 0.26 | 0.26 |
| A2-2a | Standing wave between DUT and test range antenna | 0.21 |  | 0.21 | 0.21 | U-shaped | 1.41 | 1 | 0.15 | 0.15 | 0.15 |
| A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.00 |  | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A2-4a | QZ ripple DUT | 0.09 |  | 0.09 | 0.09 | Gaussian | 1.00 | 1 | 0.09 | 0.09 | 0.09 |
| A2-12 | Frequency flatness |  | 0.25 | 0.25 | 0.25 | Gaussian | 1.00 | 1 | 0.25 | 0.25 | 0.25 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Uncertainty of the network analyzer | 0.13 |  | 0.20 | 0.20 | Gaussian | 1.00 | 1 | 0.13 | 0.20 | 0.20 |
| A2-6 | Mismatch of receiver chain | 0.13 |  | 0.33 | 0.33 | U-shaped | 1.41 | 1 | 0.09 | 0.23 | 0.23 |
| A2-3 | Insertion loss variation of receiver chain | 0.18 |  | 0.18 | 0.18 | Rectangular | 1.73 | 1 | 0.10 | 0.10 | 0.10 |
| A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.00 |  | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| C1-4 | Influence of the calibration antenna feed cable: | 0.02 |  | 0.02 | 0.02 | U-shaped | 1.41 | 1 | 0.02 | 0.02 | 0.02 |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 |  | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.29 | 0.25 | 0.25 |
| A2-1b | Misalignment positioning system | 0.00 |  | 0.00 | 0.00 | Exp. normal | 2.00 | 1 | 0.00 | 0.00 | 0.00 |
| A2-9 | Misalignment of calibration antenna and test range antenna | 0.50 |  | 0.50 | 0.50 | Exp. normal | 2.00 | 1 | 0.25 | 0.25 | 0.25 |
| A2-2b | Rotary Joints | 0.05 |  | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.03 | 0.03 | 0.03 |
| A2-4b | Standing wave between calibration antenna and test range antenna | 0.09 |  | 0.09 | 0.09 | U-shaped | 1.41 | 1 | 0.06 | 0.06 | 0.06 |
| A2-11 | QZ ripple calibration antenna | 0.01 |  | 0.01 | 0.01 | Gaussian | 1.00 | 1 | 0.01 | 0.01 | 0.01 |
| A2-13 | Switching uncertainty | 0.26 |  | 0.26 | 0.26 | Rectangular | 1.73 | 1 | 0.15 | 0.15 | 0.15 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  |  | 0.57 | 0.65 | 0.65 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  | 1.11 | 1.27 | 1.27 |

#### 9.2.3.4	MU value derivation, FR2

The MU assessment was carried out using a CATR chamber only however other chamber types are not precluded if suitable MU assessment is done.

A CATR MU budget was assessed in order to determine acceptable MU for the EIRP accuracy measurement in FR2. The CATR test setup and calibration and measurement procedures for FR2 are expected to be similar to those of FR1, although the test chamber dimensions and associated MU values will scale due to the shorter wavelengths and larger relative array apertures.

Table 9.2.3.4-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in CATR (Normal test conditions, FR2).

Table 9.2.3.4-1: CATR MU value derivation for EIRP accuracy measurements, Normal test conditions, FR2

| UID |  | Uncertainty source |  | Uncertainty value |  |  | Distribution of the probability |  | Divisor based on distribution shape |  | ci |  | Standard uncertainty ui [dB] |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | 24.25<f <29.5GHz | 37<f <40GHz |  |  |  | 24.25<f <29.5GHz | 37<f <40GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| A2-1a | Misalignment DUT & pointing error (EIRP) | 0.20 | 0.20 | Exp. normal | 2.00 | 1 | 0.10 | 0.10 |
| C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) -High power | 0.50 | 0.70 | Normal | 1.00 | 1 | 0.50 | 0.70 |
| A2-2a | Standing wave between DUT and test range antenna | 0.03 | 0.03 | U-shaped | 1.41 | 1 | 0.02 | 0.02 |
| A2-3 | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | Normal | 1.00 | 1 | 0.01 | 0.01 |
| A2-4a | QZ ripple with DUT | 0.40 | 0.40 | Normal | 1.00 | 1 | 0.40 | 0.40 |
| A2-12 | Frequency flatness | 0.25 | 0.25 | Normal | 1.00 | 1 | 0.25 | 0.25 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| C1-3 | Network Analyzer | 0.30 | 0.30 | Normal | 1.00 | 1 | 0.30 | 0.30 |
| A2-5a | Mismatch of receiver chain | 0.43 | 0.57 | U-shaped | 1.41 | 1 | 0.30 | 0.40 |
| A2-6 | Insertion loss variation in receiver chain | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Normal | 1.00 | 1 | 0.01 | 0.01 |
| A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | 1.41 | 1 | 0.15 | 0.21 |
| C1-4 | SGH Calibration uncertainty | 0.52 | 0.52 | Rectangular | 1.73 | 1 | 0.30 | 0.30 |
| A2-8 | Misalignment positioning system | 0.00 | 0.00 | Exp. normal | 2.00 | 1 | 0.00 | 0.00 |
| A2-1b | Misalignment of calibration antenna and test range antenna | 0.00 | 0.00 | Exp. normal | 2.00 | 1 | 0.00 | 0.00 |
| A2-9 | Rotary joints | 0.00 | 0.00 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | U-shaped | 1.41 | 1 | 0.06 | 0.06 |
| A2-4b | QZ ripple calibration antenna | 0.01 | 0.01 | Normal | 1.00 | 1 | 0.01 | 0.01 |
| A2-11 | Switching uncertainty | 0.10 | 0.10 | Rectangular | 1.73 | 1 | 0.06 | 0.06 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.89 | 1.06 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.74 | 2.07 |

--- Next change ---

#### 9.2.4.3	MU value derivation, FR1

Table 9.2.4.3-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in One Dimensional Compact Range.

Table 9.2.4.3-1: One Dimensional Compact Range MU value derivation for EIRP accuracy measurements, FR1

|  |  |  |  | Uncertainty value |  |  |  | Distribution of the probability |  | Divisor based on distribution shape |  | ci |  | Standard uncertainty ui [dB] |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| A4-1 | Misalignment DUT and pointing error | 0.00 | 0.00 | 0.00 | Exp. normal | 2.00 | 1 | 0.00 | 0.00 | 0.00 |
| A4-2a | Standing wave between DUT and test range antenna | 0.18 | 0.18 | 0.18 | U-shaped | 1.41 | 1 | 0.13 | 0.13 | 0.13 |
| A4-3a | Quiet zone ripple DUT | 0.03 | 0.03 | 0.03 | Gaussian | 1.00 | 1 | 0.03 | 0.03 | 0.03 |
| A4-4a | Phase curvature AAS | 0.01 | 0.01 | 0.01 | Gaussian | 1.00 | 1 | 0.01 | 0.01 | 0.01 |
| A4-5a | Polarization mismatch between DUT and receiving antenna | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 | 0.03 |
| A4-6a | Mutual coupling between DUT and receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 | 1 | 0.14 | 0.26 | 0.26 |
| A4-7 | Impedance mismatch in receiving chain | 0.01 | 0.01 | 0.01 | U-shaped | 1.41 | 1 | 0.00 | 0.01 | 0.01 |
| A4-8a | RF leakage (DUT connector terminated and test range antenna connector cable terminated) | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| A4-9 | Misalignment positioning system | 0.00 | 0.00 | 0.00 | Exp. normal | 2.00 | 1 | 0.00 | 0.00 | 0.00 |
| A4-10 | Pointing error between reference antenna and test range antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| A4-11 | Impedance mismatch in path to reference antenna | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 | 0.04 | 0.04 |
| A4-12 | Impedance mismatch in path to compact probe | 0.03 | 0.03 | 0.03 | U-shaped | 1.41 | 1 | 0.02 | 0.02 | 0.02 |
| A4-2b | Standing wave between reference antenna and receiving antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.41 | 1 | 0.06 | 0.06 | 0.06 |
| A4-3b | Quiet zone ripple reference antenna | 0.18 | 0.18 | 0.18 | Gaussian | 1.00 | 1 | 0.18 | 0.18 | 0.18 |
| A4-4b | Phase curvature refernce antenna | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A4-5b | Polarization mismatch between reference antenna and receiving antenna | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 | 0.03 |
| A4-6a | Mutual coupling between reference antenna and receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 | 1 | 0.14 | 0.26 | 0.26 |
| A4-13 | Influence of the reference antenna feed cable (flexing cables, adapters, attenuators, connector repeatability) | 0.08 | 0.08 | 0.08 | Rectangular | 1.73 | 1 | 0.05 | 0.05 | 0.05 |
| A4-14 | Mismatch of receiver chain | 0.20 | 0.30 | 0.30 | U-shaped | 1.41 | 1 | 0.14 | 0.21 | 0.21 |
| A4-15 | Insertion loss of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.73 | 1 | 0.10 | 0.10 | 0.10 |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.29 | 0.25 | 0.25 |
| A4-8b | RF leakage (SGH connector terminated and test range antenna connector cable terminated. | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  | 0.46 | 0.56 | 0.56 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | 0.90 | 1.10 | 1.10 |

--- Next change ---

#### 9.2.5.3	MU value derivation, FR1

Table 9.2.5.3-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in Near Field Test Range.

Standard uncertainty values for the signal generator, network analyzer and reference antenna are according to the test equipment uncertainty values, as captured in annex C.

Table 9.2.5.3-1: NFTR measurement uncertainty value derivation for EIRP accuracy measurements, FR1

| UID | Uncertainty source |  |  | Uncertainty value |  |  |  |  |  | Distribution of the probability |  | Divisor based on distribution shape |  | ci |  | Standard uncertainty ui [dB] |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | f<3 GHz |  | 3<f<4.2 GHz |  | 4.2<f<6 GHz |  |  |  |  |  |  |  | f<3 GHz |  | 3<f<4.2 GHz |  | 4.2<f<6 GHz |  |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| A3-1 | Axes Intersection | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-2 | Axes Orthogonality | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-3 | Horizontal Pointing | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-4 | Probe Vertical Position | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-5 | Probe H/V pointing | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-6 | Measurement Distance | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-7 | Amplitude and Phase Drift | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-8 | Amplitude and Phase Noise | 0.02 | 0.02 | 0.02 | Gaussian | 1.00 | 1 | 0.02 | 0.02 | 0.02 |
| A3-9 | Leakage and Crosstalk | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-10 | Amplitude Non-Linearity | 0.04 | 0.04 | 0.04 | Gaussian | 1.00 | 1 | 0.04 | 0.04 | 0.04 |
| A3-11 | Amplitude and Phase Shift in rotary joints | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-12 | Channel Balance Amplitude and Phase | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-13 | Probe Polarization Amplitude and Phase | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-14 | Probe Pattern Knowledge | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-15 | Multiple Reflections | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-16 | Room Scattering | 0.09 | 0.09 | 0.09 | Gaussian | 1.00 | 1 | 0.09 | 0.09 | 0.09 |
| A3-17 | DUT support Scattering | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-18 | Scan Area Truncation | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-19 | Sampling Point Offset | 0.01 | 0.01 | 0.01 | Gaussian | 1.00 | 1 | 0.01 | 0.01 | 0.01 |
| A3-20 | Spherical Mode Truncation | 0.02 | 0.02 | 0.02 | Gaussian | 1.00 | 1 | 0.02 | 0.02 | 0.02 |
| A3-21 | Positioning | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.02 | 0.02 | 0.02 |
| A3-22 | Probe Array Uniformity | 0.06 | 0.06 | 0.06 | Gaussian | 1.00 | 1 | 0.06 | 0.06 | 0.06 |
| A3-23 | Mismatch of receiver chain | 0.28 | 0.28 | 0.28 | U-Shaped | 1.41 | 1 | 0.20 | 0.20 | 0.20 |
| A3-24 | Insertion loss of receiver chain | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-25 | Uncertainty of the absolute gain of the probe antenna | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 | 1 | 0.14 | 0.26 | 0.26 |
| A3-26 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | 0.15 | Gaussian | 1.00 | 1 | 0.15 | 0.15 | 0.15 |
| A3-33 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Normal | 1.00 | 1 | 0.25 | 0.25 | 0.25 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.20 | 0.20 | Gaussian | 1.00 | 1 | 0.13 | 0.20 | 0.20 |
| A3-27 | Mismatch of receiver chain | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-28 | Insertion loss of receiver chain | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-29 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | 0.02 | U-Shaped | 1.41 | 1 | 0.01 | 0.01 | 0.01 |
| A3-30 | Influence of the calibration antenna feed cable | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A3-31 | Influence of the probe antenna cable | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.29 | 0.25 | 0.25 |
| A3-32 | Short term repeatability | 0.09 | 0.09 | 0.09 | Gaussian | 1.00 | 1 | 0.09 | 0.09 | 0.09 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  | 0.52 | 0.56 | 0.56 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | 1.01 | 1.10 | 1.10 |

--- Next change ---

#### 9.2.6.3	MU value derivation, FR1

Table 9.2.6.3-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in PWS.

Table 9.2.6.3-1: PWS measurement uncertainty value derivation for EIRP accuracy measurements, FR1

| UID |  | Uncertainty source |  | Uncertainty value |  |  |  |  |  | Distribution of the probability |  | Divisor based on distribution shape |  | ci |  | Standard uncertainty ui [dB] |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | f<3 GHz |  | 3<f<4.2 GHz |  | 4.2<f<6 GHz |  |  |  |  |  |  |  | f<3 GHz |  | 3<f<4.2 GHz |  | 4.2<f<6 GHz |  |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| A7-1a | Misalignment DUT & pointing error | 0.10 | 0.10 | 0.10 | Rectangular | 1.73 | 1 | 0.06 | 0.06 | 0.06 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 | 1 | 0.14 | 0.26 | 0.26 |
| A7-2a | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | [0.14] | Rectangular | 1.73 | 1 | 0.03 | 0.08 | [0.08] |
| A7-3 | RF leakage (calibration antenna connector terminated) | 0.09 | 0.09 | 0.09 | Normal | 1.00 | 1 | 0.09 | 0.09 | 0.09 |
| A7-4a | QZ ripple with DUT | 0.42 | 0.43 | [0.43] | Rectangular | 1.73 | 1 | 0.24 | 0.25 | [0.25] |
| A7-5 | Miscellaneous Uncertainty | 0.00 | 0.00 | 0.00 | Normal | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A7-14 | System non-linearity | [0.06] | [0.06] | [0.06] | Rectangular | 1.73 | 1 | [0.04] | [0.04] | [0.04] |
| A7-13 | Frequency Flatness | 0.13 | 0.13 | 0.13 | Rectangular | 1.73 | 1 | 0.08 | 0.08 | 0.08 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.20 | 0.20 | Gaussian | 1.00 | 1 | 0.13 | 0.20 | 0.20 |
| A7-6 | Mismatch (i.e. reference antenna, network analyzer and reference cable) | 0.13 | 0.33 | 0.33 | U-shaped | 1.41 | 1 | 0.09 | 0.23 | 0.23 |
| A7-7 | Insertion loss variation | 0.18 | 0.18 | 0.18 | Rectangular | 1.73 | 1 | 0.10 | 0.10 | 0.10 |
| A7-3 | RF leakage (calibration antenna connector terminated) | 0.09 | 0.09 | 0.09 | Normal | 1.00 | 1 | 0.09 | 0.09 | 0.09 |
| A7-8 | Influence of the calibration antenna feed cable | 0.10 | 0.10 | 0.10 | Rectangular | 1.73 | 1 | 0.06 | 0.06 | 0.06 |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.29 | 0.25 | 0.25 |
| A7-9 | Misalignment of positioning system | 0.00 | 0.00 | 0.00 | Exp. normal | 2.00 | 1 | 0.00 | 0.00 | 0.00 |
| A7-1b | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 | 0.03 |
| A7-10 | Rotary joints | 0.00 | 0.00 | 0.00 | U-shaped | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| A7-2b | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | [0.12] | Rectangular | 1.73 | 1 | 0.07 | 0.07 | [0.07] |
| A7-4a | QZ ripple with calibration antenna | 0.20 | 0.20 | 0.20 | Rectangular | 1.73 | 1 | 0.12 | 0.12 | 0.12 |
| A7-11 | Switching uncertainty | 0.02 | 0.02 | 0.02 | Rectangular | 1.73 | 1 | 0.01 | 0.01 | 0.01 |
| A7-12 | Field repeatability | 0.06 | 0.12 | [0.12] | Normal | 1.00 | 1 | 0.06 | 0.12 | [0.12] |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  | [0.50] | [0.60] | [0.60] |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | [0.98] | [1.18] | [1.18] |

--- Next change ---

### 9.2.7		Maximum accepted test system uncertainty

Maximum test system uncertainties derivation methodology was described in subclause 5.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for the EIRP accuracy test can be derived from values captured in table 9.2.7-1, separately for each of the defined frequency ranges. The common maximum accepted test system uncertainty values are applicable for all test methods addressing EIRP accuracy test requirement in Normal test conditions. Based on input values in table 9.2.7-1, the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values were derived for frequency ranges as below:

- For the frequency range up to 4.2 GHz (with the breakdown point at 3 GHz), the same MU values as for E-UTRA in TS 37.145-2 [4] were adopted also for NR operation below 4.2 GHz. It is expected that the test chamber setup, calibration and measurement procedures for E-UTRA and NR will be highly similar. All uncertainty factors were judged to be the same. The MU value was thus agreed to be 1.1 dB for up to 3 GHz bands.

- For the frequency range 4.2 - 6 GHz, all MU factors including instrumentation related MU were judged to be the same as for the 3 - 4.2 GHz range, and thus the total MU for 4.2 – 6 GHz is the same as for 3 - 4.2 GHz. This assessment was made under the assumption of testing BS designed for licensed spectrum; for unlicensed spectrum the MU may differ. The MU value was thus agreed to be 1.3 dB for 3 – 6 GHz bands. The MU in 4.2 - 6 GHz is valid for BS designed to operate in licensed spectrum.

- Based on CATR inputs in subclause 9.2.3.4, for the frequency range 24.25 < f < 29.5 GHz the MU was decided to be 1.7 dB.

- Based on CATR inputs in subclause 9.2.3.4, for the frequency range 37 < f < 40 GHz the MU was decided to be 2.0 dB.

Table 9.2.7-1: OTA test system specific measurement uncertainty values for the EIRP accuracy, FR1, Normal test conditions

|  | Expanded uncertainty [dB] |  |  |
| --- | --- | --- | --- |
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |
| Indoor Anechoic Chamber | 0.87 | 1.06 | 1.06 |
| Compact Antenna Test Range | 1.11 | 1.27 | 1.27 |
| One Dimensional Compact Range Chamber | 0.90 | 1.10 | 1.10 |
| Near Field Test Range | 1.01 | 1.10 | 1.10 |
| PWS | [0.98] | [1.18] | [1.18] |
| Common maximum accepted test system uncertainty | 1.10 | 1.30 | 1.30 |

Table 9.2.7-2: OTA test system specific measurement uncertainty values for the EIRP accuracy, FR2, Normal test conditions

|  | Expanded uncertainty [dB] |  |
| --- | --- | --- |
|  | 24.25<f <29.5GHz | 37<f <40GHz |
| Indoor Anechoic Chamber |  |  |
| Compact Antenna Test Range | 1.74 | 2.07 |
| One Dimensional Compact Range Chamber |  |  |
| Near Field Test Range |  |  |
| PWS |  |  |
| Agreed value | 1.70 | 2.00 |

An overview of the MU values for all the requirements is captured in clause 16.

--- Next change ---

#### 9.3.2.3	MU value derivation, FR1

Table 9.3.2.3-1: Indoor Anechoic Chamber measurement uncertainty value derivation for EIRP accuracy measurements in Extreme test conditions, FR1

| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.02 | 0.02 | 0.02 |
| A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.30 | 0.30 | 0.30 | Rectangular | 1.73 | 1 | 0.17 | 0.17 | 0.17 |
| A1-17 | Quality of quiet zone (extreme) | 0.60 | 0.60 | 0.60 | Gaussian | 1.00 | 1 | 0.60 | 0.60 | 0.60 |
| A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 | 0.01 |
| A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 | 0.05 | 0.05 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 | 1 | 0.14 | 0.26 | 0.26 |
| A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.41 | 1 | 0.10 | 0.23 | 0.23 |
| A1-8 | Random uncertainty | 0.10 | 0.10 | 0.10 | Rectangular | 1.73 | 1 | 0.06 | 0.06 | 0.06 |
| A1-19 | Radome loss variation | 0.40 | 0.40 | 0.40 | Rectangular | 1.73 | 1 | 0.23 | 0.23 | 0.23 |
| A1-18 | Wet radome loss variation | 0.95 | 0.95 | 0.95 | Gaussian | 1.00 | 1 | 0.95 | 0.95 | 0.95 |
| A1-20 | Change in absorber behavior | 0.10 | 0.10 | 0.10 | Gaussian | 1.00 | 1 | 0.10 | 0.10 | 0.10 |
| A1-16 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1.00 | 1 | 0.25 | 0.25 | 0.25 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 | 0.04 | 0.04 |
| A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 | 0.01 |
| A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 | 0.04 | 0.04 |
| A1-3 | Quality of quiet zone | 0.10 | 0.10 | 0.10 | Gaussian | 1.00 | 1 | 0.10 | 0.10 | 0.10 |
| A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 | 0.01 |
| A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 | 0.05 | 0.05 |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.20 | 0.20 | Gaussian | 1.00 | 1 | 0.13 | 0.20 | 0.20 |
| A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 | 0.03 |
| A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1.00 | 1 | 0.06 | 0.06 | 0.06 |
| A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 | 0.03 |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.29 | 0.25 | 0.25 |
| A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  | 1.26 | 1.29 | 1.29 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | 2.46 | 2.53 | 2.53 |

--- Next change ---

.

#### 9.3.3.3	MU value derivation, FR1

Table 9.3.3.3-1: CATR MU value derivation for EIRP accuracy measurements in Extreme test conditions

| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape |  | ci |  |  | Standard uncertainty ui [dB] |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  |  |  |  | f≤3 GHz |  |  | 3<f≤4.2 GHz |  | 4.2<f≤6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| A2-1a | Misalignment DUT & pointing error for EIRP | 0.00 | 0.00 | 0.00 | Exp. normal | 2.00 |  | 1 |  |  | 0.00 |  |  | 0.00 |  | 0.00 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 |  | 1 |  |  | 0.14 |  |  | 0.26 |  | 0.26 |
| A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.41 |  | 1 |  |  | 0.15 |  |  | 0.15 |  | 0.15 |
| A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  |  | 0.00 |  |  | 0.00 |  | 0.00 |
| A2-13 | Quality of quiet zone (extreme) | 0.60 | 0.60 | 0.60 | Gaussian | 1.00 |  | 1 |  |  | 0.60 |  |  | 0.60 |  | 0.60 |
| A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1.00 |  | 1 |  |  | 0.25 |  |  | 0.25 |  | 0.25 |
| A2-14 | Wet radome loss variation | 0.40 | 0.40 | 0.40 | Rectangular | 1.73 |  | 1 |  |  | 0.23 |  |  | 0.23 |  | 0.23 |
| A2-15 | Radome loss variation | 0.95 | 0.95 | 0.95 | Gaussian | 1.00 |  | 1 |  |  | 0.95 |  |  | 0.95 |  | 0.95 |
| A2-16 | Change in absorber behavior | 0.10 | 0.10 | 0.10 | Gaussian | 1.00 |  | 1 |  |  | 0.10 |  |  | 0.10 |  | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.20 | 0.20 | Gaussian | 1.00 |  | 1 |  |  | 0.13 |  |  | 0.20 |  | 0.00 |
| A2-6 | Mismatch of receiver chain | 0.13 | 0.33 | 0.33 | U-shaped | 1.41 |  | 1 |  |  | 0.09 |  |  | 0.23 |  | 0.23 |
| A2-3 | Insertion loss variation of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.73 |  | 1 |  |  | 0.10 |  |  | 0.10 |  | 0.10 |
| A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  |  | 0.00 |  |  | 0.00 |  | 0.00 |
| C1-4 | Influence of the calibration antenna feed cable: | 0.02 | 0.02 | 0.02 | U-shaped | 1.41 |  | 1 |  |  | 0.02 |  |  | 0.02 |  | 0.02 |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 |  | 1 |  |  | 0.29 |  |  | 0.25 |  | 0.25 |
| A2-1b | Misalignment positioning system | 0.00 | 0.00 | 0.00 | Exp. normal | 2.00 |  | 1 |  |  | 0.00 |  |  | 0.00 |  | 0.00 |
| A2-9 | Misalignment of calibration antenna and test range antenna | 0.50 | 0.50 | 0.50 | Exp. normal | 2.00 |  | 1 |  |  | 0.25 |  |  | 0.25 |  | 0.25 |
| A2-2b | Rotary Joints | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 |  | 1 |  |  | 0.03 |  |  | 0.03 |  | 0.03 |
| A2-4b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.41 |  | 1 |  |  | 0.06 |  |  | 0.06 |  | 0.06 |
| A2-11 | QZ ripple calibration antenna | 0.01 | 0.01 | 0.01 | Gaussian | 1.00 |  | 1 |  |  | 0.01 |  |  | 0.01 |  | 0.01 |
| A2-13 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.73 |  | 1 |  |  | 0.15 |  |  | 0.15 |  | 0.15 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 1.28 |  | 1.32 |  |  | 1.30 |  |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 2.51 |  | 2.58 |  |  | 2.55 |  |  |

#### 9.3.3.4	MU value derivation, FR2

Table 9.3.3.4-1: CATR MU value derivation for EIRP accuracy measurements in Extreme test conditions, FR2

| UID | Uncertainty source | Uncertainty value |  |  |  | Distribution of the probability |  |  | Divisor based on distribution shape |  |  |  | ci |  |  |  | Standard uncertainty ui [dB] |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  |  |  |  |  | 24.25<f ≤29.5GHz |  |  |  | 37<f ≤40GHz |  |  |  |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A2-1a | Misalignment DUT & pointing error (EIRP) | 0.20 |  | 0.20 |  |  | Exp. normal |  |  | 2.00 |  |  |  | 1 |  |  |  | 0.10 |  |  |  | 0.10 |  |  |
| C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) -High power | 0.50 |  | 0.70 |  |  | Gaussian |  |  | 1.00 |  |  |  | 1 |  |  |  | 0.50 |  |  |  | 0.70 |  |  |
| A2-2a | Standing wave between DUT and test range antenna | 0.03 |  | 0.03 |  |  | U-shaped |  |  | 1.41 |  |  |  | 1 |  |  |  | 0.02 |  |  |  | 0.02 |  |  |
| A2-3 | RF leakage, test range antenna cable connector terminated. | 0.01 |  | 0.01 |  |  | Gaussian |  |  | 1.00 |  |  |  | 1 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |
| A2-13 | QZ ripple with DUT (Extreme) | 0.70 |  | 0.70 |  |  | Gaussian |  |  | 1.00 |  |  |  | 1 |  |  |  | 0.70 |  |  |  | 0.70 |  |  |
| A2-12 | Frequency flatness | 0.25 |  | 0.25 |  |  | Gaussian |  |  | 1.00 |  |  |  | 1 |  |  |  | 0.25 |  |  |  | 0.25 |  |  |
| A2-15 | radome loss variation | 0.50 |  | 0.50 |  |  | Gaussian |  |  | 1.00 |  |  |  | 1 |  |  |  | 0.50 |  |  |  | 0.50 |  |  |
| A2-14 | wet radome loss variation | 0.90 |  | 0.90 |  |  | Gaussian |  |  | 1.00 |  |  |  | 1 |  |  |  | 0.90 |  |  |  | 0.90 |  |  |
| A2-16 | Change in absorber behavior | 0.50 |  | 0.50 |  |  | Gaussian |  |  | 1.00 |  |  |  | 1 |  |  |  | 0.50 |  |  |  | 0.50 |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Network Analyzer | 0.30 |  | 0.30 |  |  | Gaussian |  |  | 1.00 |  |  |  | 1 |  |  |  | 0.30 |  |  |  | 0.30 |  |  |
| A2-5a | Mismatch of receiver chain | 0.43 |  | 0.57 |  |  | U-shaped |  |  | 1.41 |  |  |  | 1 |  |  |  | 0.30 |  |  |  | 0.40 |  |  |
| A2-6 | Insertion loss variation in receiver chain | 0.00 |  | 0.00 |  |  | Rectangular |  |  | 1.73 |  |  |  | 1 |  |  |  | 0.00 |  |  |  | 0.00 |  |  |
| A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | 0.01 |  |  | Gaussian |  |  | 1.00 |  |  |  | 1 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |
| A2-7 | Influence of the calibration antenna feed cable | 0.21 |  | 0.29 |  |  | U-shaped |  |  | 1.41 |  |  |  | 1 |  |  |  | 0.15 |  |  |  | 0.21 |  |  |
| C1-4 | SGH Calibration uncertainty | 0.52 |  | 0.52 |  |  | Rectangular |  |  | 1.73 |  |  |  | 1 |  |  |  | 0.30 |  |  |  | 0.30 |  |  |
| A2-8 | Misalignment positioning system | 0.00 |  | 0.00 |  |  | Exp. normal |  |  | 2.00 |  |  |  | 1 |  |  |  | 0.00 |  |  |  | 0.00 |  |  |
| A2-1b | Misalignment of calibration antenna and test range antenna | 0.00 |  | 0.00 |  |  | Exp. normal |  |  | 2.00 |  |  |  | 1 |  |  |  | 0.00 |  |  |  | 0.00 |  |  |
| A2-9 | Rotary joints | 0.00 |  | 0.00 |  |  | U-shaped |  |  | 1.41 |  |  |  | 1 |  |  |  | 0.00 |  |  |  | 0.00 |  |  |
| A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 |  | 0.09 |  |  | U-shaped |  |  | 1.41 |  |  |  | 1 |  |  |  | 0.06 |  |  |  | 0.06 |  |  |
| A2-4b | QZ ripple calibration antenna | 0.01 |  | 0.01 |  |  | Gaussian |  |  | 1.00 |  |  |  | 1 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |
| A2-11 | Switching uncertainty | 0.10 |  | 0.10 |  |  | Rectangular |  |  | 1.73 |  |  |  | 1 |  |  |  | 0.06 |  |  |  | 0.06 |  |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 1.56 |  |  |  | 1.66 |  |  |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 3.05 |  |  |  | 3.25 |  |  |  |

.

### 9.3.4		Maximum accepted test system uncertainty

For the frequency range up to 4.2 GHz, the same MU values as for E-UTRA in TS 37.145-2 [4] were adopted. It is expected that the test chamber setup, calibration and measurement procedures for E-UTRA and NR will be highly similar. All uncertainty factors were judged to be the same.

For the frequency range 4.2 - 6 GHz, all MU factors, including instrumentation related MU were judged to be the same as for the 3 - 4.2 GHz range, and thus the total MU for 4.2 – 6 GHz is the same as for 3 - 4.2 GHz. This assessment was made under the assumption of testing BS designed for licensed spectrum; for unlicensed spectrum the MU may differ.

The MU value was thus agreed to be 2.5 dB for up to 3 GHz bands and 2.6 dB for 3 – 6 GHz bands. The MU in 4.2 - 6 GHz is valid for BS designed to operate in licensed spectrum.

For FR2, for the direct far field method the MU budget is very similar to the existing MU budget for the EIRP accuracy requirement. However there are a number of additional sources of uncertainty due to the environmental enclosure that need to be added to the budget. For FR2 only the CATR MU budget has been assessed however other suitable camber types are not precluded.

Based on the above evaluation, the MU was decided to be 3.1 dB for the frequency range 24.25 < f < 29.5 GHz and 3.3 dB for the frequency range 37 < f < 40 GHz.

Test system specific measurement uncertainty values for the EIRP accuracy in Extreme test conditions are summarised in table 9.3.4-1.

Table 9.3.4-1: Test system specific measurement uncertainty values for the EIRP accuracy in Extreme test conditions, FR1

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber | 2.46 | 2.53 | 2.53 |
| Compact Antenna Test Range | 2.51 | 2.58 | 2.55 |
| Common maximum accepted test system uncertainty | 2.5 | 2.6 | 2.6 |

Table 9.3.4-2: Test system specific measurement uncertainty values for the EIRP accuracy in Extreme test conditions, FR2

|  | Expanded uncertainty ue (dB) |  |
| --- | --- | --- |
|  | 24.25<f <29.5GHz | 37GHz < f ≦ 40 GHz |
| Compact Antenna Test Range | 3.05 | 3.25 |
| Common maximum accepted test system uncertainty | 3.1 | 3.3 |

An overview of the MU values for all the requirements is captured in clause 16.

--- Next change ---

#### 9.4.2.3	MU value derivation, FR1

As the DL RS power is an absolute measurement most of the uncertainties form the EIRP accuracy remain the same. Also it can be noted that the measured signal is a wanted signal and hence will be beam formed in the same way as the wanted signal, hence any errors which may be dependent on the beam shape will be the same as for the EIRP accuracy measurement.

Table 9.4.2.3-1: Indoor Anechoic Chamber measurement uncertainty value derivation for OTA E-UTRA DL RS power measurement

| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape |  | ci |  | Standard uncertainty ui [dB] |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  |  |  | f≤3 GHz |  |  | 3<f≤4.2 GHz |  |  | 4.2<f≤6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 |  | 1 |  | 0.02 |  |  | 0.02 |  |  | 0.02 |
| A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.30 | 0.30 | 0.30 | Rectangular | 1.73 |  | 1 |  | 0.17 |  |  | 0.17 |  |  | 0.17 |
| A1-3 | Quality of quiet zone | 0.10 | 0.10 | 0.10 | Gaussian | 1.00 |  | 1 |  | 0.10 |  |  | 0.10 |  |  | 0.10 |
| A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 |  | 1 |  | 0.01 |  |  | 0.01 |  |  | 0.01 |
| A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 |  | 1 |  | 0.00 |  |  | 0.00 |  |  | 0.00 |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1.00 |  | 1 |  | 0.05 |  |  | 0.05 |  |  | 0.05 |
| C3-1 | DL-RS MU derived from conducted spec | 0.41 | 0.56 | 0.56 | Gaussian | 1.00 |  | 1 |  | 0.41 |  |  | 0.56 |  |  | 0.56 |
| A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.41 |  | 1 |  | 0.10 |  |  | 0.23 |  |  | 0.23 |
| A1-8 | Random uncertainty | 0.10 | 0.10 | 0.10 | Rectangular | 1.73 |  | 1 |  | 0.06 |  |  | 0.06 |  |  | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 |  | 1 |  | 0.04 |  |  | 0.04 |  |  | 0.04 |
| A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 |  | 1 |  | 0.01 |  |  | 0.01 |  |  | 0.01 |
| A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 |  | 1 |  | 0.04 |  |  | 0.04 |  |  | 0.04 |
| A1-3 | Quality of quiet zone | 0.10 | 0.10 | 0.10 | Gaussian | 1.00 |  | 1 |  | 0.10 |  |  | 0.10 |  |  | 0.10 |
| A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 |  | 1 |  | 0.01 |  |  | 0.01 |  |  | 0.01 |
| A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 |  | 1 |  | 0.00 |  |  | 0.00 |  |  | 0.00 |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1.00 |  | 1 |  | 0.05 |  |  | 0.05 |  |  | 0.05 |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.20 | 0.20 | Gaussian | 1.00 |  | 1 |  | 0.13 |  |  | 0.20 |  |  | 0.20 |
| A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 |  | 1 |  | 0.03 |  |  | 0.03 |  |  | 0.03 |
| A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1.00 |  | 1 |  | 0.06 |  |  | 0.06 |  |  | 0.06 |
| A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 |  | 1 |  | 0.03 |  |  | 0.03 |  |  | 0.03 |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 |  | 1 |  | 0.29 |  |  | 0.25 |  |  | 0.25 |
| A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 |  | 1 |  | 0.00 |  |  | 0.00 |  |  | 0.00 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.59 |  | 0.73 |  |  | 0.73 |  |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.15 |  | 1.44 |  |  | 1.44 |  |  |

--- Next change ---

#### 9.4.3.3	MU value derivation, FR1

As the DL RS power is an absolute measurement most of the uncertainties form the EIRP accuracy remain the same. Also it can be noted that the measured signal is a wanted signal and hence will be beam formed in the same way as the wanted signal, hence any errors which may be dependent on the beam shape will be the same as for the EIRP accuracy measurement.

Table 6.4.3.4-1: CATR MU value derivation for OTA E-UTRA DL RS power measurement

| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz |  | 3<f≤4.2 GHz |  | 4.2<f≤6 GHz |  |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A2-1a | Misalignment DUT & pointing error for EIRP | 0.00 | 0.00 | 0.00 | Exp. normal | 2.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| C3-1 | DL-RS MU derived from conducted spec | 0.41 | 0.56 | 0.56 | Gaussian | 1.00 | 1 | 0.41 |  | 0.56 |  | 0.56 |  |
| A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.41 | 1 | 0.15 |  | 0.15 |  | 0.15 |  |
| A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A2-4a | QZ ripple DUT | 0.09 | 0.09 | 0.09 | Gaussian | 1.00 | 1 | 0.09 |  | 0.09 |  | 0.09 |  |
| A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1.00 | 1 | 0.25 |  | 0.25 |  | 0.25 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.20 | 0.20 | Gaussian | 1.00 | 1 | 0.13 |  | 0.20 |  | 0.20 |  |
| A2-5 | Mismatch of receiver chain | 0.13 | 0.33 | 0.33 | U-shaped | 1.41 | 1 | 0.09 |  | 0.23 |  | 0.23 |  |
| A2-6 | Insertion loss variation of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.73 | 1 | 0.10 |  | 0.10 |  | 0.10 |  |
| A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A2-7 | Influence of the calibration antenna feed cable: | 0.02 | 0.02 | 0.02 | U-shaped | 1.41 | 1 | 0.02 |  | 0.02 |  | 0.02 |  |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.29 |  | 0.25 |  | 0.25 |  |
| A2-8 | Misalignment positioning system | 0.00 | 0.00 | 0.00 | Exp. normal | 2.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A2-1b | Misalignment of calibration antenna and test range antenna | 0.50 | 0.50 | 0.50 | Exp. normal | 2.00 | 1 | 0.25 |  | 0.25 |  | 0.25 |  |
| A2-9 | Rotary Joints | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.03 |  | 0.03 |  | 0.03 |  |
| A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.41 | 1 | 0.06 |  | 0.06 |  | 0.06 |  |
| A2-4b | QZ ripple calibration antenna | 0.01 | 0.01 | 0.01 | Gaussian | 1.00 | 1 | 0.01 |  | 0.01 |  | 0.01 |  |
| A2-11 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.73 | 1 | 0.15 |  | 0.15 |  | 0.15 |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  |  | 0.69 |  | 0.81 |  | 0.81 |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  | 1.35 |  | 1.60 |  | 1.60 |  |

--- Next change ---

#### 9.4.4.3	MU value derivation, FR1

Table 9.4.4.3-1: NFTR measurement uncertainty value derivation for OTA E-UTRA DL RS power measurement

| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz |  | 3<f≤4.2 GHz |  | 4.2<f≤6 GHz |  |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A3-1 | Axes Intersection | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-2 | Axes Orthogonality | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-3 | Horizontal Pointing | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-4 | Probe Vertical Position | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-5 | Probe H/V pointing | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-6 | Measurement Distance | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-7 | Amplitude and Phase Drift | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-8 | Amplitude and Phase Noise | 0.02 | 0.02 | 0.02 | Gaussian | 1.00 | 1 | 0.02 |  | 0.02 |  | 0.02 |  |
| A3-9 | Leakage and Crosstalk | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-10 | Amplitude Non-Linearity | 0.04 | 0.04 | 0.04 | Gaussian | 1.00 | 1 | 0.04 |  | 0.04 |  | 0.04 |  |
| A3-11 | Amplitude and Phase Shift in rotary joints | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-12 | Channel Balance Amplitude and Phase | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-13 | Probe Polarization Amplitude and Phase | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-14 | Probe Pattern Knowledge | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-15 | Multiple Reflections | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-16 | Room Scattering | 0.09 | 0.09 | 0.09 | Gaussian | 1.00 | 1 | 0.09 |  | 0.09 |  | 0.09 |  |
| A3-17 | DUT support Scattering | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-18 | Scan Area Truncation | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-19 | Sampling Point Offset | 0.01 | 0.01 | 0.01 | Gaussian | 1.00 | 1 | 0.01 |  | 0.01 |  | 0.01 |  |
| A3-20 | Spherical Mode Truncation | 0.02 | 0.02 | 0.02 | Gaussian | 1.00 | 1 | 0.02 |  | 0.02 |  | 0.02 |  |
| A3-21 | Positioning | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.02 |  | 0.02 |  | 0.02 |  |
| A3-22 | Probe Array Uniformity | 0.06 | 0.06 | 0.06 | Gaussian | 1.00 | 1 | 0.06 |  | 0.06 |  | 0.06 |  |
| A3-23 | Mismatch of receiver chain | 0.28 | 0.28 | 0.28 | U-Shaped | 1.41 | 1 | 0.20 |  | 0.20 |  | 0.20 |  |
| A3-24 | Insertion loss of receiver chain | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-25 | Uncertainty of the absolute gain of the probe antenna | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| C3-1 | DL-RS MU derived from conducted spec | 0.41 | 0.56 | 0.56 | Gaussian | 1.00 | 1 | 0.41 |  | 0.56 |  | 0.56 |  |
| A3-26 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | 0.15 | Gaussian | 1.00 | 1 | 0.15 |  | 0.15 |  | 0.15 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.20 | 0.20 | Gaussian | 1.00 | 1 | 0.13 |  | 0.20 |  | 0.20 |  |
| A3-27 | Mismatch of receiver chain | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-28 | Insertion loss of receiver chain | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-29 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | 0.02 | U-Shaped | 1.41 | 1 | 0.01 |  | 0.01 |  | 0.01 |  |
| A3-30 | Influence of the calibration antenna feed cable | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| A3-31 | Influence of the probe antenna cable | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 |  | 0.00 |  | 0.00 |  |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.29 |  | 0.25 |  | 0.25 |  |
| A3-32 | Short term repeatability | 0.09 | 0.09 | 0.09 | Gaussian | 1.00 | 1 | 0.09 |  | 0.09 |  | 0.09 |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  |  | 0.59 |  | 0.71 |  | 0.71 |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  | 1.17 |  | 1.39 |  | 1.39 |  |

.

--- Next change ---

### 9.4.5		Maximum accepted test system uncertainty

Maximum test system uncertainties derivation methodology was described in subclause 5.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for the OTA E-UTRA DL RS power test can be derived from values captured in table 6.4.5-1, separately for each of the defined frequency ranges. The common maximum values are applicable for all test methods addressing OTA E-UTRA DL RS power test requirement. Based on the input values, the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values were derived.

Table 9.4.5-1: Test system specific measurement uncertainty values for the OTA E-UTRA DL RS power test

|  | Expanded uncertainty ue [dB] |  |  |
| --- | --- | --- | --- |
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |
| Indoor Anechoic Chamber | 1.15 | 1.44 | 1.44 |
| Compact Antenna Test Range | 1.35 | 1.60 | 1.60 |
| One Dimensional Compact Range Chamber | 1.17 | 1.39 | 1.39 |
| PWS | [1.24] | [1.53] | [1.53] |
| Common maximum accepted test system uncertainty | 1.3 | 1.5 | 1.5 |

### 9.4.6		Test Tolerance for OTA E-UTRA DL RS power

The TT was decided to be the same as the MU in FR1.

Table 9.4.6-1: Test Tolerance values for the OTA E-UTRA DL RS power

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 1.3 | 1.5 | 1.5 |

--- Next change ---

#### 9.5.2.3	MU value derivation, FR1

As the output power dynamics are relative measurements most of the uncertainties form the EIRP accuracy cancel out as the same error will be applied to both of the measured OTA signals.

This includes all calibration errors, misalignment errors, impedance mismatch and mutual coupling.

As the both the measured OTA signal will have the same beam pattern quiet zone errors, phase curvature errors also can be expected to be the same for both signals.

The uncertainty budget descriptions are the same as those in table 6.4.2.3-1 with the addition descriptions in table 6.5.2.3-1.

The MU uncertainty assessment is shown in table 6.5.2.4-1, zero values have been omitted in the table for the sake of space, but still be considered as part of the budget.

Table 9.5.2.3-1: Indoor Anechoic Chamber measurement uncertainty value derivation for OTA total power dynamic range measurement

| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape |  | ci |  | Standard uncertainty ui [dB] |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  |  |  | f≤3 GHz |  | 3<f≤4.2 GHz |  | 4.2<f≤6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| C3-2 | Total power dynamic range conducted uncertainty | 0.20 | 0.20 | 0.20 | Gaussian | 1.00 |  | 1 |  | 0.20 |  | 0.20 |  | 0.20 |
| A1-8 | Random uncertainty | 0.10 | 0.10 | 0.10 | Rectangular | 1.73 |  | 1 |  | 0.06 |  | 0.06 |  | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  | 0.21 | 0.21 | 0.21 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | 0.41 | 0.41 | 0.41 |

The same uncertainty assessments have been carried out for the UTRA dynamic range requirements, i.e.

- Power control steps,

- Power control dynamic range,

- Total power dynamic range,

- IPDL Time mask.

In each case the uncertainty for the conducted measurement is the same as that for the conducted MU in TS 25.141 [2] as follows:

- Power control steps, Uncertainty of conducted measurement = 0.1 dB, Expanded OTA uncertainty = 0.15 dB.

- Power control dynamic range, Uncertainty of conducted measurement = 1.1 dB, Expanded OTA uncertainty = 1.11 dB.

- Total power dynamic range, Uncertainty of conducted measurement = 0.3 dB, Expanded OTA uncertainty = 0.32 dB.

- IPDL Time mask, Uncertainty of conducted measurement = 0.7dB, Expanded OTA uncertainty = 0.71 dB.

--- Next change ---

#### 9.5.3.3	MU value derivation, FR1

As the output power dynamics are relative measurements most of the uncertainties form the EIRP accuracy cancel out as the same error will be applied to both of the measured OTA signals.

This includes all calibration errors, misalignment errors, impedance mismatch and mutual coupling.

As the both the measured OTA signal will have the same beam pattern quiet zone errors, phase curvature errors also can be expected to be the same for both signals.

The uncertainty budget descriptions are the same as those in table 6.4.3.3-1 with the addition descriptions in table 6.5.3.3-1.

The MU uncertainty assessment is shown in table 6.5.3.4-1, zero values have been omitted in the table for the sake of space, but still be considered as part of the budget.

Table 9.5.3.3-1: CATR MU value derivation for OTA total power dynamic range measurement

| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz |  | 4.2<f≤6 GHz |  |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| C3-2 | Total power dynamic range conducted uncertainty | 0.20 | 0.20 | 0.20 | Gaussian | 1.00 | 1 | 0.20 | 0.20 |  | 0.20 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  |  |  | 0.20 |  | 0.20 |  | 0.20 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  |  | 0.39 |  | 0.39 |  | 0.39 |

The same uncertainty assessments have been carried out for the UTRA dynamic range requirements, i.e.

- Power control steps,

- Power control dynamic range,

- Total power dynamic range,

- IPDL Time mask.

In each case the uncertainty for the conducted measurement is the same as that for the conducted MU in TS 25.141 [2] as follows:

- Power control steps: Uncertainty of conducted measurement = 0.1 dB, Expanded OTA uncertainty = 0.1 dB.

- Power control dynamic range: Uncertainty of conducted measurement = 1.1 dB, Expanded OTA uncertainty = 1.1 dB.

- Total power dynamic range: Uncertainty of conducted measurement = 0.3 dB, Expanded OTA uncertainty = 0.3 dB.

- IPDL Time mask: Uncertainty of conducted measurement = 0.7 dB, Expanded OTA uncertainty = 0.7 dB.

--- Next change ---

#### 9.5.4.3	MU value derivation, FR1

As the output power dynamics are relative measurements most of the uncertainties form the EIRP accuracy cancel out as the same error will be applied to both of the measured OTA signals.

The uncertainty budget descriptions are the same as those in table 6.4.4.4-1 (excluding uncertainties from the NF to FF transformation, since the transformation is not needed) with the addition descriptions in table 6.5.4.3-1:

The MU uncertainty assessment is shown in table 6.5.4.4-1, zero values have been omitted in the table for the sake of space, but still be considered as part of the budget.

Table 9.5.4.3-1: NFTR MU value derivation for OTA E-UTRA total power dynamic range measurement

| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| C3-2 | Total power dynamic range conducted uncertainty | 0.20 | 0.20 | 0.20 | Gaussian | 1.00 | 1 | 0.20 | 0.20 | 0.20 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  | 0.20 | 0.20 | 0.20 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | 0.39 | 0.39 | 0.39 |

--- Next change ---

### 9.5.5		Maximum accepted test system uncertainty

Maximum test system uncertainties derivation methodology was described in subclause 5.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for the OTA output power dynamics tests can be derived from values captured in tables 6.5.5-1 to 6.5.5-5, separately for each of the defined frequency ranges. The common maximum values are applicable for all test methods addressing certain OTA output power dynamics test requirement. Based on the input values, the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values were derived.

The output power dynamic range MU for FR1 for up to 4.2 GHz was agreed to be the same as for eAAS WI. It is expected that the measurement chamber setup, calibration and measurement procedures and the MU budget will be identical for E-UTRA and NR.

Furthermore, for the frequency range 4.2 – 6 GHz, the chamber and instrument uncertainties are the same as those for the frequency range 3 – 4.2 GHz, assuming testing of a BS designed for operation in licensed spectrum. The MU value was agreed to be +/- 0.4 dB for all bands up to 6 GHz.

For FR2: Similarly to FR1, since the OTA output power dynamic range requirement is a relative measurement, only the test equipment uncertainty is of importance. Based on expected test equipment uncertainty, the output power dynamic range MU for FR2 was decided to be +/- 0.4 dB (same as for FR1).

Table 9.5.5-1: Test system specific measurement uncertainty values for the OTA E-UTRA and NR dynamic range test

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber | 0.41 | 0.41 | 0.41 |
| Compact Antenna Test Range | 0.39 | 0.39 | 0.39 |
| Near Field Test Range | 0.39 | 0.39 | 0.39 |
| PWS | 0.39 | 0.39 | 0.39 |
| Common maximum accepted test system uncertainty | 0.4 | 0.4 | 0.4 |

Table 9.5.5-2: Test system specific measurement uncertainty values for the UTRA power control steps test

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber | 0.1 | 0.1 | 0.1 |
| Compact Antenna Test Range | 0.1 | 0.1 | 0.1 |
| Near Field Test Range | 0.1 | 0.1 | 0.1 |
| PWS | 0.1 | 0.1 | 0.1 |
| Common maximum accepted test system uncertainty | 0.1 | 0.1 | 0.1 |

Table 9.5.5-3: Test system specific measurement uncertainty values for the UTRA power control dynamic range test

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber | 1.1 | 1.1 | 1.1 |
| Compact Antenna Test Range | 1.1 | 1.1 | 1.1 |
| Near Field Test Range | 1.1 | 1.1 | 1.1 |
| PWS | 1.1 | 1.1 | 1.1 |
| Common maximum accepted test system uncertainty | 1.1 | 1.1 | 1.1 |

Table 9.5.5-4: Test system specific measurement uncertainty values for the UTRA total power dynamic range test

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber | 0.3 | 0.3 | 0.3 |
| Compact Antenna Test Range | 0.3 | 0.3 | 0.3 |
| Near Field Test Range | 0.3 | 0.3 | 0.3 |
| PWS | 0.3 | 0.3 | 0.3 |
| Common maximum accepted test system uncertainty | 0.3 | 0.3 | 0.3 |

Table 9.5.5-5: Test system specific measurement uncertainty values for the UTRA IPDL time mask test

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber | 0.7 | 0.7 | 0.7 |
| Compact Antenna Test Range | 0.7 | 0.7 | 0.7 |
| Near Field Test Range | 0.7 | 0.7 | 0.7 |
| PWS | 0.7 | 0.7 | 0.7 |
| Common maximum accepted test system uncertainty | 0.7 | 0.7 | 0.7 |

An overview of the MU values for all the requirements is captured in clause 16.

--- Next change ---

#### 9.7.2.3	MU value derivation, FR1

As both the wanted signal and the noise signal are at the same frequency they will be measured at the same time the requirement is effectively differential and most of the OTA chamber errors will cancel out.

The wanted signal will be beam formed and hence the errors used for the EIRP accuracy will be valid, however the co-channel noise may not be beam formed and hence could suffer different errors due to the chamber quite zone, and phase profile. These items are included in both the calibration error and the measurement error, as the requirement is differential if there is a difference between the wanted and the unwanted it will only be due to the measurement phase. The calibration errors will cancel as calibration is only done one so they will be the same for both wanted and unwanted signals.

Potentially, the EVM may vary in space due to different patterns of wanted signal and distortion. Thus for narrow beams, it may be possible that beam pointing and alignment errors could impact EVM results.

As EVM is also dependent on the phase of the calibrated path it is possible that phase ripple in the quite zone or elsewhere, which arises due to multipath reflections, may lead to frequency ripple and cause additional EVM errors which do not appear in a power accuracy analysis as done for EIRP accuracy.

The potential impacts of both beam pointing misalignment and scattering within the chamber on the received waveform and measurement accuracy were investigated. The potential deviation in the measured EVM arising from beam pointing errors was examined considering a worst case scenario, in which variation in space of EVM is maximal due to the ideal signal being correlated and the distortion uncorrelated; hence the impact of misalignment error would be the difference between array gain and element gain. Even in this circumstance, alignment errors of several degrees would not lead to a significant error in the measured EVM. Considering all likely chamber sizes, for E-UTRA any scattering would fall within the cyclic prefix of the OFDM symbol and hence not cause ISI. Furthermore, the likely delay spread of any scattering would relate to coherence bandwidths much larger than any UTRA/E-UTRA channel bandwidth. Even if the scattered energy would cause interference, the interference level would anyhow not lead to a significant EVM increase. Thus it was concluded that the impact scattering within the measurement chamber would be negligible.

The uncertainty causing by power variations when measuring OTA EVM is indicated in table 6.7.2.4-1:

Table 9.7.2.3-1: IAC MU value derivation for power uncertainty aspects of OTA EVM, FR1

| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.10 | 0.10 | 0.10 |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 |
| A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.73 | 1 | 0.06 | 0.06 | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  | 0.13 | 0.13 | 0.13 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | 0.25 | 0.25 | 0.25 |

The indoor anechoic chamber budget is carried out without consideration of the measurement equipment as this MU is given in %, converting to dB gives, for example:

2% is equivalent to 20*log10(2/100) = -33.98 dB

If the unwanted signal is 0.25 dB higher than the wanted due to the test system then this will be degraded to -33.73 dB, and

-33.73 dB is equivalent to:	10(-33.73/20) *100 = 2.06%

Additional error due to potential phase error has not been considered however the potential increase due to then OTA test equipment is well within the contribution allowable with a 1% linear MU.

NOTE:	Analysis of the phase uncertainties indicates that the contributions are not significant to affect the final MU value, however if future work indicates that phase or any other errors not related to amplitude calibration may affect the EVM measurement uncertainty the MU analysis may be re-examined.

--- Next change ---

#### 9.7.3.3 	MU value derivation, FR1

As both the wanted signal and the noise signal are at the same frequency they will be measured at the same time the requirement is effectively differential and most of the OTA chamber errors will cancel out.

The wanted signal will be beam formed and hence the errors used for the EIRP accuracy will be valid, however the co-channel noise may not be beam formed and hence could suffer different errors due to the chamber quite zone, and phase profile. These items are included in both the calibration error and the measurement error, as the requirement is differential if there is a difference between the wanted and the unwanted it will only be due to the measurement phase. The calibration errors will cancel as calibration is only done one so they will be the same for both wanted and unwanted signals.

Potentially, the EVM may vary in space due to different patterns of wanted signal and distortion. Thus for narrow beams, it may be possible that beam pointing and alignment errors could impact EVM results.

As EVM is also dependent on the phase of the calibrated path it is possible that phase ripple in the quite zone or elsewhere, which arises due to multipath reflections, may lead to frequency ripple and cause additional EVM errors which do not appear in a power accuracy analysis as done for EIRP accuracy.

The potential impacts of both beam pointing misalignment and scattering within the chamber on the received waveform and measurement accuracy were investigated. The potential deviation in the measured EVM arising from beam pointing errors was examined considering a worst case scenario, in which variation in space of EVM is maximal due to the ideal signal being correlated and the distortion uncorrelated; hence the impact of misalignment error would be the difference between array gain and element gain. Even in this circumstance, alignment errors of several degrees would not lead to a significant error in the measured EVM. Considering all likely chamber sizes, for E-UTRA any scattering would fall within the cyclic prefix of the OFDM symbol and hence not cause ISI. Furthermore, the likely delay spread of any scattering would relate to coherence bandwidths much larger than any UTRA/E-UTRA channel bandwidth. Even if the scattered energy would cause interference, the interference level would anyhow not lead to a significant EVM increase. Thus it was concluded that the impact scattering within the measurement chamber would be negligible.

The uncertainty causing by power variations when measuring OTA EVM is indicated in table 6.7.3.4-1:

Table 9.7.3.3-1: CATR MU value derivation for power uncertainty aspects of OTA EVM, FR1

| UID |  | Uncertainty source |  | Uncertainty value |  |  |  |  |  | Distribution of the probability |  | Divisor based on distribution shape |  | ci |  | Standard uncertainty ui [dB] |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | f<3 GHz |  | 3<f<4.2 GHz |  | 4.2<f<6 GHz |  |  |  |  |  |  |  | f<3 GHz |  | 3<f<4.2 GHz |  | 4.2<f<6 GHz |  |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.41 | 1 | 0.15 | 0.15 | 0.15 |
| A2-4a | QZ ripple DUT | 0.09 | 0.09 | 0.09 | Normal | 1 | 1 | 0.09 | 0.09 | 0.09 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  | 0.18 | 0.18 | 0.18 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | 0.34 | 0.34 | 0.34 |

The CATR budget is carried out without consideration of the measurement equipment as this MU is given in %, converting to dB gives, for example:

2% is equivalent to 20*log10(2/100) = -33.98 dB

If the unwanted signal is 0.35 dB higher than the wanted due to the test system then this will be degraded to -33.63 dB, and

-33.63 dB is equivalent to:	10(-33.63/20) *100 = 2.08%

Additional error due to potential phase error has not been considered however the potential increase due to then OTA test equipment is well within the contribution allowable with a 1% linear MU.

NOTE:	Analysis of the phase uncertainties indicates that the contributions are not significant to affect the final MU value, however if future work indicates that phase or any other errors not related to amplitude calibration may affect the EVM measurement uncertainty the MU analysis may be re-examined.

--- Next change ---

#### 9.10.2.3	MU value derivation, FR2

The MU assessment was carried out using a CATR chamber only. However other chamber types are not precluded if suitable MU assessment is done.

The CATR test setup and calibration and measurement procedures for FR2 are expected to be similar to those of FR1, although the test chamber dimensions and associated MU values will scale due to the shorter wavelengths and larger relative array apertures. However, it is noted that in order to achieve the test instrument uncertainties that were assumed, calibration of the spectrum analyzer may be needed.

Table 9.10.2.3-1: Compact antenna test range uncertainty assessment for EIRP measurements for transmitter OFF power and transmitter transient period

| UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape |  | ci |  | Standard uncertainty ui [dB] |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  |  |  | 24.25<f ≤29.5GHz |  | 37<f ≤40GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |
| A2-1a | Misalignment DUT & pointing error (EIRP) | 0.20 | 0.20 | Exp. normal | 2.00 |  | 1 |  | 0.10 |  | 0.10 |
| C1-10 | RF power measurement equipment standard uncertainty σ (dB) of the absolute level for a time domain wideband measurement for FR2 | 1.25 | 1.45 | Gaussian | 1.00 |  | 1 |  | 1.25 |  | 1.45 |
| A2-2a | Standing wave between DUT and test range antenna | 0.03 | 0.03 | U-shaped | 1.41 |  | 1 |  | 0.02 |  | 0.02 |
| A2-3 | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | Gaussian | 1.00 |  | 1 |  | 0.01 |  | 0.01 |
| A2-4a | QZ ripple with DUT | 0.40 | 0.40 | Gaussian | 1.00 |  | 1 |  | 0.40 |  | 0.40 |
| A2-12 | Frequency flatness | 0.25 | 0.25 | Gaussian | 1.00 |  | 1 |  | 0.25 |  | 0.25 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Network Analyzer | 0.30 | 0.30 | Gaussian | 1.00 |  | 1 |  | 0.30 |  | 0.30 |
| A2-5b | Mismatch of receiver chain for low power | 0.72 | 0.72 | U-shaped | 1.41 |  | 1 |  | 0.51 |  | 0.51 |
| A2-6 | Insertion loss variation in receiver chain | 0.00 | 0.00 | Rectangular | 1.73 |  | 1 |  | 0.00 |  | 0.00 |
| A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Gaussian | 1.00 |  | 1 |  | 0.01 |  | 0.01 |
| A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | 1.41 |  | 1 |  | 0.15 |  | 0.21 |
| C1-4 | SGH Calibration uncertainty | 0.52 | 0.52 | Rectangular | 1.73 |  | 1 |  | 0.30 |  | 0.30 |
| A2-8 | Misalignment positioning system | 0.00 | 0.00 | Exp. normal | 2.00 |  | 1 |  | 0.00 |  | 0.00 |
| A2-1b | Misalignment of calibration antenna and test range antenna | 0.00 | 0.00 | Exp. normal | 2.00 |  | 1 |  | 0.00 |  | 0.00 |
| A2-9 | Rotary joints | 0.00 | 0.00 | U-shaped | 1.41 |  | 1 |  | 0.00 |  | 0.00 |
| A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | U-shaped | 1.41 |  | 1 |  | 0.06 |  | 0.06 |
| A2-4b | QZ ripple calibration antenna | 0.01 | 0.01 | Gaussian | 1.00 |  | 1 |  | 0.01 |  | 0.01 |
| A2-11 | Switching uncertainty | 0.10 | 0.10 | Rectangular | 1.73 |  | 1 |  | 0.06 |  | 0.06 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 1.50 | 1.68 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 2.95 | 3.29 |

.

### 9.10.3	Maximum accepted test system uncertainty

The MU assessment was carried out using a CATR chamber only. However other chamber types are not precluded if suitable MU assessment is done.

Table 9.10.3-1: Maximum accepted test system uncertainty values for the EIRP accuracy, FR2

|  | 24.25 < f < 29.5 GHz | 37 < f < 40 GHz |
| --- | --- | --- |
| Maximum accepted test system uncertainty (dB) | 2.9 | 3.3 |

--- End of changes ---