## Summary
Review comments (RAN4#94)
| Expanded uncertainty [dB] | [11 empty] | A method exceeds agreed value
| IAC |  |  | CATR |  |  | Reverb |  |  | Agreed value
| 24.25<f ≤29.5GHz | 37<f ≤40GHz |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 4.2<f≤6 GHz
EIRP |  |  |  | 1.7368 | 2.0699 |  |  |  |  | 1.7 | 2 |  |  | x
EIRP extreme |  |  |  | 3.0523 | 3.2534 |  |  |  |  | 3.1 | 3.3
Power dynamics | [9 empty] | 0.4 | 0.4
EVM (%) | [9 empty] | 1 | 1
In-band TRP |  |  |  | 2.108 | 2.3899 |  | 1.8482 | 2.0827 |  | 2.1 | 2.4
ACLR- abs |  |  |  | 2.6922 | 2.7064 |  | 2.3594 | 2.3594 |  | 2.7 | 2.7
ACLR-rel |  |  |  | 2.2821 | 2.5352 |  | 2.1485 | 2.3594 |  | 2.3 | 2.6
OBUE |  |  |  | 2.7011 | 2.7153 |  | 2.3594 | 2.3594 |  | 2.7 | 2.7
COEX EM
COLO EM
TX IMD
| IAC |  |  | CATR |  |  | Reverb |  |  | Agreed value
| 18GHz<f≤26,5 GHz | 26,5GHz<f ≤40GHz | 40GHz<f ≤60GHz | 18GHz<f≤26,5 GHz | 26,5GHz<f ≤40GHz | 40GHz<f ≤60GHz | 18GHz<f≤26,5 GHz | 26,5GHz<f ≤40GHz | 40GHz<f ≤60GHz | 18GHz<f≤26,5 GHz | 26,5GHz<f ≤40GHz | 40GHz<f ≤60GHz
OOB EM | x | x | 4.9386 | x | x | 4.957 | x | x | 3.5282 | x | x | 5
RX EM | x | x | 4.9386 | x | x | 4.957 | x | x | 3.5282 | x | x | 5

## TE
Review comments (RAN4#94)
Test Equipment Errors
| Instrument | Uncertainty value (for the prob dist type) |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) -High power | 0.5 | 0.7 | Gaussian | 1 | 0.5 | 0.7
C1-8 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM) | 0.9 | 0.9 | Gaussian | 1 | 0.9 | 0.9
C1-9 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - relative (ACLR) | 0.75 | 0.9 | Gaussian | 1 | 0.75 | 0.9
C1-10 | RF power measurement equipment standard uncertainty σ (dB) of the absolute level for a time domain wideband measurement for FR2 | 1.25 | 1.45 | Gaussian | 1 | 1.25 | 1.45
C1-2 | RF signal generator
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 0.3 | 0.3
C1-4 | SGH Calibration uncertainty | 0.51962 | 0.51962 | Rectangular | 1.7321 | 0.3 | 0.3
Test Equipment Errors
| Instrument | Uncertainty value (for the prob dist type) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM) | 0.52 | 0.84 | 0.6 | Gaussian | 1 | 0.52 | 0.84 | 0.6 | RK: 40 to 60 MU is different for different methods: IAC=0.6 CATR=0.7 Reverb=1.15
C1-3 | Network Analyzer | 0.3 | 0.3 | 0.3 | Gaussian | 1 | 0.3 | 0.3 | 0.3
C1-4 | SGH Calibration uncertainty | 0.52 | 0.52 | 0.52 | Rectangular | 1.7321 | 0.30022 | 0.30022 | 0.30022
Conducted MU (minus missmatch) - up to 6GHz
| Instrument | Standard uncertainty σ (dB) |  | Probability distribution |  | Uncertainty value (for the prob dist type)
|  | f<3 GHz | 3<f<4.2 GHz |  |  | f<3 GHz | 3<f<4.2 GHz
23 | DL-RS MU derived from conducted spec | 0.41 | 0.56 | Gaussian | 1 | 0.41 | 0.56
23 | Total power dynamic range conducted unbcertainty | 0.2 | 0.2 | Gaussian | 1 | 0.2 | 0.2
23 | ACLR relative - MU of TE from conducted | 0.41 | 0.56 | Gaussian | 1 | 0.41 | 0.56
| ACLR abb, Conducted mnus missmatch) | 1.0108 | 1.2321 | Gaussian | 1 | 1.0108 | 1.2321
| OBUE, Conducted minus missmatch | 0.75239 | 0.85703 | Gaussian | 1 | 0.75239 | 0.85703
27 | Additional (COEX) emssions - Condcuted Uncertainty (minus missmatch) | 1.02 | 1.28 | Gaussian | 1 | 1.02 | 1.28
| TX IMD - conducted measurement uncertainty | 1 | 1.1 | Gaussian | 1 | 1 | 1.1
Conducted MU (minus missmatch) - Whole Spectrum
| Instrument | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
27 | OOB emsions - Conducted Uncertainty (minus missmatch) | 1.02 | 2.04 | Gaussian | 1 | 1.02 | 2.04
27 | Rx emisions - Conducted Uncertainty (minus missmatch) | 1 | 1.99 | Gaussian | 1 | 1 | 1.99
Co-location errors
| Co-lcoation Error source | Standard uncertainty σ (dB) |  | Probability distribution |  | Uncertainty value (for the prob dist type)
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz |  |  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz
| Uncertainty related to the selection of the CLTA (Note) | 1.5 | 1.5 | Rectangular | 1.7321 | 0.86603 | 0.86603
| Uncertainty related to the placement of the CLTA (Note) | 1.7 | 1.7 | Rectangular | 1.7321 | 0.9815 | 0.9815
| Uncertainty related to measuring close to noise floor - Emisions | 0.68 | 0.68 | Gaussian | 1 | 0.68 | 0.68
| Uncertainty related to measuring close to noise floor Tx OFF | 1 | 1 | Gaussian | 1 | 1 | 1
| Impedance mismatch between feeder cable and CLTA | 0.14 | 0.23 | U-Shaped | 1.4142 | 0.098995 | 0.16263
| Gain variations in LNA | 0.1 | 0.1 | Gaussian | 1 | 0.1 | 0.1
| Gain variations in measurement amplifier | 0.1 | 0.1 | Gaussian | 1 | 0.1 | 0.1

## CATR-Er
CATR | [10 empty] | Review comments (RAN4#94)
|  | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 40GHz<f ≤60GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: DUT measurement
A2-1a | Misalignment DUT & pointing error (EIRP) | 0.2 | 0.2 |  | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0
A2-18 | Misalignment DUT & pointing error (TRP) | 0.3 | 0.3 | 0.3 | Exp. normal | 2 | 1 | 0.15 | 0.15 | 0.15
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) -High power | 0.5 | 0.7 |  | Gaussian | 1 | 1 | 0.5 | 0.7 | 0
A2-2a | Standing wave between DUT and test range antenna | 0.03 | 0.03 | 0.21 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.14849 | RK:uncertainty changed to 0.03
A2-3 | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | 0.001 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.001
A2-4a | QZ ripple with DUT | 0.4 | 0.4 | 0.0928 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.0928 | RK: uncertainty value scorrect to match std uncertainty
A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25
A2-13 | QZ ripple with DUT (Extreme) | 0.7 | 0.7 |  | Gaussian | 1 | 1 | 0.7 | 0.7 | 0
A2-15 | radome loss variation | 0.5 | 0.5 |  | Gaussian | 1 | 1 | 0.5 | 0.5 | 0
A2-14 | wet radome loss variation | 0.9 | 0.9 |  | Gaussian | 1 | 1 | 0.9 | 0.9 | 0
A2-16 | Change in absorber behaviour | 0.5 | 0.5 |  | Gaussian | 1 | 1 | 0.5 | 0.5 | 0
C1-10 | RF power measurement equipment standard uncertainty σ (dB) of the absolute level for a time domain wideband measurement for FR2 | 1.25 | 1.45 |  | Gaussian | 1 | 1 | 1.25 | 1.45 | 0
A2-19 | Uncertainty of the LNA | x | x | 0 | Gaus | 1 | 1 | #VALUE! | #VALUE! | 0
A2-20 | Uncertainty of the Mixer | x | x | 2.25 | Gaus | 1 | 1 | #VALUE! | #VALUE! | 2.25
A2-21 | Miscellaneous uncertainty | x | x | 0 | Rectangular | 1.7321 | 1 | #VALUE! | #VALUE! | 0
Stage 1: Calibration measurement
C1-3 | Network Analyzer | 0.3 | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3
A2-5a | Mismatch of receiver chain | 0.43 | 0.57 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 |  | RK: calcualtion corrected - same unceratinty value
A2-5b | Mismatch of receiver chain for low power | 0.72 | 0.72 |  | U-shaped | 1.4142 | 1 | 0.50912 | 0.50912 |  |  | Nok: these are non-zero in 37.843. can set the same as in 37.843?
A2-6 | Insertion loss variation in receiver chain | 0 | 0 | 0.18 | Rectangular | 1.7321 | 1 | 0 | 0 |  | RK: uncertainty changed to 0
A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.001 | Gaussian | 1 | 1 | 0.01 | 0.01
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506
C1-4 | SGH Calibration uncertainty | 0.52 | 0.52 | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.30022 |  | RK: calcualtion corrected - same unceratinty value
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  | RK: uncertainty changed to 0 | Nok: these are non-zero in 37.843. can set the same as in 37.843?
A2-1b | Misalignment of calibration antenna and test range antenna | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  | RK: uncertainty changed to 0 | Nok: these are non-zero in 37.843. can set the same as in 37.843?
A2-9 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364
A2-4b | QZ ripple calibration antenna | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009
A2-11 | Switching uncertainty | 0.1 | 0.1 | 0.43 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 |  | RK: uncertainty changed to 0.1
Note on D10: Richard Kybett: This is most liely a typo, the std uncertainty is correct just the uncertainty value has FR1 value (0.928) in it.

## Reverb-Er
Reverberation Chamber | [10 empty] | Review comments (RAN4#94)
Stage 2: DUT measurement
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) -High power | 0.5 | 0.7 | Gaussian | 1 | 1 | 0.5 | 0.7
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142
A6-2 | Random uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5
Stage 2: DUT measurement
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz |  |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: DUT measurement
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM) | 0.52 | 0.84 | 0.6 | Gaussian | 1 | 1 | 0.52 | 0.84 | 0.6
A6-1 | Impedance mismatch in the receiving chain | 0.45 | 0.45 | 0.45 | U-shaped | 1.4142 | 1 | 0.3182 | 0.3182 | 0.3182
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.5 | 0.5 | 0.5 | Normal | 1 | 1 | 0.5 | 0.5 | 0.5
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.27 | 0.27 | 0.27 | Normal | 1 | 1 | 0.27 | 0.27 | 0.27
C1-3 | Network Analyzer | 0.3 | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.3
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Normal | 1 | 1 | 0.2 | 0.2 | 0.2
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Normal | 1 | 1 | 0.27 | 0.27 | 0.27
A6-7 | Uniformity of transfer function | 1.5 | 1.5 | 1.5 | Normal | 1 | 1 | 1.5 | 1.5 | 1.5
these overlap with above and are different, use above values

## IAC-Er
Indoor anechoic Chamber | [10 empty] | Review comments (RAN4#94)
UID | Uncertainty source | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz | Distribution of the probability | Divisor based on distribution shape | ci | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: DUT measurement
A1-1 | Positioning misalignment between the AAS BS and the reference antenna | x | x | 0.018 | Rectangular | 1.7321 | 1 | x | x | 0.010392
A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | x | x | 0 | Rectangular | 1.7321 | 1 | x | x | 0
A1-3 | Quality of quiet zone | x | x | 0.1 | Gaussian | 1 | 1 | x | x | 0.1
A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | x | x | 0.018 | Rectangular | 1.7321 | 1 | x | x | 0.010392
A1-5 | Mutual coupling between the AAS BS and the receiving antenna | x | x | 0 | Rectangular | 1.7321 | 1 | x | x | 0
A1-6 | Phase curvature | x | x | 0.05 | Gaussian | 1 | 1 | x | x | 0.05
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM) | x | x | 0.6 | Gaussian | 1 | 1 | x | x | 0.6
A1-16 | test system frequency flatness | x | x | 0.25 | Gaussian | 1 | 1 | x | x | 0.25
A1-24 | Uncertainty of the LNA | x | x | 0 | Gaussian | 1 | 1 | x | x | 0
A1.25 | Uncertainty of the Mixer | x | x | 2.25 | Gaussian | 1 | 1 | x | x | 2.25
A1-7 | Impedance mismatch in the receiving chain | x | x | 0.42 | U-shaped | 1.4142 | 1 | x | x | 0.29698
A1-8 | Random uncertainty | x | x | 0.1 | Rectangular | 1.7321 | 1 | x | x | 0.057735
Stage 1: Calibration measurement
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | x | x | 0.57 | U-shaped | 1.4142 | 1 | x | x | 0.40305
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | x | x | 0.43 | Rectangular | 1.7321 | 1 | x | x | 0.24826
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | x | x | 0.57 | U-shaped | 1.4142 | 1 | x | x | 0.40305
A1-3 | Quality of quiet zone | x | x | 0.1 | Gaussian | 1 | 1 | x | x | 0.1
A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | x | x | 0.018 | Rectangular | 1.7321 | 1 | x | x | 0.010392
A1-5 | Mutual coupling between the reference antenna and the receiving antenna | x | x | 0 | Rectangular | 1.7321 | 1 | x | x | 0
A1-6 | Phase curvature | x | x | 0.07 | Gaussian | 1 | 1 | x | x | 0.07
C1-4 | Uncertainty of the absolute gain of the reference antenna | x | x | 0.3 | Gaussian | 1 | 1 | x | x | 0.3
A1-12 | Influence of the reference antenna feed cable | x | x | 0.18 | Rectangular | 1.7321 | 1 | x | x | 0.10392
A1-13 | Reference antenna feed cable loss measurement uncertainty | x | x | 0.1 | Gaussian | 1 | 1 | x | x | 0.1
A1-14 | Influence of the receiving antenna feed cable | x | x | 0.18 | Rectangular | 1.7321 | 1 | x | x | 0.10392
C1-5 | Measurement Reciever (Co-location) | x | x | 0.52 | Rectangular | 1.7321 | 1 | x | x | 0.30022
A1-15 | Uncertainty of the absolute gain of the receiving antenna | x | x | 0 | Rectangular | 1.7321 | 1 | x | x | 0

## EIRP
Table 9.2.7-2: OTA test system specific measurement uncertainty values for the EIRP accuracy, FR2, Normal test conditions
|  | Expanded uncertainty [dB] | [8 empty] | Review comments (RAN4#94)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Indoor Anechoic Chamber
| Compact Antenna Test Range | 1.7368 | 2.0699
One Dimensional Compact Range Chamber
Near Field Test Range
PWS
| Agreed value | 1.7 | 2
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Table 9.2.3.4-1: CATR MU value derivation for EIRP accuracy measurements, Normal test conditions, FR2
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
A2-1a | Misalignment DUT & pointing error (EIRP) | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.01 | 0.01
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) -High power | 0.5 | 0.7 | Gaussian | 1 | 1 | 0.5 | 0.7 | 0.25 | 0.49
A2-2a | Standing wave between DUT and test range antenna | 0.03 | 0.03 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.00045 | 0.00045
A2-3 | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-4a | QZ ripple with DUT | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.16 | 0.16
A2-12 | Frequency flatness | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A2-5a | Mismatch of receiver chain | 0.43 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | 0.09245 | 0.16245
A2-6 | Insertion loss variation in receiver chain | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0
A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.02205 | 0.04205
C1-4 | SGH Calibration uncertainty | 0.52 | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.30022 | 0.090133 | 0.090133
A2-8 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-1b | Misalignment of calibration antenna and test range antenna | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.00405 | 0.00405
A2-4b | QZ ripple calibration antenna | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0.88614 | 1.0561 | 0.88614 | 1.0561
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 1.7368 | 2.0699 | 1.7368 | 2.0699
same as TR38.817-02
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Checked 28/1/20
same as 38.817-02

## power dynamics
| chamber | Expanded uncertainty [dB] | [9 empty] | Review comments (RAN4#94)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Indoor anechoic
CATR
Near field
1D CATR
Plane wave sythesiser
| Agreed value | 0.4 | 0.4 | same as FR1
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0 | #REF!
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0 | #REF!
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!

## EVM
| chamber | Expanded uncertainty [dB] |  | Total Evm on top of existing 1% EVM uncertainty [%]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Indoor anechoic
| CATR | [10 empty] | Review comments (RAN4#94)
Near field
1D CATR
Plane wave sythesiser
| Agreed value |  |  | 1 | 1 | Values for test equipment only (1% EVM, 12Hz Freq error and 25ns TAE)
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0 | #REF!
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!

## EIRP ex
Table 9.3.4-2: Test system specific measurement uncertainty values for the EIRP accuracy in Extreme test conditions, FR2
| chamber | Expanded uncertainty [dB] | [8 empty] | Review comments (RAN4#94)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Indoor anechoic
| CATR | 3.0523 | 3.2534
Near field
1D CATR
Plane wave sythesiser
| Agreed value | 3.1 | 3.3
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Table 9.3.3.5-1: CATR MU value derivation for EIRP accuracy measurements in Extreme test conditions, FR2
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
A2-1a | Misalignment DUT & pointing error (EIRP) | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.01 | 0.01
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) -High power | 0.5 | 0.7 | Gaussian | 1 | 1 | 0.5 | 0.7 | 0.25 | 0.49
A2-2a | Standing wave between DUT and test range antenna | 0.03 | 0.03 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.00045 | 0.00045
A2-3 | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-13 | QZ ripple with DUT (Extreme) | 0.7 | 0.7 | Gaussian | 1 | 1 | 0.7 | 0.7 | 0.49 | 0.49
A2-12 | Frequency flatness | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.0625 | 0.0625
A2-15 | radome loss variation | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.25 | 0.25
A2-14 | wet radome loss variation | 0.9 | 0.9 | Gaussian | 1 | 1 | 0.9 | 0.9 | 0.81 | 0.81
A2-16 | Change in absorber behaviour | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.25 | 0.25
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A2-5a | Mismatch of receiver chain | 0.43 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | 0.09245 | 0.16245
A2-6 | Insertion loss variation in receiver chain | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0
A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.02205 | 0.04205
C1-4 | SGH Calibration uncertainty | 0.52 | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.30022 | 0.090133 | 0.090133
A2-8 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-1b | Misalignment of calibration antenna and test range antenna | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.00405 | 0.00405
A2-4b | QZ ripple calibration antenna | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333
Combined standard uncertainty (1σ) [dB] | [6 empty] | 1.5573 | 1.6599 | 1.5573 | 1.6599
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 3.0523 | 3.2534 | 3.0523 | 3.2534
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Checked 28/1/20
same as 38.817-02

## TX OFF
| chamber | Expanded uncertainty [dB] | [8 empty] | Review comments (RAN4#94)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Indoor anechoic
| CATR | 2.9494 | 3.2939
Near field
1D CATR
Plane wave sythesiser
| Agreed value | 2.9 | 3.3
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Table 9.10.2.3-1: Compact antenna test range uncertainty assessment for EIRP measurements for transmitter OFF power and transmitter transient period
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
A2-1a | Misalignment DUT & pointing error (EIRP) | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.01 | 0.01
C1-10 | RF power measurement equipment standard uncertainty σ (dB) of the absolute level for a time domain wideband measurement for FR2 | 1.25 | 1.45 | Gaussian | 1 | 1 | 1.25 | 1.45 | 1.5625 | 2.1025
A2-2a | Standing wave between DUT and test range antenna | 0.03 | 0.03 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.00045 | 0.00045
A2-3 | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-4a | QZ ripple with DUT | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.16 | 0.16
A2-12 | Frequency flatness | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A2-5b | Mismatch of receiver chain for low power | 0.72 | 0.72 | U-shaped | 1.4142 | 1 | 0.50912 | 0.50912 | 0.2592 | 0.2592
A2-6 | Insertion loss variation in receiver chain | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0
A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.02205 | 0.04205
C1-4 | SGH Calibration uncertainty | 0.51962 | 0.51962 | Rectangular | 1.7321 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A2-8 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-1b | Misalignment of calibration antenna and test range antenna | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.00405 | 0.00405
A2-4b | QZ ripple calibration antenna | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333
Combined standard uncertainty (1σ) [dB] | [6 empty] | 1.5048 | 1.6806 | 1.5048 | 1.6806
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 2.9494 | 3.2939 | 2.9494 | 3.2939
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Checked 28/1/20
pointing error 0.2 in EIRP and 0.3 in TX OFF - why? use 0.2 fpr now.
NA is 0.3 in EIRP and 0.2 here, use 0.3 cancels out pointing error
Reciever missmatch uncertainty different to EIRP? 0.72 rather than 0.43/0.57. Could be becuase of limiter? New error description.
Final values match 38.817-02

## In-band TRP
| chamber | Expanded uncertainty [dB] | [8 empty] | Review comments (RAN4#94)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Indoor anechoic
| CATR | 2.108 | 2.3899
Near field
1D CATR
Plane wave sythesiser
| reverb | 1.8482 | 2.0827
| Agreed value | 2.1 | 2.4
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Table 11.2.3.4-1: CATR MU value derivation for OTA BS output power measurement, FR2
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
A2-1a | Misalignment DUT & pointing error (EIRP) | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.01 | 0.01
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) -High power | 0.5 | 0.7 | Gaussian | 1 | 1 | 0.5 | 0.7 | 0.25 | 0.49
A2-2a | Standing wave between DUT and test range antenna | 0.03 | 0.03 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.00045 | 0.00045
A2-3 | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-4a | QZ ripple with DUT | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.16 | 0.16
A2-12 | Frequency flatness | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A2-5a | Mismatch of receiver chain | 0.43 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | 0.09245 | 0.16245
A2-6 | Insertion loss variation in receiver chain | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0
A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.02205 | 0.04205
C1-4 | SGH Calibration uncertainty | 0.52 | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.30022 | 0.090133 | 0.090133
A2-8 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-1b | Misalignment of calibration antenna and test range antenna | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.00405 | 0.00405
A2-4b | QZ ripple calibration antenna | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0.88426 | 1.0545 | 0.88426 | 1.0545
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 1.7331 | 2.0668 | 1.7331 | 2.0668
TRP Summation error | [6 empty] | 1.2 | 1.2
Total MU | [6 empty] | 2.108 | 2.3899
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
Table 11.2.5.4-1: Reverberation chamber MU value derivation for OTA BS output power, FR2
Reverberation Chamber
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) -High power | 0.5 | 0.7 | Gaussian | 1 | 1 | 0.5 | 0.7 | 0.25 | 0.49
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement | [8 empty] | 0 | 0
A6-3 | Reference antenna radiation efficiency | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.0729 | 0.0729
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.25 | 0.25
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0.94294 | 1.0626 | 0.94294 | 1.0626
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 1.8482 | 2.0827 | 1.8482 | 2.0827
Checked 28/1/20
CATR
-pointing error is different again 0.2 for EIRP, 0.35 here
-RX chain missmatch calc is worng (corrected)
-Misalignment of positioning system, has calc error (std uncertainty is zero though which matches EIRP)
-Misalignment of SGH and pointing system is zero in TRP/TX OFF but is 0.5 here - why?
-Switching uncertainty different (uses EIRP uncertainty as std uncertainty)
-With all above corrected get 1.74/2.07 compared to 1.78/2.06 - so total is ok
Reverb
- NA uncertainty shoudl be 0.3 not 0.2, small incraes in total but under agreed value

## ACLR-abs
Test system specific MU values for the absolute OTA ACLR, Normal test conditions, FR2
| chamber | Expanded uncertainty [dB] | [8 empty] | Review comments (RAN4=94)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Indoor anechoic
| CATR | 2.6922 | 2.7064
Near field
1D CATR
Plane wave sythesiser
| Reverberation Chamber | 2.3594 | 2.3594
| Agreed Vlaue | 2.7 | 2.7
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Table 11.3.3.4-1: CATR MU value derivation for the EIRP measurement of the absolute OTA ACLR, FR2
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
A2-1a | Misalignment DUT & pointing error (EIRP) | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.01 | 0.01
C1-8 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM) | 0.9 | 0.9 | Gaussian | 1 | 1 | 0.9 | 0.9 | 0.81 | 0.81
C1-7 | Standing wave between DUT and test range antenna | 0.03 | 0.03 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.00045 | 0.00045
A2-2a | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-3 | QZ ripple with DUT | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.16 | 0.16
A2-4a | Frequency flatness | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A2-21 | Mismatch of receiver chain for low power | 0.72 | 0.72 | U-shaped | 1.4142 | 1 | 0.50912 | 0.50912 | 0.2592 | 0.2592
A2-6 | Insertion loss variation in receiver chain | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0
C1-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-5a | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.02205 | 0.04205
C1-4 | SGH Calibration uncertainty | 0.51962 | 0.51962 | Rectangular | 1.7321 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A2-6 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-3 | Misalignment of calibration antenna and test range antenna | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-7 | Rotary joints | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0
C1-4 | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.00405 | 0.00405
A2-8 | QZ ripple calibration antenna | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05
A2-1b | Switching uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333
Combined standard uncertainty (1σ) [dB] | [6 empty] | 1.2296 | 1.2377 | 1.2296 | 1.2377
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 2.41 | 2.4259 | 2.41 | 2.4259
TRP Summation error | [6 empty] | 1.2 | 1.2 |  |  | RK: corrected after Nokia comment
Total MU | [6 empty] | 2.6922 | 2.7064
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Table 11.3.5.4-2: Reverberation chamber MU value derivation for absolute ACLR measurement, FR2
Reverberation Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
C1-8 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM) | 0.9 | 0.9 | Gaussian | 1 | 1 | 0.9 | 0.9 | 0.81 | 0.81
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.0729 | 0.0729
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.25 | 0.25
Combined standard uncertainty (1σ) [dB] | [6 empty] | 1.2038 | 1.2038 | 1.2038 | 1.2038
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 2.3594 | 2.3594 | 2.3594 | 2.3594
same as TR 38.817-02
Checked 28/1/20
CATR
-pointing error is different again 0.2 for EIRP, 0.3 here
- missalignment values different to EIRP (where they are 0)
- with EIRP values used the final value is 2.41/2.43 rather than 2.37/2.37 - propose using EIRP values
Reverb
- NA is wrong (0.2 not 0,.3)
total very similar to TR38.817-02 (2.32) smaller than agreed value so OK

## ACLR-rel
Table 11.3.7-4: Test system specific MU values for the relative OTA ACLR, Normal test conditions, FR2
| chamber | Expanded uncertainty [dB] | [8 empty] | Review comments (RAN4#94)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Indoor anechoic
| CATR | 2.2821 | 2.5352
Near field
1D CATR
Plane wave sythesiser
| Reverberation Chamber | 2.1485 | 2.3594
| Agreed Value | 2.3 | 2.6
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 0 | 0 | 0 | 0
TRP Summation error | [6 empty] | 0.75 | 0.75
Total MU | [6 empty] | 0.75 | 0.75
Table 11.3.3.4-2: CATR MU value derivation for the EIRP measurement of the relative OTA ACLR, FR2
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
A2-1a | Misalignment DUT & pointing error (EIRP) | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.01 | 0.01
C1-9 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - relative (ACLR) | 0.75 | 0.9 | Gaussian | 1 | 1 | 0.75 | 0.9 | 0.5625 | 0.81
A2-4a | QZ ripple with DUT | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.16 | 0.16
A2-12 | Frequency flatness | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A2-5a | Mismatch of receiver chain | 0.43 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | 0.09245 | 0.16245
A2-6 | Insertion loss variation in receiver chain | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0
A2-11 | Switching uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 0 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0.99035 | 1.1394 | 0.99035 | 1.1394
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 1.9411 | 2.2333 | 1.9411 | 2.2333
TRP Summation error | [6 empty] | 1.2 | 1.2 |  |  | [rk] corrected
Total MU | [6 empty] | 2.2821 | 2.5352
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Table 11.3.5.4-1: Reverberation chamber MU value derivation for relative ACLR measurement, FR2
Reverberation Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
C1-9 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - relative (ACLR) | 0.75 | 0.9 | Gaussian | 1 | 1 | 0.75 | 0.9 | 0.5625 | 0.81
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement | [8 empty] | 0 | 0
A6-3 | Reference antenna radiation efficiency | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.0729 | 0.0729
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.25 | 0.25
Combined standard uncertainty (1σ) [dB] | [6 empty] | 1.0962 | 1.2038 | 1.0962 | 1.2038
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 2.1485 | 2.3594 | 2.1485 | 2.3594
Checked 28/1/20
CATR
-pointing error is different again 0.2 for EIRP, 0.3 here
- reciever missmatch different to EIRP
-reciever insertion loss variation different (EIRP=0, here, 0.18/0.2)
- Using EIRP value, final value is close 1.94/2.23 vs. 1.96/2.31 propose using EIRP values
Reverb
- NA uncertainty shoudl be 0.3 not 0.2, small incraes in total but under agreed value

## OBUE
Table 11.4.7-2: Test system specific MU values for the OTA OBUE measurement, FR2
| chamber | Expanded uncertainty [dB] | [8 empty] | Review comments (RAN4#94)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Indoor anechoic
| CATR | 2.7011 | 2.7153
Near field
1D CATR
Plane wave sythesiser
| Reverberation Chamber | 2.3594 | 2.3594
| Agreed Vlaue | 2.7 | 2.7
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Table 11.4.3.4-1: CATR MU value derivation for OTA OBUE measurement, FR2
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
A2-18 | Misalignment DUT & pointing error (TRP) | 0.3 | 0.3 | Exp. normal | 2 | 1 | 0.15 | 0.15 | 0.0225 | 0.0225
C1-8 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM) | 0.9 | 0.9 | Gaussian | 1 | 1 | 0.9 | 0.9 | 0.81 | 0.81
A2-2a | Standing wave between DUT and test range antenna | 0.03 | 0.03 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.00045 | 0.00045
A2-3 | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-4a | QZ ripple with DUT | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.16 | 0.16
A2-12 | Frequency flatness | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A2-5b | Mismatch of receiver chain for low power | 0.72 | 0.72 | U-shaped | 1.4142 | 1 | 0.50912 | 0.50912 | 0.2592 | 0.2592
A2-6 | Insertion loss variation in receiver chain | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0
A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.02205 | 0.04205
C1-4 | SGH Calibration uncertainty | 0.51962 | 0.51962 | Rectangular | 1.7321 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A2-8 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-1b | Misalignment of calibration antenna and test range antenna | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.00405 | 0.00405
A2-4b | QZ ripple calibration antenna | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333
Combined standard uncertainty (1σ) [dB] | [6 empty] | 1.2347 | 1.2427 | 1.2347 | 1.2427
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 2.4199 | 2.4357 | 2.4199 | 2.4357
TRP Summation error | [6 empty] | 1.2 | 1.2 |  |  | [rk] corrected
Total MU | [6 empty] | 2.7011 | 2.7153
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Combined standard uncertainty (1σ) [dB] | [6 empty] | 0 | 0
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Table 11.4.5.4-1: Reverberation chamber MU value derivation for OTA OBUE measurement, FR2
Reverberation Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: DUT measurement
C1-8 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM) | 0.9 | 0.9 | Gaussian | 1 | 1 | 0.9 | 0.9 | 0.81 | 0.81
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.0729 | 0.0729
C1-3 | Network Analyzer | 0.3 | 0.3 | Gaussian | 1 | 0.3 | 0.3 | 0.3 | 0.09 | 0.09
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.25 | 0.25
Combined standard uncertainty (1σ) [dB] | [6 empty] | 1.2038 | 1.2038 | 1.2038 | 1.2038
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 2.3594 | 2.3594 | 2.3594 | 2.3594
Checked 29/1/20
CATR
- NA is wrong
-Influence of calibration feed cable is different to ACS abs/EIRP
-numbetr of other minor differences
- corrected total 2.53 vs 2.37 below agreed value so ok
Reverb
- NA is wrong (0.2 not 0,.3)
total very similar to TR38.817-02 (2.32) smaller than agreed value so OK

## OOB EM
| chamber | [6 empty] | Review comments (RAN4#94)
40GHz<f ≤60GHz
| Indoor anechoic chamber | 4.9386
| CATR | 4.957
| Reverberation Chamber | 3.5282
| Agreed Vlaue | 5 | error in 38.817-02, this is recorded as 4.9
Table 12.2.3.x-1: IAC value derivation for TX spurious emissions, FR2
Indoor anechoic chamber
UID | Uncertainty source |  | Distribution of the probability | Divisor based on distribution shape | ci
|  | 40GHz<f ≤60GHz |  |  |  | 40GHz<f ≤60GHz
Stage 2: DUT measurement
A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.018 | Rectangular | 1.7321 | 1 | 0.010392 | 0.000108
A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0 | Rectangular | 1.7321 | 1 | 0 | 0
A1-3 | Quality of quiet zone | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.01
A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.018 | Rectangular | 1.7321 | 1 | 0.010392 | 0.000108
A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | Rectangular | 1.7321 | 1 | 0 | 0
A1-6 | Phase curvature | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.0025
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM) | 0.6 | Gaussian | 1 | 1 | 0.6 | 0.36
A1-16 | test system frequency flatness | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.0625
A1-24 | Uncertainty of the LNA | 0 | Gaussian | 1 | 1 | 0 | 0
A1.25 | Uncertainty of the Mixer | 2.25 | Gaussian | 1 | 1 | 2.25 | 5.0625
A1-7 | Impedance mismatch in the receiving chain | 0.42 | U-shaped | 1.4142 | 1 | 0.29698 | 0.0882
A1-8 | Random uncertainty | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.0033333
Stage 1: Calibration measurement | [6 empty] | 0
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.57 | U-shaped | 1.4142 | 1 | 0.40305 | 0.16245
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.43 | Rectangular | 1.7321 | 1 | 0.24826 | 0.061633
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.57 | U-shaped | 1.4142 | 1 | 0.40305 | 0.16245
A1-3 | Quality of quiet zone | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.01
A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.018 | Rectangular | 1.7321 | 1 | 0.010392 | 0.000108
A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0 | Rectangular | 1.7321 | 1 | 0 | 0
A1-6 | Phase curvature | 0.07 | Gaussian | 1 | 1 | 0.07 | 0.0049
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.09
A1-12 | Influence of the reference antenna feed cable | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.0108
A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.01
A1-14 | Influence of the receiving antenna feed cable | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.0108
C1-4 | SGH Calibration uncertainty | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.090133
A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | Rectangular | 1.7321 | 1 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [5 empty] | 2.4905 | 2.4905
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [5 empty] | 4.8814 | 4.8814
TRP Summation error | [5 empty] | 0.75
Total MU | [5 empty] | 4.9386
Table 12.2.3.4-1: CATR value derivation for TX spurious emissions, FR2
CATR
UID | Uncertainty source |  | Distribution of the probability | Divisor based on distribution shape | ci
|  | 40GHz<f ≤60GHz |  |  |  | 40GHz<f ≤60GHz
Stage 2: DUT measurement
A2-18 | Misalignment DUT & pointing error (TRP) | 0.3 | Exp. normal | 2 | 1 | 0.15 | 0.0225
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM) | 0.6 | Gaussian | 1 | 1 | 0.6 | 0.36
A2-19 | Uncertainty of the LNA | 0 | Gaus | 1 | 1 | 0 | 0
A2-20 | Uncertainty of the Mixer | 2.25 | Gaus | 1 | 1 | 2.25 | 5.0625
A2-2a | Standing wave between DUT and test range antenna | 0.21 | Gaussian | 1 | 1 | 0.21 | 0.0441
A2-3 | RF leakage, test range antenna cable connector terminated. | 0.001 | Gaussian | 1 | 1 | 0.001 | 1e-06
A2-4a | QZ ripple with DUT | 0.0928 | Gaussian | 1 | 1 | 0.0928 | 0.0086118
A2-12 | Frequency flatness | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.0625
A2-21 | Miscellaneous uncertainty | 0 | Rectangular | 1.7321 | 1 | 0 | 0
Stage 1: Calibration measurement | [6 empty] | 0
C1-3 | Network Analyzer | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.09
A2-5a | Mismatch of receiver chain | 0.57 | U-shaped | 1.4142 | 1 | 0.40305 | 0.16245
A2-5b | Mismatch of receiver chain for low power | 0 | U-shaped | 1.4142 | 1 | 0 | 0
A2-6 | Insertion loss variation in receiver chain | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.0108
A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.001 | Gaussian | 1 | 1 | 0.001 | 1e-06
A2-7 | Influence of the calibration antenna feed cable | 0.29 | U-shaped | 1.4142 | 1 | 0.20506 | 0.04205
C1-4 | SGH Calibration uncertainty | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.090133
A2-8 | Misalignment positioning system | 0 | Exp. normal | 2 | 1 | 0 | 0
A2-1b | Misalignment of calibration antenna and test range antenna | 0 | Exp. normal | 2 | 1 | 0 | 0
A2-9 | Rotary joints | 0 | U-shaped | 1.4142 | 1 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.00405
A2-4b | QZ ripple calibration antenna | 0.009 | Gaussian | 1 | 1 | 0.009 | 8.1e-05
A2-11 | Switching uncertainty | 0.43 | Rectangular | 1.7321 | 1 | 0.24826 | 0.061633
Combined standard uncertainty (1σ) [dB] | [5 empty] | 2.4539 | 2.4539
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [5 empty] | 4.8096 | 4.8096
TRP Summation error | [5 empty] | 1.2
Total MU | [5 empty] | 4.957
Table 12.2.3.4-1-2: Reverberation Chamber MUperpoint value derivation for TX spurious emissions, 18 GHz – 60 GHz
Reverberation Chamber
UID | Uncertainty source |  | Distribution of the probability | Divisor based on distribution shape | ci
|  | 40GHz<f ≤60GHz |  |  |  | 40GHz<f ≤60GHz
Stage 2: DUT measurement
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM) | 0.6 | Gaussian | 1 | 1 | 0.6 | 0.36
A6-1 | Impedance mismatch in the receiving chain | 0.45 | U-shaped | 1.4142 | 1 | 0.3182 | 0.10125
A6-2 | Random uncertainty | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.5 | Normal | 1 | 1 | 0.5 | 0.25
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.27 | Normal | 1 | 1 | 0.27 | 0.0729
C1-3 | Network Analyzer | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.09
A6-5 | Influence of the reference antenna feed cable | 0.2 | Normal | 1 | 1 | 0.2 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | Normal | 1 | 1 | 0.27 | 0.0729
A6-7 | Uniformity of transfer function | 1.5 | Normal | 1 | 1 | 1.5 | 2.25
Combined standard uncertainty (1σ) [dB] | [5 empty] | 1.8001 | 1.8001
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [5 empty] | 3.5282 | 3.5282
Checked 29/1/20
Below 40GHz some discrepancies biut suggest using in-band TRP limiits (this is whjat is done in 38.817-02)
IAB
-power measurement different for 3 methods
CATR
-power measurement different for 3 methods
Reverb
power measurement equipment much larger than other methods - should synchronise

## RX EM
| chamber | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
| Indoor anechoic chamber | x | x | x
| CATR | x | x | x
| Reverberation Chamber | x | x | x
| Agreed Vlaue | x | x | 5
Indoor anechoic chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz |  |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz |  |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Reverberation Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz |  |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Checked 29/1/20
same as TX emssisons