3GPP TSG-RAN WG4 Meeting #97-e	R4-2016370

Electronic Meeting, 2nd – 13th November 2020

Source:		Rohde & Schwarz

Title:	Plane Wave Synthesizer – Pending MU terms from 4.2GHz to 6GHz

Agenda Item:		7.17.2

Document for:	Discussion

# Introduction

Following the effort from previous meetings where Plane Wave Synthesizer (PWS) was included as accepted method in TR 37.941 [1] and the MU assessment was introduced together with remaining methods, this contribution presents an analysis of the MU terms that were still pending for the frequency range from 4.2GHz to 6GHz.

# MU budgets

The following is the MU budget table applicable to PWS for EIRP measurements as agreed in [1], where the MU terms that required further analysis for the frequency range 4.2 to 6GHz are highlighted:

#### 9.2.6.3	MU value derivation, FR1

Table 9.2.6.3-1 captures the uncertainty budget contributors and table 9.2.6.3-2 captures the derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in PWS.

Table 9.2.6.3-1: PWS measurement accuracy contributors for EIRP accuracy measurements, FR1

| UID / Details in annex | Uncertainty source |
| --- | --- |
| Stage 2: BS measurement |  |
| A7-1a | Misalignment BS & pointing error |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) |
| A7-2a | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for BS antenna |
| A7-3 | RF leakage (calibration antenna connector terminated) |
| A7-4a | QZ ripple with BS |
| A7-5 | Miscellaneous Uncertainty |
| A7-14 | System non-linearity |
| A7-13 | Frequency Flatness |
| Stage 1: Calibration measurement |  |
| C1-3 | Uncertainty of the network analyzer |
| A7-6 | Mismatch (i.e. reference antenna, network analyser and reference cable) |
| A7-7 | Insertion loss variation |
| A7-3 | RF leakage (calibration antenna connector terminated) |
| A7-8 | Influence of the calibration antenna feed cable |
| C1-4 | Uncertainty of the absolute gain of the reference antenna |
| A7-9 | Misalignment of positioning system |
| A7-1b | Misalignment of calibration antenna & pointing error |
| A7-10 | Rotary joints |
| A7-2b | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna |
| A7-4b | QZ ripple with calibration antenna |
| A7-11 | Switching uncertainty |
| A7-12 | Field repeatability |

# Error contribution descriptions

The following sections describe in detail the corresponding evaluation, based on measurements in the range of 4.2GHz to 6GHz on the actual R&S implementation of PWS, for each of the MU terms highlighted in previous tables.

## Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna/calibration antenna

As described in [4], this MU term covers both the standing wave between DUT or calibration antenna and the test range antenna as per section but it also counts for the PWS imperfect field synthesis over distance.

It is evaluated by means of an S21 measurement between the PWS and the DUT or calibration antenna, where the measurement distance is varied ±35mm from the reference at 1500mm. This variation corresponds to ~λ02 for the lowest frequency range (i.e. 2300 MHz).

The maximum coupling deviation for each center frequency is calculated as follows, assuming a rectangular distribution:

ufc,m=maxdi20logS21di,fc,mS211500,fc,m

| UID | Uncertainty Source | Value 4.2 < f ≤ 6 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) 4.2 < f ≤ 6 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A7-2a | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.20 | dB | Rectangular | √3 | 1 | 0.12 |
| A7-2b | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.15 | dB | Rectangular | √3 | 1 | 0.09 |

## QZ ripple with DUT/calibration antenna

Similar to what was described for lower frequencies in [5], additional testing has been performed and the metric has been selected to evaluate the QZ amplitude ripple as per the most recent draft for IEEE Std. 149 [7].

This amplitude ripple MU term is measured by determining the variation of amplitude about the 2nd degree polynomial fit, reported in dB. Example of the amplitude ripple is shown in the following figure:

Figure 3.1-1: 1D amplitude ripple is 0.22 dB; f = 5000 MHz

Since the field probe measurements are performed over the reference plane where the EUT is placed, there are several possible 1D curves around the axis. The worst case per frequency range among all 1D curves is selected as MU value, and a rectangular distribution is assumed.

| UID | Uncertainty Source | Value 4.2 < f ≤ 6 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) 4.2 < f ≤ 6 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A7-4a | Normalized QZ ripple with DUT | 0.57 | dB | Rectangular | √3 | 1 | 0.33 |

The field distribution measurements (magnitude) are performed using a broadband dual polarized probe R&S TC-TA18 and a Vector Network Analyzer (e.g. R&S ZVA8) in order to get results like the one shown in figure 3.1-2.

Figure 3.1-2: QZ field distribution, f = 5000 MHz.

## System non-linearity

As described in TR 37.941 [1], clause A.7, this uncertainty term is calculated as RSS of the following items, assuming a rectangular distribution:

- System non-linearity in time. This is assessed by repeated measurements over a period of time (e.g. 60 minutes) for the same reference power transmitted by the reference antenna. The largest difference between the results is recorded as the uncertainty.

- System non-linearity in power. This is assessed by repeated measurements over a range of transmitted powers. The largest delta between the increments on the receiving side versus the transmitting side is recorded as the uncertainty.

Following the same approach presented in [6] a test campaign was performed on R&S PWC200 to obtain the corresponding MU estimation for the frequency range between 4.2GHz and 6GHz.

According to the results from this measurement campaign, system non-linearity uncertainty is calculated using the worst case for each frequency range and combined as RSS for the two items:

| UID | Uncertainty Source | Value 4.2 < f ≤ 6 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) 4.2 < f ≤ 6 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A7-14 | System non-linearity | 0.15 | dB | Rectangular | √3 | 1 | 0.09 |

## Field repeatability

As described in [4], this MU term is assessed by repeated measurements after field calibration and the uncertainty terms is defined based on the maximum error between two individual measurements as:

u=maxi20logS21i-mini20logS21i

| UID | Uncertainty Source | Value 4.2 < f ≤ 6 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) 4.2 < f ≤ 6 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A7-14 | Field calibration repeatability | 0.15 | dB | Normal | 1 | 1 | 0.15 |

# Conclusion

In this contribution, proposed values for the pending MU terms for the Plane Wave Synthesizer has been provided to finalize the MU for 4.2GHz to 6GHz frequency range.

The corresponding changes to TR 37.941 [1] are provided in the separate CR [2] and [3].

# References

- 3GPP TR 37.941 v15.1.0 (2020-09)

- R4-2016466, “CR to TR 37.941: Completion of MU terms for PWS”, Rohde & Schwarz, RAN4#97-e, November 2020

- R4-2016467, “Mirror CR to TR 37.941: Completion of MU terms for PWS”, Rohde & Schwarz, RAN4#97-e, November 2020

- R4-1907993, “2D Compact Range for BS Testing – MU budget analysis”, Rohde & Schwarz, RAN4#92, August 2019

- R4-1914796, “Plane Wave Synthesizer – Pending MU terms”, Rohde & Schwarz, RAN4#93, November 2019

- R4-2011203, “Plane Wave Synthesizer – Pending MU term”, Rohde & Schwarz, RAN4#96-e, August 2020

- P149™/D8: “Draft Recommended Practice for Antenna Measurements”, Antennas and Propagation Standards Committee of the IEEE Antennas and Propagation Society, April 2020.

- R.  Haupt,  “Generating  a  Plane  Wave  in  the  Near  Field  with  a  Planar Array Antenna”, Microwave Journal, September 2003

- C. Rowell, A. Tankielun, “Plane Wave Converter for 5G Massive MIMO Basestation Measurements”, the 12th European Conference on Antennas and Propagation (EuCAP2018), London, UK, 9 - 13 April, 2018

- B. Derat, “Creating Far-field Condition at a Tenth of Far-field Distance: An Innovative Technique for 5G OTA measurement”, Electronic Design Innovation Conference (EDICON), Beijing, China, 20 – 22 March, 2018