## summary
| Expanded uncertainty [dB] | [5 empty] | A method exceeds agreed value
| IAC |  | CATR |  | Agreed value
| 24.25<f<29.5GHz | 37<f<40GHz | 24.25<f<29.5GHz | 37<f<40GHz | 24.25<f<29.5GHz | 37<f<40GHz | 24.25<f<29.5GHz | 37<f<40GHz
EIS | 2.3347 | 2.455 | 2.2511 | 2.3266 | 2.4 | 2.4 |  | x

## TE
Review comments (RAN4#94)
Test Equipment Errors
UID | Instrument | Uncertainty value (for the prob dist type) |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 24.25<f<29.5GHz | 37<f<40GHz |  |  | 24.25<f<29.5GHz | 37<f<40GHz
C1-1
C1-2 | Uncertainty of the RF signal generator | 0.9 | 0.9 | Gaussian | 1 | 0.9 | 0.9 | Ericsson use 0.57, MVG and Keysight use 0.9. FR1 value was 0.46. As TE use Keysight number
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | Gaussian | 1 | 0.3 | 0.3 | Ericsson use 0.3, this is same as agreed in Tx. MVG use 0.4, Keysight 0.2.
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 | 0.52 | Rectangular | 1.7321 | 0.30022 | 0.30022 | MVG No is 1.6 much higher than0.52 used by Ericsson and Keysight, TX ref antenna also used 0.52

## IA-Er
Indoor anachoic | [8 empty] | Review comments (RAN4#94)
|  | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f<29.5GHz | 37<f<40GHz |  |  |  | 24.25<f<29.5GHz | 37<f<40GHz
Stage 2: DUT measurement
B1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392
B1-2 | Pointing misalignment between the AAS BS and the transmitting antenna | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392
B1-3 | Quality of quiet zone | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1
B1-4a | Polarization mismatch between the AAS BS and the transmitting antenna | 0.018 | 0.018 | Rectangular | 1.7321 | 1 | 0.010392 | 0.010392
B1-5a | Mutual coupling between the AAS BS and the transmitting antenna | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0
B1-6 | Phase curvature | 0.07 | 0.07 | Gaussian | 1 | 1 | 0.07 | 0.07
C1-2 | Uncertainty of the RF signal generator | 0.9 | 0.9 | Gaussian | 1 | 1 | 0.9 | 0.9
B1-7 | Impedance mismatch in the transmitting chain | 0.42 | 0.42 | U-shaped | 1.4142 | 1 | 0.29698 | 0.29698
B1-8 | Random uncertainty | 0.18 | 0.25 | Rectangular | 1.7321 | 1 | 0.10392 | 0.14434
Stage 1: Calibration measurement
B1-9 | Impedance mismatch between the transmitting antenna and the network analyzer | 0.43 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305
B1-10 | Positioning and pointing misalignment between the reference antenna and the transmitting antenna | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.24826 | 0.24826
B1-11 | Impedance mismatch between the reference antenna and network analyzer | 0.43 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305
B1-3 | Quality of quiet zone | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1
B1-4b | Polarization mismatch between the reference antenna and the transmitting antenna | 0.018 | 0.018 | Rectangular | 1.7321 | 1 | 0.010392 | 0.010392
B1-5b | Mutual coupling between the reference antenna and the transmitting antenna | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0
B1-6 | Phase curvature | 0.07 | 0.07 | Gaussian | 1 | 1 | 0.07 | 0.07
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3
B1-12 | Influence of the reference antenna feed cable | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392
B1-13 | Reference antenna feed cable loss measurement uncertainty | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1
B1-14 | Influence of the transmitting antenna feed cable | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.30022
B1-15 | Uncertainty of the absolute gain of the transmitting antenna | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0

## CATR-Er
CATR | [8 empty] | Review comments (RAN4#94)
|  | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f<29.5GHz | 37<f<40GHz |  |  |  | 24.25<f<29.5GHz | 37<f<40GHz
Stage 2: DUT measurement
B2-1a | Misalignment DUT & pointing error | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | Ericsson 0.2, MVG, Keysignt 0.1. 0.2 is consiatnt with same error in Tx
B2-2 | Standing wave between DUT and test range antenna | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | MVG 0.1, Keysight 0.03, Ericsson 0.21. . O.21 is consiatnt with Tx
C1-2 | Uncertainty of the RF signal generator | 0.9 | 0.9 | Gaussian | 1 | 1 | 0.9 | 0.9
B2-3 | RF leakage & dynamic range, test range antenna cable connector terminated. | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | MVG & Keysight 0.1, Ericsson 0.01. . O.o1 is consiatnt with Tx
B2-4a | QZ ripple with DUT | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | All agree 0.4
B2-9 | Miscellaneous uncertainty | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | All agree 0.4
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3
B2-5 | Mismatch of transmit chain (i.e. between transmitting measurement antenna and DUT) | 0.43 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | MVG 0.5, Keysight 0.74, Ericsson 0.43/0.57., equivilment missmatch in Tx (Rx chain) is same as Ericsson
B2-6 | Insertion loss of transmitter chain | 0.12 | 0.12 | Rectangular | 1.7321 | 1 | 0.069282 | 0.069282 | MVG 0, Keysight 0.12, Ericsson 0.26, No equivilment missmatch in Tx, take mean 0.12
B2-7 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | MVG & Keysight 0.1, Ericsson 0.01. . 0.01 is consiatnt with Tx
B2-8 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | MVG & Keysight 0, Ericsson 0.21. . In TX we use 0.21/0.29 use same here
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.30022
B2-11 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | This is the same error as B2-11, should be misalignmnet of the positioning system? Zero
B2-4b | QZ ripple with SGH | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | MVG 0.2, Keysight 0.072, Ericsson 0.1, Tx uses 0.01!, use 0.1
B2-10 | Rotary joints | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | All 0
B2-11b | Misalignment calibration antenna & pointing error | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | MVG & Keysight 0, Ericsson 0.3 . In TX we use 0
| Misalignment calibration system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | All 0, not used in any other budget - remove
B2-12 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | MVG 0.1, Keysight 0.03, Ericsson 0.09, Tx uses 0.09!, use 0.09
B2-13 | Switching uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | MVG 0, Keysight 0.1, Ericsson 0.43, Tx uses 0.1, use 0.1

## EIS
| chamber | Expanded uncertainty [dB] | [8 empty] | Review comments (RAN4#94)
|  | 24.25<f<29.5GHz | 37<f<40GHz
| Indoor anachoic | 2.3347 | 2.455
| CATR | 2.2511 | 2.3266
| Agreed value | 2.4 | 2.4
Indoor anachoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f<29.5GHz | 37<f<40GHz |  |  |  | 24.25<f<29.5GHz | 37<f<40GHz
Stage 2: DUT measurement
B1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.0108 | 0.0108
B1-2 | Pointing misalignment between the AAS BS and the transmitting antenna | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.0108 | 0.0108
B1-3 | Quality of quiet zone | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.01 | 0.01
B1-4a | Polarization mismatch between the AAS BS and the transmitting antenna | 0.018 | 0.018 | Rectangular | 1.7321 | 1 | 0.010392 | 0.010392 | 0.000108 | 0.000108
B1-5a | Mutual coupling between the AAS BS and the transmitting antenna | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0
B1-6 | Phase curvature | 0.07 | 0.07 | Gaussian | 1 | 1 | 0.07 | 0.07 | 0.0049 | 0.0049
C1-2 | Uncertainty of the RF signal generator | 0.9 | 0.9 | Gaussian | 1 | 1 | 0.9 | 0.9 | 0.81 | 0.81
B1-7 | Impedance mismatch in the transmitting chain | 0.42 | 0.42 | U-shaped | 1.4142 | 1 | 0.29698 | 0.29698 | 0.0882 | 0.0882
B1-8 | Random uncertainty | 0.18 | 0.25 | Rectangular | 1.7321 | 1 | 0.10392 | 0.14434 | 0.0108 | 0.020833
Stage 1: Calibration measurement | [8 empty] | 0 | 0
B1-9 | Impedance mismatch between the transmitting antenna and the network analyzer | 0.43 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | 0.09245 | 0.16245
B1-10 | Positioning and pointing misalignment between the reference antenna and the transmitting antenna | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.24826 | 0.24826 | 0.061633 | 0.061633
B1-11 | Impedance mismatch between the reference antenna and network analyzer | 0.43 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | 0.09245 | 0.16245
B1-3 | Quality of quiet zone | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.01 | 0.01
B1-4b | Polarization mismatch between the reference antenna and the transmitting antenna | 0.018 | 0.018 | Rectangular | 1.7321 | 1 | 0.010392 | 0.010392 | 0.000108 | 0.000108
B1-5b | Mutual coupling between the reference antenna and the transmitting antenna | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0
B1-6 | Phase curvature | 0.07 | 0.07 | Gaussian | 1 | 1 | 0.07 | 0.07 | 0.0049 | 0.0049
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
B1-12 | Influence of the reference antenna feed cable | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.0108 | 0.0108
B1-13 | Reference antenna feed cable loss measurement uncertainty | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.01 | 0.01
B1-14 | Influence of the transmitting antenna feed cable | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.0108 | 0.0108
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.30022 | 0.090133 | 0.090133
B1-15 | Uncertainty of the absolute gain of the transmitting antenna | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [6 empty] | 1.1912 | 1.2526 | 1.1912 | 1.2526
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 2.3347 | 2.455 | 2.3347 | 2.455
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 24.25<f<29.5GHz | 37<f<40GHz |  |  |  | 24.25<f<29.5GHz | 37<f<40GHz
Stage 2: DUT measurement
B2-1a | Misalignment DUT & pointing error | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0.01 | 0.01
B2-2 | Standing wave between DUT and test range antenna | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.02205 | 0.02205
C1-2 | Uncertainty of the RF signal generator | 0.9 | 0.9 | Gaussian | 1 | 1 | 0.9 | 0.9 | 0.81 | 0.81
B2-3 | RF leakage & dynamic range, test range antenna cable connector terminated. | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
B2-4a | QZ ripple with DUT | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.16 | 0.16
B2-9 | Miscellaneous uncertainty | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3 | 0.09 | 0.09
B2-5 | Mismatch of transmit chain (i.e. between transmitting measurement antenna and DUT) | 0.43 | 0.57 | U-shaped | 1.4142 | 1 | 0.30406 | 0.40305 | 0.09245 | 0.16245
B2-6 | Insertion loss of transmitter chain | 0.12 | 0.12 | Rectangular | 1.7321 | 1 | 0.069282 | 0.069282 | 0.0048 | 0.0048
B2-7 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
B2-8 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506 | 0.02205 | 0.04205
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.30022 | 0.090133 | 0.090133
B2-11 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
B2-4b | QZ ripple with SGH | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.01 | 0.01
B2-10 | Rotary joints | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0
B2-11b | Misalignment calibration antenna & pointing error | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
0 | Misalignment calibration system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
B2-12 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.00405 | 0.00405
B2-13 | Switching uncertainty | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333
Combined standard uncertainty (1σ) [dB] | [6 empty] | 1.1485 | 1.187 | 1.1485 | 1.187
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [6 empty] | 2.2511 | 2.3266 | 2.2511 | 2.3266