## Summary
[7 empty] | A method exceeds agreed value |  |  | Review comments (RAN4#94)
| General chamber |  |  | Common maximum accepted test system uncertainty
| f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
COLO EM | 3.0472 | 3.2874 | 3.3437 | 3.2 | 3.4 | 3.4
COLO OFF | 3.3691 | 3.5878 | 3.6395 | 3.4 | 3.6 | 3.6
TX IMD | 3.2364 | 3.3682 | 3.4996 | 3.2 | 3.4 | 3.5
COLO BLK | 3.2423 | 3.3739 | 3.505 | See below |  |  |  | x | x
OTA Blocking final MU values
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
EIS |  | 0.7 | 1 | 1.22
COLO BLK |  | 3.2423 | 3.3739 | 3.505
NoIse contribution |  | 0.1 | 0.1 | 0.1
Co-location blocing MU (dB) |  | Wanted signal operating band
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Interferer Fequency (MHz) | f<3 GHz | 3.417 | 3.493 | 3.5642
| 3<f<4.2 GHz | 3.5457 | 3.619 | 3.8113
| 4.2<f<6 GHz | 3.6742 | 3.7449 | 3.8113

## TE
Review comments (RAN4#94)
Test Equipment Errors
UID | Instrument | Uncertainty value (for the prob dist type) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 0.14 | 0.26 | 0.26
C1-2 | Uncertainty of the RF signal generator | 0.46 | 0.46 | 0.46 | Gaussian | 1 | 0.46 | 0.46 | 0.46
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 0.13 | 0.2 | 0.2
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 0.29 | 0.25 | 0.25
C1-5 | Measurement Reciever (Co-location) | 0.41 | 0.74 | 0.8 | Gaussian | 1 | 0.41 | 0.74 | 0.8
C1-6 | Noise figure measurememt accuracy | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 0.2 | 0.2 | 0.2
Test Equipment Errors (same errors and values are consistant with above, buit greater freq range)
UID | Instrument | Uncertainty value (for the prob dist type) |  |  |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | f<3 GHz | 3<f<3.8 GHz | 3.8<f<12.75 GHz | 12.75<f<19 GHz | 19<f<26 GHz |  |  | f<3 GHz | 3<f<3.8 GHz | 3.8<f<12.75 GHz | 12.75<f<19 GHz | 19<f<26 GHz
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | 0.37 | 0.37 | Gaussian | 1 | 0.14 | 0.26 | 0.26 | 0.37 | 0.37
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 0.13 | 0.2 | 0.2 | 0.2 | 0.2
Conducted MU (minus missmatch) - up to 6GHz
| Instrument | Uncertainty value (for the prob dist type) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
C3-1 | DL-RS MU derived from conducted spec | 0.41 | 0.56 | 0.56 | Gaussian | 1 | 0.41 | 0.56 | 0.56
C3-2 | Total power dynamic range conducted unbcertainty | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 0.2 | 0.2 | 0.2
C3-3 | ACLR relative - MU of TE from conducted | 0.41 | 0.61 | 0.61 | Gaussian | 1 | 0.41 | 0.56 | 0.56 | these 3 are not used, remove if current method is agreed
C3-4 | ACLR absolute, Conducted minus mismatch | 1.0108 | 1.2321 | 1.2321 | Gaussian | 1 | 1.0108 | 1.2321 | 1.2321
C3-5 | OBUE/SEM, Conducted minus mismatch | 0.75239 | 0.85703 | 0.85703 | Gaussian | 1 | 0.75239 | 0.85703 | 0.85703
C3-8 | Additional (COEX) emssions - Condcuted Uncertainty (minus missmatch) | 1.02 | 1.28 | 1.53 | Gaussian | 1 | 1.02 | 1.28 | 1.53
C3-7 | TX IMD - conducted measurement uncertainty | 1 | 1.1 | 1.2 | Gaussian | 1 | 1 | 1.1 | 1.2
C3-9 | Colocation blocking - conducted measurement uncertainty | 1 | 1.1 | 1.2 | Gaussian | 1 | 1 | 1.1 | 1.2
Conducted MU (minus missmatch) - Whole Spectrum
| Instrument | Uncertainty value (for the prob dist type) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz |  |  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz
C3-6 | OOB emsions - Conducted Uncertainty (minus missmatch) | 1 | 2 | 2 | Gaussian | 1 | 1 | 2 | 2
C3-7 | Rx emisions - Conducted Uncertainty (minus missmatch) | 1 | 2 | 2 | Gaussian | 1 | 1 | 2 | 2
Co-location errors
| Co-lcoation Error source | Uncertainty value (for the prob dist type) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 4.2<f<6 GHz |  |  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 4.2<f<6 GHz
C2-9 | Uncertainty related to the selection of the CLTA (Note) | 1.5 | 1.5 | 1.5 | Rectangular | 1.7321 | 0.86603 | 0.86603 | 0.86603
C2-10 | Uncertainty related to the placement of the CLTA (Note) | 1.7 | 1.7 | 1.7 | Rectangular | 1.7321 | 0.9815 | 0.9815 | 0.9815
C2-11 | Uncertainty related to measuring close to noise floor - Emisions | 0.68 | 0.68 | 0.68 | Gaussian | 1 | 0.68 | 0.68 | 0.68
C2-12 | Uncertainty related to measuring close to noise floor Tx OFF | 1 | 1 | 1 | Gaussian | 1 | 1 | 1 | 1
C2-13 | Impedance mismatch between feeder cable and CLTA | 0.14 | 0.23 | 0.25 | U-Shaped | 1.4142 | 0.098995 | 0.16263 | 0.17678
C2-14 | Gain variations in LNA | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 0.1 | 0.1 | 0.1
C2-15 | Gain variations in measurement amplifier | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 0.1 | 0.1 | 0.1

## Gen-Er
General Directional Chamber | [10 empty] | Review comments (RAN4#94)
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz |  |  |  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz
Stage 2: DUT measurement
A5-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321
A5-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
A5-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1
A5-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735
A5-5a | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
A5-6a | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05
C3-6 | OOB emsions - Conducted Uncertainty (minus missmatch) | 1 | 2 | 2 | Gaussian | 1 | 1 | 1 | 2 | 2
A5-7 | Impedance mismatch in the receiving chain | 0.2 | 0.45 | 0.45 | U-shaped | 1.4142 | 1 | 0.14142 | 0.3182 | 0.3182
A5-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735
A5-17 | Measurement antenna frequency variation | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735
A5-18 | FSPL estimation error | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A5-16 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25
A5-19 | Measurement system dynamic range uncertainty | 0.51 | 0.51 | 0.51 | Gaussian | 1 | 1 | 0.51 | 0.51 | 0.51
A5-20 | Reflections in anechoic chamber | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01
Stage 1: Calibration measurement
A5-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355
A5-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735
A5-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355
A5-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1
A5-4b | Polarization mismatch for reference antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735
A5-5b | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
A5-6b | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2
A5-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868
A5-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06
A5-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25
A5-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0

## COLO OFF
Table 13.2.5-1: Common maximum accepted test system uncertainty value derivation for TDD OFF power level measurement
| chamber | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
| General Directional Chamber | 3.3691 | 3.5878 | 3.6395
CATR
Near field
1D CATR
Plane wave sythesiser
| Common maximum accepted test system uncertainty | 3.4 | 3.6 | 3.6
Table 13.2.4.3-1: General OTA chamber MU value derivation for TDD OFF power level measurement
General Directional Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement | [10 empty] | 0 | 0 | 0
C2-9 | Uncertainty related to the selection of the CLTA (Note) | 1.5 | 1.5 | 1.5 | Rectangular | 1.7321 | 1 | 0.86603 | 0.86603 | 0.86603 | 0.75 | 0.75 | 0.75
C2-10 | Uncertainty related to the placement of the CLTA (Note) | 1.7 | 1.7 | 1.7 | Rectangular | 1.7321 | 1 | 0.9815 | 0.9815 | 0.9815 | 0.96333 | 0.96333 | 0.96333
C2-12 | Uncertainty related to measuring close to noise floor Tx OFF | 1 | 1 | 1 | Gaussian | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
C2-13 | Impedance mismatch between feeder cable and CLTA | 0.14 | 0.23 | 0.25 | U-Shaped | 1.4142 | 1 | 0.098995 | 0.16263 | 0.17678 | 0.0098 | 0.02645 | 0.03125
C2-14 | Gain variations in LNA | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A5-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | #REF! | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
C1-5 | Measurement Reciever (Co-location) | 0.41 | 0.74 | 0.8 | Gaussian | 1 | 1 | 0.41 | 0.74 | 0.8 | 0.1681 | 0.5476 | 0.64
A5-20 | Reflections in anechoic chamber | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01 | 0.0001 | 0.0001 | 0.0001 | RK:was reflections in anachoic chamber ? Repleced with quaility of quiet zone as already defned, value is 0.1 rathetr than 0.01 but make no difference to total
C2-15 | Gain variations in measurement amplifier | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-6 | Noise figure measurememt accuracy | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
Combined standard uncertainty (1σ) [dB] | [7 empty] | 1.7189 | 1.8305 | 1.8569 | 1.7189 | 1.8305 | 1.8569
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 3.3691 | 3.5878 | 3.6395 | 3.3691 | 3.5878 | 3.6395
same value as TR37.843
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
|  | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
check 28/1/20
OK-same as 37.843

## COLO-EM
Table 13.3.5-1: Common maximum accepted test system uncertainty value derivation for co-location emissions level measurement
| chamber | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
| General Directional Chamber | 3.0472 | 3.2874 | 3.3437
CATR
Near field
1D CATR
Plane wave sythesiser
| Common maximum accepted test system uncertainty | 3.2 | 3.4 | 3.4
Table 13.3.2.3-1: General OTA chamber MU value derivation for co-location emissions level measurement
General Directional Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement | [10 empty] | 0 | 0 | 0
C2-9 | Uncertainty related to the selection of the CLTA (Note) | 1.5 | 1.5 | 1.5 | Rectangular | 1.7321 | 1 | 0.86603 | 0.86603 | 0.86603 | 0.75 | 0.75 | 0.75
C2-10 | Uncertainty related to the placement of the CLTA (Note) | 1.7 | 1.7 | 1.7 | Rectangular | 1.7321 | 1 | 0.9815 | 0.9815 | 0.9815 | 0.96333 | 0.96333 | 0.96333
C2-11 | Uncertainty related to measuring close to noise floor - Emisions | 0.68 | 0.68 | 0.68 | Gaussian | 1 | 1 | 0.68 | 0.68 | 0.68 | 0.4624 | 0.4624 | 0.4624
C2-13 | Impedance mismatch between feeder cable and CLTA | 0.14 | 0.23 | 0.25 | U-Shaped | 1.4142 | 1 | 0.098995 | 0.16263 | 0.17678 | 0.0098 | 0.02645 | 0.03125
C2-14 | Gain variations in LNA | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A5-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | #REF! | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
C1-5 | Measurement Reciever (Co-location) | 0.41 | 0.74 | 0.8 | Gaussian | 1 | 1 | 0.41 | 0.74 | 0.8 | 0.1681 | 0.5476 | 0.64
A5-20 | Reflections in anechoic chamber | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01 | 0.0001 | 0.0001 | 0.0001 | RK:was reflections in anachoic chamber ? Repleced with quaility of quiet zone as already defned, value is 0.1 rathetr than 0.01 but make no difference to total
C2-15 | Gain variations in measurement amplifier | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-6 | Noise figure measurememt accuracy | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
Combined standard uncertainty (1σ) [dB] | [7 empty] | 1.5547 | 1.6773 | 1.706 | 1.5547 | 1.6773 | 1.706
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 3.0472 | 3.2874 | 3.3437 | 3.0472 | 3.2874 | 3.3437
same as 37.843
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
|  | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
check 28/1/20
OK-same as 37.843

## TXIMD
Table 13.4.5-1: Common maximum accepted test system uncertainty value derivation for co-location TX IMD interferer signal level
| chamber | Expanded uncertainty [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
| General Chamber | 3.2364 | 3.3682 | 3.4996
CATR
Near field
1D CATR
Plane wave sythesiser
| Common maximum accepted test system uncertainty | 3.2 | 3.4 | 3.5
Table 13.4.4.3-1 MU for OTA transmitter intermodulation interferer signal
General Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
C2-9 | Uncertainty related to the selection of the CLTA (Note) | 1.5 | 1.5 | 1.5 | Rectangular | 1.7321 | 1 | 0.86603 | 0.86603 | 0.86603 | 0.75 | 0.75 | 0.75
C2-10 | Uncertainty related to the placement of the CLTA (Note) | 1.7 | 1.7 | 1.7 | Rectangular | 1.7321 | 1 | 0.9815 | 0.9815 | 0.9815 | 0.96333 | 0.96333 | 0.96333
C2-13 | Impedance mismatch between feeder cable and CLTA | 0.14 | 0.23 | 0.25 | U-Shaped | 1.4142 | 1 | 0.098995 | 0.16263 | 0.17678 | 0.0098 | 0.02645 | 0.03125
A5-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
A5-20 | Reflections in anechoic chamber | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01 | 0.0001 | 0.0001 | 0.0001 | error not used before, replaced with quality of quiet zone.
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C3-7 | TX IMD - conducted measurement uncertainty | 1 | 1.1 | 1.2 | Gaussian | 1 | 1 | 1 | 1.1 | 1.2 | 1 | 1.21 | 1.44
Combined standard uncertainty (1σ) [dB] | [7 empty] | 1.6512 | 1.7185 | 1.7855 | 1.6512 | 1.7185 | 1.7855
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 3.2364 | 3.3682 | 3.4996 | 3.2364 | 3.3682 | 3.4996
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
|  | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
check 28/1/20
OK-same as 37.843

## COLO-BLK
Table 13.5.5-1: Common maximum accepted test system uncertainty value derivation for co-location blocking interferer signal level
| chamber | Expanded uncertainty [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
| General Chamber | 3.2423 | 3.3739 | 3.505
CATR
Near field
1D CATR
Plane wave sythesiser
Common maximum accepted test system uncertainty
Table 13.5.4.3-1: MU for co-location blocking interferer signal
General Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
C2-9 | Uncertainty related to the selection of the CLTA (Note) | 1.5 | 1.5 | 1.5 | Rectangular | 1.7321 | 1 | 0.86603 | 0.86603 | 0.86603 | 0.75 | 0.75 | 0.75
C2-10 | Uncertainty related to the placement of the CLTA (Note) | 1.7 | 1.7 | 1.7 | Rectangular | 1.7321 | 1 | 0.9815 | 0.9815 | 0.9815 | 0.96333 | 0.96333 | 0.96333
C2-13 | Impedance mismatch between feeder cable and CLTA | 0.14 | 0.23 | 0.25 | U-Shaped | 1.4142 | 1 | 0.098995 | 0.16263 | 0.17678 | 0.0098 | 0.02645 | 0.03125
A5-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
A5-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01 | error not used before, replaced with quality of quiet zone.
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C3-9 | Colocation blocking - conducted measurement uncertainty | 1 | 1.1 | 1.2 | Gaussian | 1 | 1 | 1 | 1.1 | 1.2 | 1 | 1.21 | 1.44
Combined standard uncertainty (1σ) [dB] | [7 empty] | 1.6542 | 1.7214 | 1.7883 | 1.6542 | 1.7214 | 1.7883
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 3.2423 | 3.3739 | 3.505 | 3.2423 | 3.3739 | 3.505
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
|  | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Final values come from the RSS calcaultion with the EIS MU