3GPP TSG RAN Meeting #89-e	 	R5-206192

Electronic Meeting, 9th  – 20th November 2020

Source:		Rohde & Schwarz

Title:	Estimate of Measurement Uncertainty for n259

Agenda Item:		5.3.2.17

Document for:	Discussion and Endorsement

# Introduction

At RAN5#88-e, there was a request in [1] to provide an MU estimation for band n259 [1].

In this document, we state our preliminary estimate for MU in band n259. We emphasize that the estimate is non-binding and not intended for MTSU discussion since for this a more detailed analysis is required.

# Discussion

While the upper frequency limit of band n259 is 43.5 GHz, the measurement uncertainty is required up to 44.3 GHz due to the out-of-band test case SEM, where a channel bandwidth of 400 MHz is assumed.

## Frequency error

Since the frequency MU is stated as a relative uncertainty in the specification (0.01 ppm), we think that it can be reused for band n259.

Observation 1: For Frequency error there is no relaxation due to testability issue. Applying usual TT approach is sufficient.

## ACLR

Due to the higher frequency the transmission between the center of quiet zone to the measurement device is degraded which leads to a higher impact of noise. The mismatch uncertainty will be increased since the VSWR is becoming worse with increasing frequency.

NR ACLR minimum conformance requirement for band n259 is identical to that of band n260. As a rigorous MU analysis of ACLR for band n259 could not be performed yet, in the meantime an oversimplified approach is to extrapolate the frequency dependence of the MU. Since the frequency dependency of the MU is complex and unknown, and since for the MU of a specific test case only two frequency points are available, we apply a linear extrapolation. Since the ordinate values (MU) are not given for discrete frequency values but for ranges some ambiguity in choosing the frequency values for the extrapolation exists. Since the MU is valid until the upper limit of each range we choose these as reference frequencies for the interpolation.

Since the extrapolation approach is very rough and unsafe, we add 0.5 dB to the values as a systematic uncertainty due to uncertain assumptions and limitations of extrapolation:

Table 1: MU for Tx test cases

| Test Case | MU (dB) |  | extrapolated MU (dB) | MU estimation including margin (dB) |
| --- | --- | --- | --- | --- |
|  | FR2a | FR2b | FR2c | FR2c |
| reference frequency (GHZ) | 31.25 | 40.8 | 44.3 |  |
| 6.5.2.3 Adjacent channel leakage ratio | 5.09 | 5.09 | 5.09 | 5.59 |

We note that ACLR has testability issues for frequencies smaller than or equal to 40.8 GHz and will have the same or more severe issues for band n259 due to the increased path loss and noise level.

# Conclusion

Proposal: Apply the MU values in Table 1 as basis solely for TT derivation and not to use them for MTSU

# References

- R5-204851, Update of WF on Band n259, NTT Docomo, Inc., 3GPP TSG RAN WG 5 Meeting #88-e, August 2020