3GPP TSG RAN Meeting #92-e	 	R5-215476

Electronic Meeting, 16th – 27th August 2021

Source:		Rohde & Schwarz

Title:	On the MU of Tx Modulation Quality test cases

Agenda Item:		5.3.2.17

Document for:	Discussion and Endorsement

# Introduction

At RAN5 #91-e the discussion of EVM MU has been started [1, 2]. However, no agreement could be achieved. In this updated resubmission of Ref. [1], we present the results of our analysis on the MU and testability for EVM.

# Discussion

## Measurement Uncertainty

The EVM test case for FR2 applies several modulations and wave forms with Outer_Full and Inner_Full allocations. We focus on the Outer_Full allocation since it has larger MPR than the Inner_Full allocation and represents the worst case with respect to SNR and intrinsic EVM of the test system.

The EVM test case applies UE maximum output power. Due to a possible mismatch of UE and TE polarization planes, the power at the TE for the polarization with higher power than the other one can be reduced by 3dB in the worst case.  In the analysis the minimum Peak EIRP for PC3, MBR, TT, MPR, T(MPR) and 3 dB due to possible polarization mismatch have been taken into account to calculate the minimum power level in the QZ for each CHBW and modulation. TT is considered in the power level calculation in order not to fail an UE in the EVM TC. The latter may happen if the EVM MU of the test system is determined for the higher power levels disregarding TT. In that case an UE that passes the MPR test case may fail the EVM TC not due to low modulation quality but because of its power level and the according higher TE EVM MU.

The intrinsic EVM of the test system has been determined by measuring the EVM of a modulated FR2 5G signal provided by a vector signal generator with the test system.

In Tables 1 and 2, we present our results for EVM MU for PC3 for the PUSCH test.

Table 1: EVM MU (%) for PC3, FR2a IFF 30cm Quiet Zone

|  | EVM MU (%) |  |  |  | Minimum Conformance requirement (%) | EVM increase for 400 MHz (%) | EVM increase / minimum conf. requ. (%) | Proposed TT for 400 MHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | CHBW |  |  |  |  |  |  |  |
| Modulation | 400 MHz | 200 MHz | 100 MHz | 50 MHz |  |  |  |  |
| Pi/2 BPSK DFT-s-OFDM | 2.2 | 1.8 | 1.7 | 1.7 | 30.0 | 0.08 | 0.27 | 0 |
| QPSK DFT-s-OFDM | 2.2 | 2.1 | 2.0 | 2.0 | 17.5 | 0.14 | 0.79 | 0 |
| 16 QAM DFT-s-OFDM | 3.2 | 2.3 | 2.0 | 2.0 | 12.5 | 0.40 | 3.2 | 0 |
| 64 QAM DFT-s-OFDM | 3.2 | 2.3 | 2.0 | 2.0 | 8.0 | 0.62 | 7.7 | 0.62 |
| QPSK CP-OFDM | 3.0 | 2.6 | 2.5 | 2.4 | 17.5 | 0.26 | 1.5 | 0 |
| 16 QAM CP-OFDM | 4.2 | 2.9 | 2.5 | 2.4 | 12.5 | 0.69 | 5.49 | 0 |
| 64QAM CP-OFDM | 4.2 | 2.9 | 2.5 | 2.4 | 8.0 | 1.04 | 12.9 | 1.0 |

Table 2: EVM MU (%) for PC3, FR2b IFF 30cm Quiet Zone

|  | EVM MU (%) |  |  |  | Minimum Conformance requirement (%) | EVM increase for 400 MHz (%) | EVM increase / minimum conf. requ. (%) | Proposed TT for 400 MHz | Proposed TT for 200 MHz | Proposed TT for 100 MHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | CHBW |  |  |  |  |  |  |  |  |  |
| Modulation | 400 MHz | 200 MHz | 100 MHz | 50 MHz |  |  |  |  |  |  |
| Pi/2 BPSK DFT-s-OFDM | 2.8 | 2.2 | 2.0 | 1.9 | 30.0 | 0.13 | 0.43 | 0.00 | 0.00 | 0.00 |
| QPSK DFT-s-OFDM | 3.0 | 2.3 | 2.0 | 1.9 | 17.5 | 0.26 | 1.46 | 0.00 | 0.00 | 0.00 |
| 16 QAM DFT-s-OFDM | 4.8 | 3.5 | 2.7 | 2.4 | 12.5 | 0.89 | 7.12 | 0.00 | 0.00 | 0.00 |
| 64 QAM DFT-s-OFDM | 4.8 | 3.5 | 2.7 | 2.4 | 8.0 | 1.33 | 16.62 | 1.33 | 0.73 | 0.00 |
| QPSK CP-OFDM | 4.3 | 2.9 | 2.4 | 2.2 | 17.5 | 0.52 | 2.97 | 0.00 | 0.00 | 0.00 |
| 16 QAM CP-OFDM | 6.4 | 4.3 | 3.2 | 2.7 | 12.5 | 1.54 | 12.35 | 0.12 | 0.00 | 0.00 |
| 64 QAM CP-OFDM | 6.4 | 4.3 | 3.2 | 2.7 | 8.0 | 2.24 | 28.06 | 2.24 | 1.08 | 0.62 |

We propose to use the values of Tables 1 and 2 as MU for EVM PUSCH and PUCCH PC3 for FR2a and FR2b IFF 30cm Quiet Zone, respectively.

Proposal 1: Apply the EVM MU values of Tables 1 and 2 as MU for EVM PUSCH and PUCCH PC3 for FR2a and FR2b IFF 30cm Quiet Zone, respectively.

## Testability and TT discussion

In the FR1 EVM TC, TT has been specified for 256QAM. The following tables summarize the test requirements, TT and MTSU for EVM in FR1:

Table 6.4.2.1.5-1: Test requirements for Error Vector Magnitude

| Parameter | Unit | Average EVM Level |
| --- | --- | --- |
| Pi/2-BPSK | % | 30 + TT |
| QPSK | % | 17.5 + TT |
| 16 QAM | % | 12.5 + TT |
| 64 QAM | % | 8 + TT |
| 256 QAM | % | 3.5 + TT |
| Note 1: TT is defined in Table 6.4.2.1.5-2. |  |  |

Table 6.4.2.1.5-2: Test Tolerance

| Parameter | Unit | Average EVM Level |
| --- | --- | --- |
| Pi/2-BPSK | % | 0 |
| QPSK | % | 0 |
| 16 QAM | % | 0 |
| 64 QAM | % | 0 |
| 256 QAM | % | 0.3 for 15 dBm < PUL 0.8 for -25 dBm < PUL≤ 15 dBm 1.1 for -40dBm ≤ PUL ≤ -25dBm |

Table F.1.2-1: Maximum Test System Uncertainty for transmitter tests

| Subclause | Maximum Test System Uncertainty | Derivation of Test System Uncertainty |
| --- | --- | --- |
| 6.4.2.1 Error Vector Magnitude | For up to 256QAM: f ≤ 6.0GHz, BW ≤ 100MHz 15 dBm < PUL PUSCH, PUCCH, PRACH: ±1.5 % -25 dBm < PUL ≤ 15 dBm PUSCH, PUCCH, PRACH: ±2.5 % -40dBm ≤ PUL ≤ -25dBm PUSCH, PUCCH, PRACH: ±3.0 % . |  |

Observation 1: TT has been applied in FR1 for 256 QAM, since the EVM MU is a significant fraction of the minimum conformance requirement.

We note that the value measured by the test equipment is the root sum square (rss) of the DUT EVM and the intrinsic TE EVM since the error vectors of UE and TE are statistically independent:

EVMmeasured=EVMDUT2+EVMTE2			(1)

In FR1, the TT has been calculated so that an UE complying with the minimum conformance requirement does not fail the test due to the intrinsic TE EVM:

TT=EVMconf.requ.2+EVMTE2-EVMconf.requ. ,	(2)

where EVMconf.requ. is the minimum conformance requirement and EVMTE is the intrinsic EVM of the test system.

Except for 256QAM there is no TT defined in FR1, i.e., for 8% EVM minimum conformance requirement the MU impact on the result is tolerated which is for low power values 0.54% which is a fraction of 6.8% of the minimum requirement. For 256 QAM and PUL>15 dB, the increase of measured EVM due to intrinsic TE EVM is 0.3% and it is compensated by TT of the same size. The fraction of the increase compare to the minimum conformance requirement is 8.6%. Summed up an increase of the measured EVM value relative to the conformance requirement of 6.8% has been tolerated whereas and increase by 8.6% has triggered definition of TT.

Observation 2: For FR1, an increase of the measured EVM value relative to the conformance requirement of 6.8% has been tolerated whereas an increase by 8.6% has triggered definition of TT.

In the following, we apply a similar condition as in FR1 for introducing TT: When the increase of the measured EVM due to intrinsic TE EVM is below 7.5% of the minimum conformance requirement, no TT is applied. In the other case TT is calculated via Eq. (2).

Table 1 and Table 2 summarize the TT. Modulation and CHBW combinations with green shading do not require any test tolerance (TT=0). For the others, proposed TT is stated in the same table.

We note that all test points are testable and that the proposed TT does not exceed the ratio of TT relative to the minimum conformance requirement in FR1.

Observation 3: All test points are testable. The proposed TT relative to the minimum conformance requirement does not exceed the values specified for FR1.

Proposal 2: Apply the TT values of Tables 1 and 2 as TT for EVM PUSCH and PUCCH PC3 for FR2a and FR2b IFF 30cm Quiet Zone, respectively.

Accompanying CRs [4] & [5] provide a documentation of the MU in TR 38.903 [6] and of the MTSU and TT in TS 38.521-2 [7], respectively.

# Assumptions

| ID | Description | Assumption |
| --- | --- | --- |
| #1 | Frequency ranges under consideration | All Rel-15 FR2 bands for in-band measurements |
| #2 | Size of QZ for IFF | 30 cm |
| #3 | UE power class | PC3 |
| #4 | Max EIRP | EIRP MAX= +43 dBm (rms) |
| #5 | Temperature range of the test equipment | 20°C – 35°C |

# Conclusion

In summary, EVM MU and testability have been discussed and MU and TT values have been proposed for FR2a, FR2b IFF with Quiet Zone size of 30cm.

Proposal 1: Apply the EVM MU values of Tables 1 and 2 as MU for EVM PUSCH and PUCCH PC3 for FR2a and FR2b IFF 30cm Quiet Zone, respectively.

Observation 1: TT has been applied in FR1 for 256 QAM, since the EVM MU is a significant fraction of the minimum conformance requirement.

Observation 2: For FR1, an increase of the measured EVM value relative to the conformance requirement of 6.8% has been tolerated whereas an increase by 8.6% has triggered definition of TT.

Observation 3: All test points are testable. The proposed TT relative to the minimum conformance requirement does not exceed the values specified for FR1.

Proposal 2: Apply the TT values of Tables 1 and 2 as TT for EVM PUSCH and PUCCH PC3 for FR2a and FR2b IFF 30cm Quiet Zone, respectively.

# References

- R5-213310, On the MU of FR2 EVM, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #91-e, May 2021

- R5-213817, On FR2 EVM measurement uncertainties and test tolerances, Keysight Technologies, 3GPP TSG RAN WG5 Meeting #91-e, May 2021

- R5-202852, Discussion on the power level for transmit modulation quality test cases, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #87-e, May 2020

- R5-215477, Update of transmit modulation quality test cases, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #92-e, August 2021

- R5-215478, MU for Tx modulation test cases, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #92-e, August 2021

- TR 38.903, V16.8.0, 3GPP TSG RAN WG5, June 2021

- TS 38.521-2, V16.8.0, 3GPP TSG RAN WG5, June 2021