3GPP TSG-RAN5 Meeting #92-e	R5-214848

# Online, , 16th Aug 2021 - 27th Aug 2021

| CR-Form-v12.1 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.903 | CR | 0244 | rev | - | Current version: | 16.8.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | Introducing EIRP UL Absolute Power MU for FR2 RRM |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Anritsu |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | 5GS_NR_LTE-UEConTest |  |  |  |  |  | Date: |  |  | 2021-08-06 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-16 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-15 (Release 15) Rel-16 (Release 16) Rel-17 (Release 17) Rel-18 (Release 18) |  |
| Reason for change: |  | UL Absolute Power MU for FR2 RRM is not yet specified. |  |  |  |  |  |  |  |  |
| Summary of change: |  | UL Absolute Power MU for FR2 RRM test is specified for DFF, IFF, enhanced IFF and IFF+DFF. Impact of noise is set as TBD at this moment as it could depends on the UL power level to be measured in the test case which needs further study. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | UL Absolute Power MU for FR2 RRM test case is kept missing. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | E.3 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  |  | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  |  | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  |  | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  |  |  |  |  |  |  |  |  |  |
| This CR's revision history: |  |  |  |  |  |  |  |  |  |  |

<<Unchaged sections skipped>>

E.3	Uncertainty assessment for RRM MU quantities.

RRM measurement uncertainty analysis shall define the values for the following MU quantities:

-	DL AWGN absolute power or wanted DL signal absolute power

-	DL applied SNR

-	DL Fading profile uncertainty

-	DL AWGN and signal flatness

-	UL absolute power measurement

-	UL relative power measurement

-	UL signal transmit timing relative to DL

-	Relative transmit timing accuracy during UE timing adjustment

E.3.1	Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power

Table E.3.1-1 summarizes the MU threshold for DL AWGN absolute power for RRM FR2 test cases. The origin MU values for different test setups with varies parameters can be found in following subclauses.

Table E.3.1-1: MU threshold for DL AWGN absolute power for RRM FR2

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | As configured in the test case | 5.65 dB2 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.65 dB2 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | As configured in the test case | FFS |
|  | 32.125GHz < f <= 40.8GHz |  |  | FFS |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table E.3.1.3-2 for PC3 UEs and Table FFS for PC1 UEs NOTE 2: If the TT analysis for a specific test case based on this MU value results in an unsolvable conflict, making the test case untestable, even after the alternative solutions listed in clause A.4 have been considered for the test case in TS 38.133 [6] Annex A, the TT analysis shall be repeated using a lower MU value, taking into account the lower values defined in this clause. The test case will be applicable for the subset of the test systems meeting this reduced MU Threshold. |  |  |  |  |

The types of test setup are defined in clause 7.1.3.2 of TS 38.508-1 [18]

E.3.1.1	Uncertainty budget format and assessment for DFF test setup

The uncertainty contributions that may impact the overall MU value are listed in Table E.3.1.1-1.

Table E.3.1.1-1: Uncertainty contributions for DL AWGN absolute power or wanted DL signal absolute power

| UID |  | Description of uncertainty contribution |  | Details in annex |  |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 |  | Positioning misalignment |  | B.2.1.1 |  |
| 2 |  | Measure distance uncertainty |  | B.2.1.2 |  |
| 3 |  | Quality of Quiet Zone |  | B.2.1.3 |  |
| 4 |  | Mismatch |  | B.2.1.4 |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  | B.2.1.5 |  |
| 6 |  | gNB emulator uncertainty |  | B.2.1.17 |  |
| 7 |  | Phase curvature |  | B.2.1.7 |  |
| 8 |  | Amplifier uncertainties |  | B.2.1.8 |  |
| 9 |  | Random uncertainty |  | B.2.1.9 |  |
| 10 |  | Influence of the XPD |  | B.2.1.10 |  |
| 11 |  | Insertion Loss Variation |  | B.2.1.11 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  | B.2.1.12 |  |
| 13 |  | Multiple measurement antenna uncertainty |  | B.2.1.25 |  |
| 14 |  | DUT repositioning |  | B.2.1.26 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 |  | Mismatch |  | B.2.1.4 |  |
| 16 |  | Amplifier Uncertainties |  | B.2.1.8 |  |
| 17 |  | Misalignment of positioning System |  | B.2.1.13 |  |
| 18 |  | Uncertainty of the Network Analyzer |  | B.2.1.14 |  |
| 19 |  | Uncertainty of the absolute gain of the calibration antenna |  | B.2.1.15 |  |
| 20 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  | B.2.1.16 |  |
| 21 |  | Phase centre offset of calibration antenna |  | B.2.1.18 |  |
| 22 |  | Quality of quiet zone for calibration process |  | B.2.1.19 |  |
| 23 |  | Standing wave between reference calibration antenna and measurement antenna |  | B.2.1.20 |  |
| 24 |  | Influence of the calibration antenna feed cable |  | B.2.1.21 |  |
| 25 |  | Insertion Loss Variation |  | B.2.1.11 |  |
| Systematic uncertainties |  |  |  |  |  |
| 26 |  | Systematic error related to beam peak search |  | B.2.1.28 |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz}.

-	The uncertainty assessment is applicable for 1AoA and 2AoA test cases

-	The uncertainty assessment is provided in Table E.3.1.1-2.

Table E.3.1.1-2: Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.15 | Rectangular | 1.73 | 0.08 |
| 3 | Quality of Quiet Zone (NOTE 4) | 1.2 | Actual | 1.00 | 1.2 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.06 | U-shaped | 1.41 | 0.043 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 16 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 17 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 19 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 20 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 21 | Phase centre offset of calibration antenna | 0.47 | Rectangular | 1.73 | 0.27 |
| 22 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 23 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 24 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 25 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 26 | Systematic error related to beam peak search |  |  |  | 0.5 |
| Total measurement uncertainty |  |  |  |  | Value |
| DL AWGN absolute power or wanted DL signal absolute power expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.65 |
| NOTE 1: The analysis was done only for the case of operating in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in Annex D.2 of TR 38.810 [13] for Quiet Zone size less or equal to 30 cm. NOTE 5: The values in this table have been derived for DL powers above and equal to REFSENS. The values might need to be revisited for power levels below REFSENS |  |  |  |  |  |

E.3.1.2	Uncertainty budget format and assessment for Simplified DFF test setup

[FFS]

E.3.1.3	Uncertainty budget format and assessment for IFF test setup

The uncertainty contributions that may impact the overall MU value are listed in Table E.3.1.3-1.

Table E.3.1.3-1: Uncertainty contributions for DL AWGN absolute power or wanted DL signal absolute power

| UID |  | Description of uncertainty contribution |  | Details in annex |  |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 |  | Positioning misalignment |  | B.2.2.1 |  |
| 2 |  | Measure distance uncertainty |  | B.2.2.2 |  |
| 3 |  | Quality of Quiet Zone |  | B.2.2.3 |  |
| 4 |  | Mismatch |  | B.2.2.4 |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  | B.2.2.5 |  |
| 6 |  | gNB emulator uncertainty |  | B.2.2.17 |  |
| 7 |  | Phase curvature |  | B.2.2.7 |  |
| 8 |  | Amplifier uncertainties |  | B.2.2.8 |  |
| 9 |  | Random uncertainty |  | B.2.2.9 |  |
| 10 |  | Influence of the XPD |  | B.2.2.10 |  |
| 11 |  | Insertion Loss Variation |  | B.2.2.11 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  | B.2.2.12 |  |
| 13 |  | Multiple measurement antenna uncertainty |  | B.2.2.25 |  |
| 14 |  | DUT repositioning |  | B.2.2.26 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 |  | Mismatch |  | B.2.2.4 |  |
| 16 |  | Amplifier Uncertainties |  | B.2.2.8 |  |
| 17 |  | Misalignment of positioning System |  | B.2.2.13 |  |
| 18 |  | Uncertainty of the Network Analyzer |  | B.2.2.14 |  |
| 19 |  | Uncertainty of the absolute gain of the calibration antenna |  | B.2.2.15 |  |
| 20 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  | B.2.2.16 |  |
| 21 |  | Phase centre offset of calibration antenna |  | B.2.2.18 |  |
| 22 |  | Quality of quiet zone for calibration process |  | B.2.2.19 |  |
| 23 |  | Standing wave between reference calibration antenna and measurement antenna |  | B.2.2.20 |  |
| 24 |  | Influence of the calibration antenna feed cable |  | B.2.2.21 |  |
| 25 |  | Insertion Loss Variation |  | B.2.2.11 |  |
| Systematic uncertainties |  |  |  |  |  |
| 26 |  | Systematic error related to beam peak search |  | B.2.2.28 |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz}.

-	The uncertainty assessment is applicable for 1AoA test cases-	The uncertainty assessment is provided in Table E.3.1.3-2.

Table E.3.1.3-2: Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 16 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 17 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 19 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 20 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 21 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 22 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 23 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 24 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 25 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 26 | Systematic error related to beam peak search |  |  |  | 0.5 |
| Total measurement uncertainty |  |  |  |  | Value |
| DL AWGN absolute power expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.19 |
| NOTE 1: The analysis was done only for the case of operating in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in Annex D.2 of TR 38.810 [13] for Quiet Zone size less or equal to 30 cm. NOTE 5: The values in this table have been derived for DL powers above and equal to REFSENS. The values might need to be revisited for power levels below REFSENS |  |  |  |  |  |

E.3.1.4	Uncertainty budget format and assessment for Enhanced IFF test setup

The uncertainty contributions that may impact the overall MU value are listed in Table E.3.1.4-1.

Table E.3.1.4-1: Uncertainty contributions for DL AWGN absolute power or wanted DL signal absolute power

| UID |  | Description of uncertainty contribution |  | Details in annex |  |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 to 14 |  | See 1-14 of Table E.3.1.3-1 |  | N/A |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 to 25 |  | See 15-25 of Table E.3.1.3-1 |  | N/A |  |
| Systematic uncertainties |  |  |  |  |  |
| 26 |  | See 26 of Table E.3.1.3-1 |  | N/A |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz}.

-	The uncertainty assessment is applicable for 1AoA and 2AoA test cases

-	The uncertainty assessment is provided in Table E.3.1.4-2.

Table E.3.1.4-2: Uncertainty assessment for fDL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 28 | Systematic error related to beam peak search |  |  |  | 0.5 |
| Total measurement uncertainty |  |  |  |  | Value |
| DL AWGN absolute power expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.25 |
| NOTE 1: The analysis was done only for the case of operating in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in Annex M of TR 38.521-2 [7] for Quiet Zone size less or equal to 30 cm. NOTE 5: The values in this table have been derived for DL powers above and equal to REFSENS. The values might need to be revisited for power levels below REFSENS |  |  |  |  |  |

E.3.1.5	Uncertainty budget format and assessment for IFF+DFF Hybrid test setup

For DFF probe, Uncertainty shall be evaluated using the Uncertainty budget format as specified in E.3.1.1.

For IFF probe, Uncertainty shall be evaluated using the Uncertainty budget format as specified in E.3.1.3.

The overall uncertainty of the IFF+DFF Hybrid test set up shall be calculated with the max(Total DFF probe MU, Total IFF probe MU).

E.3.2	Uncertainty assessment for DL applied SNR

E.3.3	Uncertainty assessment for DL Fading profile uncertainty

E.3.4	Uncertainty assessment for DL AWGN and signal flatness

E.3.5	Uncertainty assessment for UL absolute power measurement

Following tables summarize the MU threshold for EIRP UL absolute power measurement in FR2 RRM test cases. The origin MU values for different test setups with varies parameters can be found in following clauses.

Table B.3.5-1: MU threshold for EIRP UL absolute power measurement

TBD

E.3.5.1	Uncertainty budget format and assessment for DFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.3.5.1-1.

Table B.3.5.1-1: Uncertainty contributions for EIRP UL absolute power measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.1.1 |
| 2 | Measure distance uncertainty | B.2.1.2 |
| 3 | Quality of quiet zone | B.2.1.3 |
| 4 | Mismatch | B.2.1.4 |
| 5 | Standing Wave Between the DUT and measurement antenna | B.2.1.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.1.6 |
| 7 | Phase curvature | B.2.1.7 |
| 8 | Amplifier uncertainties | B.2.1.8 |
| 9 | Random uncertainty | B.2.1.9 |
| 10 | Influence of the XPD | B.2.1.10 |
| 11 | Insertion Loss Variation | B.2.1.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.1.12 |
| 13 | Influence of beam peak search grid | B.2.1.23 |
| 14 | Multiple measurement antenna uncertainty | B.2.1.25 |
| 15 | DUT repositioning | B.2.1.26 |
| Stage 1: Calibration measurement |  |  |
| 16 | Mismatch | B.2.1.4 |
| 17 | Amplifier uncertainties | B.2.1.8 |
| 18 | Misalignment of positioning System | B.2.1.13 |
| 19 | Uncertainty of the Network Analyzer | B.2.1.14 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | B.2.1.15 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.1.16 |
| 22 | Phase centre offset of calibration antenna | B.2.1.18 |
| 23 | Quality of quiet zone for calibration process | B.2.1.19 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | B.2.1.20 |
| 25 | Influence of the calibration antenna feed cable | B.2.1.21 |
| 26 | Insertion Loss Variation | B.2.1.11 |
| Systematic uncertainties |  |  |
| 27 | Influence of noise | B.2.1.27 |
| 28 | Systematic error related to beam peak search | B.2.1.28 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz}.

-	The uncertainty assessment is applicable for 1AoA and 2AoA test cases

-	The uncertainty assessment is provided in Table B.3.5.1-2.

Table B.3.5.3-2: Uncertainty assessment for EIRP measurement  (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID |  |  | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 |  |  | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |  |  |
| 2 |  |  | Measure distance uncertainty | 0.15 | Rectangular | 1.73 | 0.08 |  |  |
| 3 |  |  | Quality of quiet zone (NOTE 1) | 1.2 | Actual | 1.00 | 1.2 |  |  |
| 4 |  |  | Mismatch | 1.30 | Actual | 1.00 | 1.30 |  |  |
| 5 |  |  | Standing Wave Between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |  |  |
| 6 |  |  | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.50 | Normal | 2.00 | 1.25 |  |  |
| 7 |  |  | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |  |  |
| 8 |  |  | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |  |  |
| 9 |  |  | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |  |  |
| 10 |  |  | Influence of the XPD | 0.06 | U-shaped | 1.41 | 0.043 |  |  |
| 11 |  |  | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |  |  |
| 12 |  |  | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |  |  |
| 13 |  |  | Influence of beam peak search grid | 0.00 | Actual | 1 | 0.00 |  |  |
| 14 |  |  | Multiple measurement antenna uncertainty(NOTE 4) | 0.15 | Actual | 1 | 0.15 |  |  |
| 15 |  |  | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |
| 16 |  |  | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |  |  |
| 17 |  |  | Amplifier uncertainties | 0.00 | Normal | 2.00 | 0.00 |  |  |
| 18 |  |  | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |  |  |
| 19 |  |  | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |  |  |
| 20 |  |  | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |  |  |
| 21 |  |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |  |  |
| 22 |  |  | Phase centre offset of calibration antenna | 0.47 | Rectangular | 1.73 | 0.27 |  |  |
| 23 |  |  | Quality of quiet zone for calibration process (NOTE 1) | 0.4 | Actual | 1.00 | 0.4 |  |  |
| 24 |  |  | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |  |  |
| 25 |  |  | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |  |  |
| 26 |  |  | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |  |  |
|  |  | Systematic uncertainties (NOTE 3) |  |  |  |  |  | Value |  |
| 27 |  | Influence of noise |  |  |  |  |  | TBD |  |
| 28 |  | Systematic error related to beam peak search |  |  |  |  |  | 0.5 |  |
| Total measurement uncertainty |  |  |  |  |  |  |  | Value |  |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | TBD |  |
| NOTE 1: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The assessment assumes minimum output power level. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |  |  |  |  |

E.3.5.2	TBD

E.3.5.3	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.3.5.3-1.

Table E.3.5.3-1: Uncertainty contributions for EIRP UL absolute power measurement

| UID | Description of uncertainty contribution | Details in clause |  |
| --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.2.1 |  |
| 2 | Measure distance uncertainty | B.2.2.2 |  |
| 3 | Quality of Quiet Zone | B.2.2.3 |  |
| 4 | Mismatch | B.2.2.4 |  |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |  |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |  |
| 7 | Phase curvature | B.2.2.7 |  |
| 8 | Amplifier uncertainties | B.2.2.8 |  |
| 9 | Random uncertainty | B.2.2.9 |  |
| 10 | Influence of the XPD | B.2.2.10 |  |
| 11 | Insertion Loss Variation | B.2.2.11 |  |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |  |
| 13 | Influence of beam peak search grid | B.2.2.23 |  |
| 14 | Multiple measurement antenna uncertainty | B.2.2.25 |  |
| 15 | DUT repositioning | B.2.2.26 |  |
| Stage 1: Calibration measurement |  |  |
| 16 | Mismatch | B.2.2.4 |  |
| 17 | Amplifier Uncertainties | B.2.2.8 |  |
| 18 | Misalignment of positioning System | B.2.2.13 |  |
| 19 | Uncertainty of the Network Analyzer | B.2.2.14 |  |
| 20 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |  |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |  |
| 22 | Phase centre offset of calibration antenna | B.2.2.18 |  |
| 23 | Quality of quiet zone for calibration process | B.2.2.19 |  |
| 24 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |  |
| 25 | Influence of the calibration antenna feed cable | B.2.2.21 |  |
| 26 | Insertion Loss Variation | B.2.2.11 |  |
| Systematic uncertainties |  |  |
| 27 | Influence of noise | B.2.2.27 |  |
| 28 | Systematic error related to beam peak search | B.2.2.28 |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz}.

-	The uncertainty assessment is applicable for 1AoA test cases

-	The uncertainty assessment for EIRP is provided in Table B.3.5.3-2 for PC3 UEs

Table B.3.5.3-2: Uncertainty assessment for EIRP UL absolute power measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.50 | Normal | 2.00 | 1.25 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of beam peak search grid | 0.00 | Actual | 1 | 0.00 |
| 14 | Multiple measurement antenna uncertainty (NOTE 4) | 0.15 | Actual | 1 | 0.15 |
| 15 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 1) | 0.4 | Actual | 1.00 | 0.4 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 27 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | TBD |
| 28 | Systematic error related to beam peak search |  |  |  | 0.5 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | TBD |
| NOTE 1: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The assessment assumes minimum output power level. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

E.3.5.4	Uncertainty budget format and assessment for Enhanced IFF test setup

The uncertainty contributions that may impact the overall MU value are listed in Table E.3.5.4-1.

Table E.3.5.4-1: Uncertainty contributions for EIRP UL absolute power measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |
| 1 to 15 | See 1-15 of Table E.3.5.3-1 | N/A |
| Stage 1: Calibration measurement |  |  |  |  |
| 16 to 26 | See 16-26 of Table E.3.5.3-1 | N/A |
| Systematic uncertainties |  |  |  |  |
| 27 to 28 | See 27-28 of Table E.3.5.3-1 | N/A |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz}.

-	The uncertainty assessment is applicable for 1AoA and 2AoA test cases

-	The uncertainty assessment is provided in Table E.3.5.4-2.

Table B.3.5.4-2: Uncertainty assessment for EIRP UL absolute power measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.50 | Normal | 2.00 | 1.25 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of beam peak search grid | 0.00 | Actual | 1 | 0.00 |
| 14 | Multiple measurement antenna uncertainty (NOTE 4) | 0.15 | Actual | 1 | 0.15 |
| 15 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 1) | 0.4 | Actual | 1.00 | 0.4 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  |  | Systematic uncertainties (NOTE 3) |  |  |  |  | Value |  |
| 27 |  | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  |  | TBD |  |
| 28 |  | Systematic error related to beam peak search |  |  |  |  | 0.5 |  |
| Total measurement uncertainty |  |  |  |  |  |  | Value |  |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | TBD |  |
| NOTE 1: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The assessment assumes minimum output power level. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |  |  |  |

E.3.5.5	Uncertainty budget format and assessment for IFF+DFF Hybrid test setup

For DFF probe, Uncertainty shall be evaluated using the Uncertainty budget format as specified in E.3.5.1.

For IFF probe, Uncertainty shall be evaluated using the Uncertainty budget format as specified in E.3.5.3.

The overall uncertainty of the IFF+DFF Hybrid test set up shall be calculated with the max(Total DFF probe MU, Total IFF probe MU).

E.3.6	Uncertainty assessment for UL relative power measurement

E.3.7	Uncertainty assessment for UL signal transmit timing relative to DL

E.3.8	Uncertainty assessment for Relative transmit timing accuracy during UE timing adjustment

Annex F: Change history

<<End of change>>