3GPP TSG-RAN5 Meeting #103	R5-243104

# Fukuoka City, Fukuoka, Japan, 20th May 2024 - 24th May 2024

| CR-Form-v12.3 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.903 | CR | 0745 | rev | - | Current version: | 18.2.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | Documentation of MU for UL MIMO |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | ROHDE & SCHWARZ |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | TEI15_Test, 5GS_NR_LTE-UEConTest |  |  |  |  |  | Date: |  |  | 2024-05-10 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-18 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-17 (Release 17) Rel-18 (Release 18) Rel-19 (Release 19) Rel-20 (Release 20) |  |
| Reason for change: |  | The MU value for the UL MIMO spurious emission test cases is not specified. |  |  |  |  |  |  |  |  |
| Summary of change: |  | The MU value for the UL MIMO spurious emission test cases has been defined. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | The UL MIMO spurious emission test cases will remain incomplete. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | B.2.2.27, B.18, B18.2 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  | X |  | Test specifications |  |  |  | TS/TR 38.521-2 CR 1052 TS/TR 38.521-3 CR 1784 |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | Discussion in R5-243103, SA test case update in R5-243105, NSA test case update in CR R5-243106 |  |  |  |  |  |  |  |  |
| This CR's revision history: |  |  |  |  |  |  |  |  |  |  |

< Unchanged sections omitted >

### B.2.2.27	Influence of noise

See B.2.1.27.

The uncertainty value of influence of noise is estimated as below table and used across clause B.

Editor’s Note: For ACLR, all applicable configurations need to be added.

Table B.2.2.27-1: Uncertainty value for influence of noise for PC3, SISO for IFF

| Test case | Frequency range | Noise floor | Minimum requirement | Estimated SNRtotal [dB/400MHz] | Relaxation | Influence of noise |
| --- | --- | --- | --- | --- | --- | --- |
| MOP-EIRP | FR2a | N/A | 20.7dBm/ChBW (22.4-1.7) | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b | N/A | 18.9dBm/ChBW (20.6-1.7) | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2c | N/A | 18.2dBm/ChBW (18.7-0.5) | 11.45 (NOTE 1) | 0 | 0.3 |
| MOP-TRP | FR2a | N/A | 23dBm/ChBW | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b, FR2c | N/A | 23dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
| MOP-Spherical | FR2a | N/A | 9.75dBm/ChBW (Spherical – MBR= 11.5-1.75) | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2b | N/A | 7.6dBm/ChBW (Spherical – MBR=8-0.4) | 6.37 (NOTE 1) | 0 | 0.9 |
|  | FR2c | N/A | 5.4dBm/ChBW (Spherical – MBR=5.8-0.4) | 5.86 (NOTE 1) | 0 | 1.0 |
| MPR | FR2a | -7.6dBm/400MHz | 7.65dBm/ChBW (EIRP-MPB-MPR-T(MPR)=22.4-0.75-9-5) | 15.17 (NOTE 1) | 0 | 0.13 |
|  | FR2b | -5.5dBm/400MHz | 5.85dBm/ChBW (EIRP-MPB-MPR-T(MPR)=20.6-0.75-9-5) | 11.30 (NOTE 1) | 0 | 0.31 |
|  | FR2c | [-4.5]dBm/400MHz | 4.2dBm/ChBW (EIRP-MPB-MPR-T(MPR)=18.7-0.5-9-5) | [8.70] (NOTE 1) | [0] | [0.55] |
| Configured output power with power boost | FR2a | N/A | 21.7dBm/ChBW (22.4-1.7+1) | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b | N/A | 19.9dBm/ChBW (20.6-1.7+1) | 11.45 (NOTE 1) | 0 | 0.3 |
| Minimum output power | FR2a | -10.6dBm/400MHz | -13dBm | -2.54 (NOTE 1) | 8.4 | 1.0 (with relaxation) |
|  | FR2b | -5.5dBm/400MHz | -13dBm | -7.64 (NOTE 1) | 13.5 | 1.0 (with relaxation) |
|  | FR2c | [-4.5]dBm/400MHz | -13dBm | [-8.64] (NOTE 1) | [14.5] | [1.0] (with relaxation) |
| OFF power – TRP | FR2a | N/A | -35dBm/ChBW | -24.54 (NOTE 2) | 30.4 | 1.0 (with relaxation) |
|  | FR2b | N/A |  | <-24.54 (NOTE 2) | N/A | Propose not to test |
| OFF power – EIRP | FR2a | -7.6dBm/400MHz | -30dBm/ChBW | -22.54 (NOTE 2) | 28.4 | 1.0 (with relaxation) |
|  | FR2b | -5.5dBm/400MHz |  | -24.64 (NOTE 2) | 30.5 | 1.0 (with relaxation) |
| Absolute power tolerance | Same as Minimum output power |  |  |  |  |  |
| Relative power tolerance | FR2a | -13.6dBm/100MHz | -7.6dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
|  | FR2b | -11.5dBm/100MHz | -5.5dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
| Aggregate power tolerance | Same as Relative power tolerance |  |  |  |  |  |
| Aggregate power tolerance | FR2a | -13.6dBm/100MHz | -7.6dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
|  | FR2b | -11.5dBm/100MHz | -5.5dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
| SEM | FR2a | N/A | -13dBm/1MHz | 8.14 (NOTE 1) | 0 | 0.62 |
|  | FR2b, FR2c | N/A |  | 5.86 (NOTE 1) | 0 | 1.0 |
| ACLR (CP) | FR2a | -7.6dBm/400MHz | Highest testable MPR for 400MHz: 3dB 16.65dBm/ChBW (EIRP-MPB-MPR-T(MPR) =22.4-0.75-3-2) Actual lowest: 7.65dBm/ChBW (EIRP-MPB-MPR-T(MPR)=22.4-0.75-9-5) | 22.86 (with 3dB MPR) (NOTE 1) | 0 | N/A |
|  | FR2b | -5.5dBm/400MHz | Highest testable MPR for 400MHz: 2dB 16.35dBm/ChBW (EIRP-MPB-MPR-T(MPR) =20.6-0.75-2-1.5) Actual lowest: 5.85dBm/ChBW (EIRP-MPB-MPR-T(MPR)=20.6-0.75-9-5) | 21.86 (with 2dB MPR) (NOTE 1) | 0 | N/A |
| ACLR (ACP) | FR2a | -7.6dBm/400MHz | Highest testable MPR for 400MHz: 3dB -0.35dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=22.4-0.75-3-2-17) Actual lowest: -9.35 dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=22.4-0.75-9-5-17) | 5.86 (NOTE 1) | 0 | 1.0 |
|  | FR2b | -5.5dBm/400MHz | Highest testable MPR for 400MHz: 2dB 0.35dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=20.6-0.75-2-1.5-16) Actual lowest: -10.15 dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=20.6-0.75-9-5-16) | 5.86 (with 2dB MPR) (NOTE 1) | 0 | 1.0 |
| General Tx spurious | 6GHz <=f<=23.45GHz | N/A | -13dBm/1MHz | 10.0 (NOTE 1) | 0 | 0.41 |
|  | 23.45GHz<=f<=40GHz | N/A | -13dBm/1MHz | 10.0 (NOTE 1) | 0 |  |
|  | 40GHz<=f<=80GHz | N/A | -13dBm/1MHz | 10.0 (NOTE 1) | 0 |  |
|  | 80GHz<f<=87GHz | N/A | -13dBm/1MHz | 6.8 (NOTE 1) | 0 | 0.83 |
| Tx spurious Co-existence | n260 (Aggressor band : n257, n261) | -23 | -2dBm/100MHz (-22dBm/MHz) | 0.86 (NOTE 2) | 5 | 1.0 (with relaxation) |
|  | n257, n261 (Aggressor band : n260) | -27.7 | -5dBm/100MHz (-25dBm/MHz) | 2.56 (NOTE 2) | 3.3 | 1.0 (with relaxation) |
|  | 23.6 GHz ≤ f ≤ 24.0GHz | -27.7 | 1dBm/200MHz (-22 dBm/MHz) | 5.56 (NOTE 2) | 0.3 | 1.0 (with relaxation) |
|  | 36 GHz ≤ f ≤ 37GHz | -23dBm/MHz | 7dBm/1000MHz (-23dBm/MHz) | -0.14 (NOTE 2) | 6 | 1.0 (with relaxation) |
|  | 57 GHz ≤ f ≤ 66GHz | N/A | 2dBm/100MHz (-18dBm/MHz) | 5.86 (NOTE 1) | 0 | 1.0 |
| Additional spurious emission | NS_202 (7.25GHz <=f <=12.75GHz) | -40 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | 10 (NOTE 1) | 0 | 0.41 |
|  | NS_202 (12.75GHz <=f <=23.45GHz) | -23 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | -7.14 (NOTE 2) | 13 | 1.0 (with relaxation) |
|  | NS_202 (23.6GHz <=f <=24.0GHz) | -27.7 dBm/MHz | 1dBm/200MHz (-22 dBm/MHz) | 5.56 (NOTE 2) | 0.3 | 1.0 (with relaxation) |
|  | NS_202 (23.45GHz <=f <=40.8GHz) | -23 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | -7.14 (NOTE 2) | 13 | 1.0 (with relaxation) |
|  | NS_202 (40.8GHz <=f <=66GHz) | -23 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | -7.14 (NOTE 2) | 13 | 1.0 (with relaxation) |
|  | NS_203 (23.6GHz <=f <=24.0GHz) | -27.7 dBm/MHz | +1dBm/200MHz (-22dBm/MHz) | 5.56 (NOTE 2) | 0.3 | 1.0 (with relaxation) |
| Rx spurious | 6GHz <=f<=20GHz |  | -47dBm/1MHz | -2.2 (NOTE 3) | 10.2 | 1.0 dB for 23.45~40.8GHz and 80~87 GHz, 0.64dB for 6~23.45 and 40.8~80 GHz. |
|  | 20GHz<=f<=40GHz |  | -47dBm/1MHz | -11.34 (NOTE 2) | 17.2 |  |
|  | 40GHz<=f<=80GHz |  | -47dBm/1MHz | -25.1 (NOTE 3) | 33.1 |  |
|  | 80GHz<f<=87GHz |  | -47dBm/1MHz | -27.1 (NOTE 2) | 33.1 |  |
| NOTE 1: Estimated SNR is calculated based on agreed influence of noise. NOTE 2: Estimated SNR is calculated based on agreed relaxation value: Estimated SNR = 6dB - relaxation. NOTE 3: Estimated SNR is calculated based on agreed relaxation value: Estimated SNR = 8dB - relaxation. |  |  |  |  |  |  |

Table B.2.2.27-1a: Uncertainty value for influence of noise for PC3, MIMO for IFF

| Test case | Frequency range | Noise floor | Minimum requirement | Estimated SNRtotal [dB/400MHz] | Relaxation | Influence of noise |
| --- | --- | --- | --- | --- | --- | --- |
| MOP-EIRP | FR2a | N/A | 22.4 dBm/ChBW | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b | N/A | 20.6 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2c | N/A | 18.7 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
| MOP-TRP | FR2a | N/A | 23 dBm/ChBW | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b, FR2c | N/A | 23 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
| MOP-Spherical | FR2a | N/A | 11.5 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2b | N/A | 8 dBm/ChBW | 6.37 (NOTE 1) | 0 | 0.9 |
| Minimum output power | FR2a | -10.6dBm/400MHz | -10 dBm | -0.46 (NOTE 1) | 5.4 | 1.0 (with relaxation) |
|  | FR2b | -5.5dBm/400MHz | -10 dBm | -4.64 (NOTE 1) | 10.5 | 1.0 (with relaxation) |
| OFF power – TRP | FR2a | N/A | -35dBm/ChBW | -24.54 (NOTE 2) | 30.4 | 1.0 (with relaxation) |
|  | FR2b | N/A |  | <-24.54 (NOTE 2) | N/A | Propose not to test |
| General Tx spurious | Same as SISO in Table B.2.2.27-1 |  |  |  |  |  |
| Tx spurious Co-existence | Same as SISO in Table B.2.2.27-1 |  |  |  |  |  |
| Additional spurious emission | Same as SISO in Table B.2.2.27-1 |  |  |  |  |  |
| NOTE 1: Estimated SNR is calculated based on agreed influence of noise. |  |  |  |  |  |  |

Table B.2.2.27-2: Uncertainty value for influence of noise for PC1, SISO for IFF

| Test case | Frequency range | Relaxation [dB] | Influence of noise [dB] |
| --- | --- | --- | --- |
| MOP-EIRP | FR2a | 0 | 0.13 |
|  | FR2b | 0 | 0.2 |
| MOP-TRP | FR2a | 0 | 0.13 |
|  | FR2b | 0 | 0.27 |
| MOP-Spherical | FR2a | 0 | 0.2 |
|  | FR2b | 0 | 0.35 |
| MPR | FR2a | 0 | 0.13 |
|  | FR2b | 0 | 0.3 |
| Minimum output power | FR2a | 0 | 0.3 |
|  | FR2b | 0 | 0.6 |
| OFF power – TRP | FR2a | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | FR2b | Same as defined for PC3 in Table B.2.2.27-1 |  |
| SEM | FR2a | 0 | 1.81 (NOTE 1, 2) |
|  | FR2b | 0 | TBD |
| ACLR (ACP) | FR2a | 0 | 0.95 |
|  | FR2b | 0 | 0.95 |
| General Tx spurious | 6GHz <=f<12.75GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 12.75GHz<=f<66GHz | 0 | 1.08 (NOTE 1, 3) |
|  | 66GHz<=f<=80GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
| Tx spurious Co-existence | n260 (Aggressor band : n257, n261) | 5 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | n257, n261 (Aggressor band : n260) | 3.3 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | 23.6 GHz ≤ f ≤ 24.0GHz | 0.3 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | 57 GHz ≤ f ≤ 66GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
| Additional spurious emission | NS_202 (7.25GHz <=f <=12.75GHz) | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | NS_202 (12.75GHz <=f <=23.45GHz) | 13 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | NS_202, NS_203 (23.6 GHz ≤ f ≤ 24.0GHz) | 0.3 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | NS_202 (23.45GHz <=f <=40.8GHz) | 13 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | NS_202 (40.8GHz <=f <=66GHz) | 13 (Same as PC3) | 2.34 (NOTE 1, 4) |
| Rx spurious | 6GHz <=f<=20GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 20GHz<=f<=40GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 40GHz<=f<=80GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
| NOTE 1: values assuming up to 6% of grid points with EIRP > 43dBm. NOTE 2: values assuming SNR = -7.9dB for points with EIRP > 43dBm / SNR = 8.14dB otherwise. NOTE 3: values assuming SNR = -5dB for points with EIRP > 43dBm / SNR = 10dB otherwise. NOTE 4: values assuming SNR = -9dB for points with EIRP > 43dBm / SNR = 6dB otherwise and considering the proposed relaxation. |  |  |  |

Table B.2.2.27-2a: Uncertainty value for influence of noise for PC1, MIMO for IFF

| Test case | Frequency range | Relaxation [dB] | Influence of noise [dB] |
| --- | --- | --- | --- |
| MOP-EIRP | FR2a | 0 | 0.13 |
|  | FR2b | 0 | 0.2 |
| MOP-TRP | FR2a | 0 | 0.13 |
|  | FR2b | 0 | 0.27 |
| MOP-Spherical | FR2a | 0 | 0.2 |
|  | FR2b | 0 | 0.35 |
| Minimum output power | FR2a | 0 | 0.15 |
|  | FR2b | 0 | 0.3 |
| General Tx spurious | Same as SISO in Table B.2.2.27-2 |  |  |
| Tx spurious Co-existence | Same as SISO in Table B.2.2.27-2 |  |  |
| Additional spurious emission | Same as SISO in Table B.2.2.27-2 |  |  |

< Unchanged sections omitted >

# B.18	Spurious emissions

Editor’s Note:

-	MU value analysis and offset value analysis for PC1, 2 and 4 are not complete.

-	MU offset value analysis for PC2 and 4 are not complete.

-	MU value analysis for various test setups in clause B.18.x is not complete for above 87 GHz.

Test procedure of general spurious emission comprises 2 stages: coarse TRP measurement and fine TRP measurement BW. Coarse TRP measurement is introduced to reduce the measurement time by applying sparser grids and/or wider measurement BW than fine TRP measurement while having offset dB more stringent test requirement in order not to cause additional misjudgement risk. For the frequency ranges for which coarse TRP measurement does not PASS, the measurement is continued with fine TRP measurement procedure.

Tables B.18-1,  B.18-1a, B.18-1b summarizes the MU threshold for fine TRP measurements for General spurious emissions, spurious emission band UE co-existence and additional spurious emission, respectively. The origin MU values for fine TRP measurement for different test setups can be found in following subclauses.

Table B.18-1: MU threshold for TRP measurement for general spurious emission

| Power Class | Diversity |  | Frequency |  | In-band BW |  | In-band Power (NOTE2) |  | Threshold MU value [dB] (NOTE1) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| PC3 | SISO, MIMO |  | 6 GHz <= f <=12.75 GHz |  | BW <= 400MHz |  | P = Max Output Power |  | 5.29 |  |
|  |  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  |  |  | 5.25 |  |
|  |  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  |  |  | 5.41 |  |
|  |  |  | 40.8 GHz <= f <= 66 GHz |  |  |  |  |  | 7.42 |  |
|  |  |  | 66 GHz <= f <= 80 GHz |  |  |  |  |  | 7.72 |  |
|  |  |  | 80 GHz < f <= 87 GHz |  |  |  |  |  | 8.14 |  |
| PC1 | SISO, MIMO |  | 6 GHz <= f <=12.75 GHz |  | BW <= 400MHz |  | P = Max Output Power |  | 5.28 |  |
|  |  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  |  |  | 5.91 |  |
|  |  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  |  |  | 6.07 |  |
|  |  |  | 40.8 GHz <= f <= 66 GHz |  |  |  |  |  | 8.09 |  |
|  |  |  | 66 GHz <= f <= 80 GHz |  |  |  |  |  | 7.71 |  |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.18.2-3 to Table B.18.2-11 for PC3 UEs and in Table B.18.2-12 to Table B.18.2-16 for PC1 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |  |  |  |  |  |  |

Table B.18-1a: MU threshold for TRP measurement for spurious emission band UE co-existence

| Power Class | Diversity |  | Frequency |  |  | In-band BW |  | In-band Power (NOTE2) | Threshold MU value [dB] (NOTE1) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| PC3 | SISO, MIMO |  | n257, n260, n261 |  |  | BW <= 400MHz |  | P = Max Output Power | 6.00 |  |
|  | SISO, MIMO |  | 23.6 GHz < f <= 24.0 GHz |  |  |  |  |  | 6.00 |  |
|  | SISO, MIMO |  | 36 GHz <= f <= 37 GHz |  |  |  |  |  | 6.00 |  |
|  | SISO, MIMO |  | 57 GHz <= f <= 66 GHz |  |  |  |  |  | 8.01 |  |
| PC1 | SISO, MIMO |  | n257, n261 |  |  | BW <= 400MHz |  | P = Max Output Power | 7.32 |  |
|  | SISO, MIMO |  | n260 |  |  |  |  |  | 7.32 |  |
|  | SISO, MIMO |  | 23.6 GHz < f <= 24.0 GHz |  |  |  |  |  | 7.32 |  |
|  | SISO, MIMO |  | 57 GHz <= f <= 66 GHz |  |  |  |  |  | 8.00 |  |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.18.2-3 to Table B.18.2-11 for PC3 UEs and in Table B.18.2-12 to Table B.18.2-16 for PC1 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |  |  |  |  |  |  |

Table B.18-1b: MU threshold for TRP measurement for additional spurious emission

| Power Class | Diversity | Frequency | In-band BW | In-band Power (NOTE2) | Threshold MU value [dB] (NOTE1) |
| --- | --- | --- | --- | --- | --- |
| PC3 | SISO, MIMO | 6 GHz <= f <=12.75 GHz NS_202 | BW <= 400MHz | P = Max Output Power | 5.29 |
|  |  | 12.75 GHz <= f <= 23.45 GHz NS_202 |  |  | 5.84 |
|  |  | 23.45 GHz <= f <= 40.8 GHz NS_202, NS_203 |  |  | 6.00 |
|  |  | 40.8 GHz <= f <= 2nd harmonic of the upper frequency edge of the UL operating band NS_202 |  |  | 8.01 |
| PC1 | SISO, MIMO | 6 GHz <= f <=12.75 GHz NS_202 | BW <= 400MHz | P = Max Output Power | 5.28 |
|  |  | 12.75 GHz <= f <= 23.45 GHz NS_202 |  |  | 7.16 |
|  |  | 23.45 GHz <= f <= 40.8 GHz NS_202, NS_203 |  |  | 7.32 |
|  |  | 40.8 GHz <= f <= 2nd harmonic of the upper frequency edge of the UL operating band NS_202 |  |  | 9.34 |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.18.2-3 to Table B.18.2-11 for PC3 UEs and in Table B.18.2-12 to Table B.18-2.16 for PC1 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |  |

## B.18.1	Uncertainty budget format and assessment for DFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.18.1-1.

Table B.18.1-1: Uncertainty contributions for TRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.1.1 |
| 2 | Measure distance uncertainty | B.2.1.2 |
| 3 | Quality of quiet zone | B.2.1.3 |
| 4 | Mismatch | B.2.1.4 |
| 5 | Standing Wave Between the DUT and measurement antenna | B.2.1.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.1.6 |
| 7 | Phase curvature | B.2.1.7 |
| 8 | Amplifier uncertainties | B.2.1.8 |
| 9 | Random uncertainty | B.2.1.9 |
| 10 | Influence of the XPD | B.2.1.10 |
| 11 | Insertion Loss Variation | B.2.1.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.1.12 |
| 13 | Influence of TRP measurement grid | B.2.1.22 |
| 14 | Influence of beam peak search grid | B.2.1.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.1.25 |
| 16 | DUT repositioning | B.2.1.26 |
| Stage 1: Calibration measurement |  |  |
| 17 | Mismatch | B.2.1.4 |
| 18 | Amplifier uncertainties | B.2.1.8 |
| 19 | Misalignment of positioning System | B.2.1.13 |
| 20 | Uncertainty of the Network Analyzer | B.2.1.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.1.15 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.1.16 |
| 23 | Phase centre offset of calibration antenna | B.2.1.18 |
| 24 | Quality of quiet zone for calibration process | B.2.1.19 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | B.2.1.20 |
| 26 | Influence of the calibration antenna feed cable | B.2.1.21 |
| 27 | Insertion Loss Variation | B.2.1.11 |
| Systematic uncertainties |  |  |
| 28 | Systematic error due to TRP calculation/quadrature | B.2.1.24 |
| 29 | Influence of noise | B.2.1.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {6 GHz to 87 GHz}, P = [Maximum output power].

-	The uncertainty assessment for TRP is provided in Table B.18.1-2 to B.18.1-xx

Table B.18.1-2: Uncertainty assessment for TRP measurement (f=TBD, D=TBD)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment |  |  |  |  |
| 2 | Measure distance uncertainty |  |  |  |  |
| 3 | Quality of quiet zone (NOTE 2) |  |  |  |  |
| 4 | Mismatch (NOTE 3) |  |  |  |  |
| 5 | Standing Wave Between the DUT and measurement antenna |  |  |  |  |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 4) |  |  |  |  |
| 7 | Phase curvature |  |  |  |  |
| 8 | Amplifier uncertainties |  |  |  |  |
| 9 | Random uncertainty |  |  |  |  |
| 10 | Influence of the XPD |  |  |  |  |
| 11 | Insertion Loss Variation |  |  |  |  |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |  |
| 13 | Influence of TRP measurement grid (NOTE 5) |  |  |  |  |
| 14 | Influence of beam peak search grid (NOTE 6) |  |  |  |  |
| 15 | Multiple measurement antenna uncertainty |  |  |  |  |
| 16 | DUT repositioning |  |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch |  |  |  |  |
| 18 | Amplifier uncertainties |  |  |  |  |
| 19 | Misalignment of positioning System |  |  |  |  |
| 20 | Uncertainty of the Network Analyzer |  |  |  |  |
| 21 | Uncertainty of the absolute gain of the calibration antenna |  |  |  |  |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |  |
| 23 | Phase centre offset of calibration antenna |  |  |  |  |
| 24 | Quality of quiet zone for calibration process (NOTE 2) |  |  |  |  |
| 25 | Standing wave between reference calibration antenna and measurement antenna |  |  |  |  |
| 26 | Influence of the calibration antenna feed cable |  |  |  |  |
| 27 | Insertion Loss Variation |  |  |  |  |
| TRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |
| Systematic uncertainties (NOTE 7) |  |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 5) |  |  |  |  |
| 29 | Influence of noise |  |  |  |  |
| Total measurement uncertainty |  |  |  |  |  |
| TRP total measurement uncertainty [dB] |  |  |  |  |  |
| NOTE 1: The impact of phase variation on EIRP is FFS. NOTE 2: The quality of quiet zone is different for EIRP and TRP. For TRP, the standard uncertainty is FFS; for EIRP, the standard uncertainty of quiet zone is FFS. NOTE 3: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 4: The assessment assumes maximum DUT output power. NOTE 5: This contributor shall only be considered for TRP measurements. NOTE 6: This contributor shall only be considered for EIRP measurements. NOTE 7: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 8: Void |  |  |  |  |  |

## B.18.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.18.2-1.

Table B.18.2-1: Uncertainty contributions for TRP measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | # B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of TRP measurement grid | B.2.2.22 |
| 14 | Influence of beam peak search grid | B.2.2.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 16 | DUT repositioning | B.2.2.26 |
| 17 | Misalignment of DUT due to change of DUT orientation | B.2.2.31 |
| Stage 1: Calibration measurement |  |  |
| 18 | Mismatch | B.2.2.4 |
| 19 | Amplifier Uncertainties | B.2.2.8 |
| 20 | Misalignment of positioning System | B.2.2.13 |
| 21 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 24 | Phase centre offset of calibration antenna | B.2.2.18 |
| 25 | Quality of quiet zone for calibration process | B.2.2.19 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 27 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 28 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 29 | Systematic error due to TRP calculation/quadrature | B.2.2.24 |
| 30 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet zone size ≤ 30 cm, f = {6 GHz to 87 GHz}, P = Maximum output power.

-	The uncertainty assessment for TRP is provided from Table B.18.2-2 to Table B.18.2-11 for PC3 UEs and from Table B.18.2-12 to Table B.18.2-16 for PC1 UEs.

Table B.18.2-2: Void

Table B.18.2-3: Uncertainty assessment for TRP measurement (f=6 GHz to 12.75GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.00 | Normal | 2.00 | 1.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (6 GHz < f <= 12.75 GHz) [dB] (a) |  |  |  | 4.88 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | General spurious emissions Influence of noise (c1) (6 GHz < f <= 12.75 GHz) |  |  |  | 0.41 |
| 30 | Additional spurious emissions Influence of noise (c2) NS_202 (6 GHz < f <= 12.75 GHz) |  |  |  | 0.41 |
|  | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (6 GHz < f <= 12.75 GHz) |  |  |  |  | 5.29 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c2) [dB] NS_202 (6 GHz < f <= 12.75 GHz) |  |  |  |  | 5.29 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: The analysis is valid for SISO and MIMO. |  |  |  |  |  |

Table B.18.2-4: Void

Table B.18.2-5: Uncertainty assessment for TRP measurement (f=12.75 GHz to 23.45 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (12.75 GHz < f <= 23.45 GHz) [dB] (a) |  |  |  | 4.84 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | General spurious emissions Influence of noise (c1) (12.75 GHz < f <= 23.45 GHz) |  |  |  | 0.41 |
| 30 | Additional spurious emissions Influence of noise (c2) NS_202 (12.75 GHz < f <= 23.45 GHz) |  |  |  | 1.0 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (12.75 GHz < f <= 23.45 GHz) |  |  |  |  | 5.25 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c2) [dB] NS_202 (12.75 GHz < f <= 23.45 GHz) |  |  |  |  | 5.84 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: The analysis is valid for SISO and MIMO. |  |  |  |  |  |

Table B.18.2-6: Void

Table B.18.2-7: Uncertainty assessment for TRP measurement (f=23.45 GHz to 40.8 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.40 | Actual | 1.00 | 1.40 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.73 | Normal | 2.00 | 1.37 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.6 | Normal | 2.00 | 0.3 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (23.45 GHz < f <= 40.8 GHz) [dB] (a) |  |  |  | 5.00 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | General spurious emissions Influence of noise (c1) (23.45 GHz < f <= 40.8 GHz) |  |  |  | 0.41 |
| 30 | Spurious emission band UE co-existence Influence of noise (c2) (f within NR Bands n257, n260 or n261) |  |  |  | 1.0 |
| 30 | Spurious emission band UE co-existence Influence of noise (c3) (36 GHz <= f <= 37 GHz) |  |  |  | 1.0 |
| 30 | Additional spurious emissions Influence of noise (c4) NS_202 (23.6 GHz < f <= 24.0 GHz) |  |  |  | 1.0 |
| 30 | Additional spurious emissions Influence of noise (c5) NS_202 (23.45 GHz < f <= 40.8 GHz) |  |  |  | 1.0 |
| 30 | Additional spurious emissions Influence of noise (c6) NS_203 (23.6 GHz < f <= 24.0 GHz) |  |  |  | 1.0 |
| 30 | Spurious emission band UE co-existence Influence of noise (c7) (23.6 GHz < f <= 24.0 GHz) |  |  |  | 1.0 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (23.45 GHz < f <= 40.8 GHz) |  |  |  |  | 5.41 |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c2) [dB] (f within NR Bands n257, n260 or n261) |  |  |  |  | 6.00 |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c3) [dB] (36 GHz <= f <= 37 GHz) |  |  |  |  | 6.00 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c4) [dB] NS_202 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | 6.00 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c5) [dB] NS_202 (23.45 GHz < f <= 40.8 GHz) |  |  |  |  | 6.00 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c6) [dB] NS_203 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | 6.00 |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c7) [dB] (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | 6.00 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: The analysis is valid for SISO and MIMO. |  |  |  |  |  |

Table B.18.2-8: Void

Table B.18.2-9: Uncertainty assessment for TRP measurement (f= 40.8 GHz to 66 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.0 | Normal | 2.00 | 0.0 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 2.30 | Actual | 1.00 | 2.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 4.0 | Normal | 2.00 | 2.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.7 | Normal | 2.00 | 0.85 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 1.70 | Normal | 2.00 | 0.85 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.28 | Normal | 2.00 | 0.14 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty ( 40.8 GHz < f <= 66 GHz) [dB] (a) |  |  |  | 7.01 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.00 |
| 30 | General spurious emissions Influence of noise (c1) (40.8 GHz < f <= 66 GHz) |  |  |  | 0.41 |
| 30 | Spurious emission band UE co-existence Influence of noise (c2) (57 GHz <= f <= 66 GHz) |  |  |  | 1.0 |
| 30 | Additional spurious emissions Influence of noise (c3) NS_202 (40.8 GHz < f <= 2nd harmonic of the upper frequency edge of the UL operating band) |  |  |  | 1.0 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (40.8 GHz < f <= 66 GHz) |  |  |  |  | 7.42 |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c2) [dB] (57 GHz <= f <= 66 GHz) |  |  |  |  | 8.01 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c3) [dB] NS_202 (40.8 GHz < f <= 2nd harmonic of the upper frequency edge of the UL operating band) |  |  |  |  | 8.01 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: Void NOTE 7: The analysis is valid for SISO and MIMO. |  |  |  |  |  |

Table B.18.2-10: Void

Table B.18.2-11: Uncertainty assessment for TRP measurement (f= 66 GHz to 87 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 2.30 | Actual | 1.00 | 2.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 4.00 | Normal | 2.00 | 2.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 3.0 | Normal | 2.00 | 1.50 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.70 | Normal | 2.00 | 0.85 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 1.70 | Normal | 2.00 | 0.85 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.28 | Normal | 2.00 | 0.14 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (66 GHz < f <= 87 GHz) [dB] (a) |  |  |  | 7.31 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.00 |
| 30 | General spurious emissions Influence of noise (c1) (66 GHz < f <= 80 GHz) |  |  |  | 0.41 |
| 30 | General spurious emissions Influence of noise (c1) (80 GHz < f <= 87 GHz) |  |  |  | 0.83 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (66 GHz < f <= 80 GHz) |  |  |  |  | 7.72 |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (80 GHz < f <= 87 GHz) |  |  |  |  | 8.14 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: The analysis is valid for SISO and MIMO. |  |  |  |  |  |

Table B.18.2-12: Uncertainty assessment for TRP measurement (f=6 GHz to 12.75GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.00 | Normal | 2.00 | 1.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (6 GHz < f <= 12.75 GHz) [dB] (a) |  |  |  | 4.87 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.00 |
| 30 | General spurious emissions Influence of noise (c1) (6 GHz < f <= 12.75 GHz) |  |  |  | 0.41 |
| 30 | Additional spurious emissions Influence of noise (c2) NS_202 (6 GHz < f <= 12.75 GHz) |  |  |  | 0.41 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (6 GHz < f <= 12.75 GHz) |  |  |  |  | 5.28 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c2) [dB] NS_202 (6 GHz < f <= 12.75 GHz) |  |  |  |  | 5.28 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: The analysis is valid for SISO and MIMO. |  |  |  |  |  |

Table B.18.2-13: Uncertainty assessment for TRP measurement (f=12.75 GHz to 23.45 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (12.75 GHz < f <= 23.45 GHz) [dB] (a) |  |  |  | 4.83 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.00 |
| 30 | General spurious emissions Influence of noise (c1) (12.75 GHz < f <= 23.45 GHz) |  |  |  | 1.08 |
| e30 | Additional spurious emissions Influence of noise (c2) NS_202 (12.75 GHz < f <= 23.45 GHz) |  |  |  | 2.34 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c) [dB] (12.75 GHz < f <= 23.45 GHz) |  |  |  |  | 5.91 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c2) [dB] NS_202 (12.75 GHz < f <= 23.45 GHz) |  |  |  |  | 7.16 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: The analysis is valid for SISO and MIMO. |  |  |  |  |  |

Table B.18.2-14: Uncertainty assessment for TRP measurement (f=23.45 GHz to 40.8 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.40 | Actual | 1.00 | 1.40 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.73 | Normal | 2.00 | 1.37 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.6 | Normal | 2.00 | 0.3 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (23.45 GHz < f <= 40.8 GHz) [dB] (a) |  |  |  | 4.99 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.00 |
| 30 | General spurious emissions Influence of noise (c1) (23.45 GHz < f <= 40.8 GHz) |  |  |  | 1.08 |
| 30 | Spurious emission band UE co-existence Influence of noise (c2) (f within NR Bands n257, n260 or n261) |  |  |  | 2.34 |
| 30 | Additional spurious emissions Influence of noise (c4) NS_202 (23.6 GHz < f <= 24.0 GHz) |  |  |  | 2.34 |
| 30 | Additional spurious emissions Influence of noise (c5) NS_202 (23.45 GHz < f <= 40.8 GHz) |  |  |  | 2.34 |
| 30 | Additional spurious emissions Influence of noise (c6) NS_203 (23.6 GHz < f <= 24.0 GHz) |  |  |  | 2.34 |
| 30 | Spurious emission band UE co-existence Influence of noise (c7) (23.6 GHz < f <= 24.0 GHz) |  |  |  | 2.34 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (23.45 GHz < f <= 40.8 GHz) |  |  |  |  | 6.07 |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c2) [dB] (f within NR Bands n257, n260 or n261) |  |  |  |  | 7.32 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c4) [dB] NS_202 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | 7.32 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c5) [dB] NS_202 (23.45 GHz < f <= 40.8 GHz) |  |  |  |  | 7.32 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c6) [dB] NS_203 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | 7.32 |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c7) [dB] (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | 7.32 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: The analysis is valid for SISO and MIMO. |  |  |  |  |  |

Table B.18.2-15: Uncertainty assessment for TRP measurement (f= 40.8 GHz to 66 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 2.30 | Actual | 1.00 | 2.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 4.0 | Normal | 2.00 | 2.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.7 | Normal | 2.00 | 0.85 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 1.70 | Normal | 2.00 | 0.85 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.28 | Normal | 2.00 | 0.14 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty ( 40.8 GHz < f <= 66 GHz) [dB] (a) |  |  |  | 7.00 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.00 |
| 30 | General spurious emissions Influence of noise (c1) (40.8 GHz < f <= 66 GHz) |  |  |  | 1.08 |
| 30 | Spurious emission band UE co-existence Influence of noise (c2) (57 GHz <= f <= 66 GHz) |  |  |  | 1 |
| 30 | Additional spurious emissions Influence of noise (c3) NS_202 (40.8 GHz < f <= 2nd harmonic of the upper frequency edge of the UL operating band) |  |  |  | 2.34 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (40.8 GHz < f <= 66 GHz) |  |  |  |  | 8.09 |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c2) [dB] (57 GHz <= f <= 66 GHz) |  |  |  |  | 8.00 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c3) [dB] NS_202 (40.8 GHz < f <= 2nd harmonic of the upper frequency edge of the UL operating band) |  |  |  |  | 9.34 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: Void NOTE 7: The analysis is valid for SISO and MIMO. |  |  |  |  |  |

Table B.18.2-16: Uncertainty assessment for TRP measurement (f= 66 GHz to 80 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 4 | Mismatch | 2.30 | Actual | 1.00 | 2.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 4.00 | Normal | 2.00 | 2.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 3.00 | Normal | 2.00 | 1.50 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.06 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.70 | Normal | 2.00 | 0.85 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 1.70 | Normal | 2.00 | 0.85 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.28 | Normal | 2.00 | 0.14 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (66 GHz < f <= 80 GHz) [dB] (a) |  |  |  | 7.30 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0 |
| 30 | General spurious emissions Influence of noise (c1) (66 GHz < f <= 80 GHz) |  |  |  | 0.41 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (66 GHz < f <= 80 GHz) |  |  |  |  | 7.71 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: The analysis is valid for SISO and MIMO. |  |  |  |  |  |

NOTE: MU assessment for additional spurious in Table B.18.2-3 to Table B.18.2-16 is based on the following relaxations:

Table B.18.2-17: Transmitter Spurious emissions relaxation considered in MU assessment (Quiet Zone size ≤ 30 cm)

| Power Class | Frequency | Relaxation |
| --- | --- | --- |
| PC1 | 6 GHz < f <= 12.75 GHz | 0 dB |
|  | 12.75 GHz < f <= 23.45 GHz | 0 dB |
|  | 23.45GHz <= f <= 40.8GHz | 0 dB |
|  | 40.8 GHz < f <= 66 GHz | 0 dB |
|  | 66 GHz < f <= 80 GHz | 0 dB |
| PC2 | 6 GHz < f <= 12.75 GHz | FFS |
|  | 12.75 GHz < f <= 23.45 GHz | FFS |
|  | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
|  | 66 GHz < f <= 87 GHz | FFS |
| PC3 | 6 GHz < f <= 12.75 GHz | 0 dB |
|  | 12.75 GHz < f <= 23.45 GHz | 0 dB |
|  | 23.45GHz <= f <= 40.8GHz | 0 dB |
|  | 40.8 GHz < f <= 66 GHz | 0 dB |
|  | 66 GHz < f <= 87 GHz | 0 dB |
| PC4 | 6 GHz < f <= 12.75 GHz | FFS |
|  | 12.75 GHz < f <= 23.45 GHz | FFS |
|  | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
|  | 66 GHz < f <= 80 GHz | FFS |

Table B.18.2-18: Spurious emissions band UE co-existence relaxation considered in MU assessment (Quiet Zone size ≤ 30 cm)

| Power Class | Frequency | Relaxation |
| --- | --- | --- |
| PC1 | 23.45GHz <= f <= 40.8GHz | 3.3 dB (for protected bands n257, n261) 5 dB (for protected band n260) 0.3 dB (for 23.6 GHz ≤ f ≤ 24.0 GHz) |
|  | 40.8 GHz < f <= 66 GHz | 0 dB (for 57.0 GHz ≤ f ≤ 66.0 GHz) |
| PC2 | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
| PC3 | 23.45GHz <= f <= 40.8GHz | 3.3 dB (for protected bands n257, n261) 5 dB (for protected band n260) 0.3 dB (for 23.6 GHz ≤ f ≤ 24.0 GHz) |
|  | 40.8 GHz < f <= 66 GHz | 6 dB (for 36.0 GHz ≤ f ≤ 37.0 GHz) 0 dB (for 57.0 GHz ≤ f ≤ 66.0 GHz) |
| PC4 | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |

Table B.18.2-19: Additional Spurious emissions relaxation considered in MU assessment (Quiet Zone size ≤ 30 cm)

| Power Class | Frequency | Relaxation |
| --- | --- | --- |
| PC1 | 6 GHz < f <= 12.75 GHz | 0 dB (NS_202) |
|  | 12.75 GHz < f <= 23.45 GHz | 13 dB (NS_202) |
|  | 23.45GHz <= f <= 40.8GHz | 13 dB (whole frequency range for NS_202) 0.3 dB (for 23.6 GHz ≤ f ≤ 24.0 GHz for NS_202 & NS_203) |
|  | 40.8 GHz < f <= 66 GHz | 13 dB (NS_202) |
|  | 66 GHz < f <= 80 GHz | FFS |
| PC2 | 6 GHz < f <= 12.75 GHz | FFS |
|  | 12.75 GHz < f <= 23.45 GHz | FFS |
|  | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
|  | 66 GHz < f <= 80 GHz | FFS |
| PC3 | 6 GHz < f <= 12.75 GHz | 0 dB (NS_202) |
|  | 12.75 GHz < f <= 23.45 GHz | 13 dB (NS_202) |
|  | 23.45GHz <= f <= 40.8GHz | 13 dB (whole frequency range for NS_202) 0.3 dB (for 23.6 GHz ≤ f ≤ 24.0 GHz for NS_202 & NS_203) |
|  | 40.8 GHz < f <= 66 GHz | 13 dB (NS_202) |
| PC4 | 6 GHz < f <= 12.75 GHz | FFS |
|  | 12.75 GHz < f <= 23.45 GHz | FFS |
|  | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
|  | 66 GHz < f <= 80 GHz | FFS |

< End of changes >