3GPP TSG-RAN WG4 Meeting #117												R4-2520214

Dallas, USA, 17 – 21 November, 2025

Agenda item:			7.3.3

Source:	Huawei, HiSilicon

Title:	on VSAT OTA tests

Document for:	approval

# Introduction

WF [1] from RAN4 #116bis highlights current issues. This contribution further discusses them.

# Discussion

Sub-topic 2-4 is on DUT size and weight. A search on current commercial products [2, 3, 4] shows the following sizes in flat panel configurations. Some products are housed within aerodynamic radomes, e.g. used on high speed trains.

Product 1: 45 x 45 (cm) [2]

Product 2: 230 x 115, 140 x 80,  65 x 60, 45 x 40 (cm) [3]

Product 3: 29 x 60 (cm) [4]

VSAT over two meters in diagnoal would require a quiet zone of similar size. A quiet zone over two meters would necessitate a substantial reflector in IFF configuration (aka compact antenna test range). Typically the ratio between the reflector antenna in IFF and quiet zone is 2:1. Some designs with serrated edges to reduce diffraction effect in quiet zone may achieve a smaller ratio.

Proposal 1: given the sizes of commercial VSAT range from 0.5m to 2.5m, RAN4 should choose a maximum testable VSAT size rather than use the largest commercially available VSAT size.

Most performance metrics defined in TS 38.101-5 [5] clause 9 for Tx and clause 10 for Rx can be tested with exceptions such as antenna gain as already identified in the last meeting WF [1]. Beam peak search is needed after locking a beam for performance measurement. 3dB beamwidth for a square antenna array of size D operating at wavelength l is estimated to be  l/D with the assumption that the antenna aperture is uniformly illuminated. This gives rise to the range of 3dB beamwidths based on the VSAT antenna sizes in proposal 1 as shown in Table 1. The 3dB beamwidths in Table 1 are estimated based on uniform aperture field distribution. Tapered aperture field distributions, e.g. 5dB tapering to reduce sidelobe levels, would broaden the 3dB beamwidth by a factor of 1.2 to 1.5. In addition, rectangular apertures would have two different 3dB beamwidths in the orthogonal planes. Regardless of the above mentioned minor differences, the range of 3dB beamwidths can be used to estimate the required peak search accuracy.

| wavelength at 14GHz (cm) | 2.1 |  |  |  |
| --- | --- | --- | --- | --- |
| wavelength at 35GHz (cm) | 0.9 |  |  |  |
| antenna size (cm) | 230 | 140 | 65 | 40 |
| 3dB BW 14GHz (degree) | 0.53 | 0.88 | 1.89 | 3.07 |
| 3dB BW at 35GHz (degree) | 0.21 | 0.35 | 0.76 | 1.23 |

Table 1  3dB beamwidths at 14GHz and 35GHz of various antenna sizes

Table 1 clearly shows that Ku and Ka bands have different 3dB beamwidths and would require different search step sizes.

Proposal 2: use l/D (where D is antenna size, l is wavelength) to estimate 3dB beamwidth required for beam peak search and use l/(4*D) as beam peak search step size.

Proposal 3: use different peak search step sizes for Ku and Ka bands due to the frequency gap between them.

For VSAT arrays with uniform or symmetrical aperture phase distributions, the direction of beam peak is normally perpendicular to the antenna plane. For example, if VSAT is placed in XY plane in a Cartesian coordinate with its centre at (0, 0), the direction of beam peak would be in the Z-axis direction. Under beam locking mode, it would be useful to assume the direction of beam peak is perpendicular to the VSAT antenna plane unless declared otherwise by manufacturers. This would minimise the beam peak search time.

Proposal 4: the direction of beam peak is assumed to be perpendicular to the VSAT antenna plane under beam locking mode unless declared otherwise by manufacturers.

The typical FR1 performance metrics such as TRP and TRS do not seem to apply to VSAT due to its higher frequency range. As a result, the performance tests should be selected from those specified in clause 9 and clause 10 in 38.101-5.

Proposal 5: VSAT performance tests should be selected from those specified in clause 9 and clause 10 in 38.101-5.

It is noted that VSAT antenna gain has been considered impossible to measure in the last RAN4 meeting. To avoid RAN4 producing untestable performance metrics, antenna gain requirements could be represented equivalently by similar EIRP requirements as EIRP and antenna gain are related.

Proposal 6: consider adding equivalent antenna gain requirements in the form of EIRP in TS 38.101-5 to avoid untestable requirements.

# Conclusion

This contribution provides the following proposals.

Proposal 1: given the sizes of commercial VSAT range from 0.5m to 2.5m, RAN4 should choose a maximum testable VSAT size rather than use the largest commercially available VSAT size.

Proposal 2: use l/D (where D is antenna size, l is wavelength) to estimate 3dB beamwidth required for beam peak search and use l/(4*D) as beam peak search step size.

Proposal 3: use different peak search step sizes for Ku and Ka bands due to the frequency gap between them.

Proposal 4: the direction of beam peak is assumed to be perpendicular to the VSAT antenna plane under beam locking mode unless declared otherwise by manufacturers.

Proposal 5: VSAT performance tests should be selected from those specified in clause 9 and clause 10 in 38.101-5.

Proposal 6: consider adding equivalent antenna gain requirements in the form of EIRP in TS 38.101-5 to avoid untestable requirements.

# References

- R4-2514820 WF on NR_UE_OTA_Enh

- Intelsat VSAT: https://www.intelsat.com/wp-content/uploads/2021/03/intelsat-flexmove-Paradigm-Swarm-techspecs-002.pdf

- SatixFy VSAT: https://www.satixfy.com/wp-content/uploads/2025/02/Ka-Ku-Terminals.pdf

- hiSky VSAT: https://hiskysat.com/product/

- TS 38.101-5 v19.2 User Equipment (UE) radio transmission and reception; Part 5: Satellite access Radio Frequency (RF) and performance requirements