3GPP TSG-RAN WG4 Meeting #106	R4-2302983
Athens, Greece, February 27 – March 3, 2023

Agenda item:	9.15.2.2

Source:	Bluetest AB

Title:	TP to TR 38.870 Addition of Reverberation Chamber Method (Alternate Method)

Release:	Rel-18

Document for:	Approval

# 1	Introduction

In RAN#104 meeting, the WF for Rel 18 TRP TRS WI outlined the agreements for the Reverberation Chamber (RC) test methodology [1].

In this contribution, we propose initial text to section 8 of TR 38.870 for capturing these agreements within the skeleton framework previously approved.

# 2	Background References

- R4-2220610, “WF for Rel-18 TRP TRS WI”, vivo, RAN#105 meeting, Nov 2022.

- 3GPP TR 37.941, “Universal Mobile Telecommunications System (UMTS); LTE; 5G; Radio Frequency (RF) conformance testing background for radiated Base Station (BS) requirements”, 2022.

- 3GPP TS 34.114, “User Equipment (UE) / Mobile Station (MS) Over The Air (OTA) antenna performance”, 2016.

- 3GPP TR 25.914, “Technical Specification 3rd Generation Partnership Project; Technical Specification Group Radio Access Networks; Measurements of Radio Performances for UMTS Terminals in Speech Mode”, 2022.

- 3GPP TR 37.902, “Technical Specification Group Radio Access Network; Measurements of User Equipment (UE) radio performances for LTE/UMTS terminals; Total Radiated Power (TRP) and Total Radiated Sensitivity (TRS) test methodology”, 2022.

# 3		TP to TR 38.870

The following text for TR 38.870 is proposed for approval:

<<  Start of text proposal, unchanged sections omitted >>

# 2	References

The following documents contain provisions which, through reference in this text, constitute provisions of the present document.

-	References are either specific (identified by date of publication, edition number, version number, etc.) or non-specific.

-	For a specific reference, subsequent revisions do not apply.

-	For a non-specific reference, the latest version applies. In the case of a reference to a 3GPP document (including a GSM document), a non-specific reference implicitly refers to the latest version of that document in the same Release as the present document.

<Editor’s note: reference to be added >

[1]	3GPP TR 21.905: "Vocabulary for 3GPP Specifications".

…

[5]	3GPP TS 38.521-1: "NR; User Equipment (UE) conformance specification; Radio transmission and reception; Part 1: Range 1 standalone"

[6]	3GPP TS 38.521-3: "NR; User Equipment (UE) conformance specification; Radio transmission and reception; Part 3: Range 1 and Range 2 Interworking operation with other radios"

[7]	3GPP TS 38.508-1: ”5GS; User Equipment (UE) conformance specification; Part 1: Common test environment”

[8]	3GPP TR 37.941: “Radio Frequency (RF) conformance testing background for radiated Base Station (BS) requirements”

…

[x]	<doctype> <#>[ ([up to and including]{yyyy[-mm]|V<a[.b[.c]]>}[onwards])]: "<Title>".

### 5.1.2	Definition of the Total Radiated Power (TRP) for RC method

TRP in the reverberation chamber is a measurement of transmitter performance in an isotropic Rayleigh fading environment that is based on sampling the radiated power of the UE/MS for a discrete number of field combinations. The average value of these statistically distributed samples is proportional to the TRP and by calibrating the average power transfer function, an absolute value of the TRP can be obtained.  The TRP with Reverberation Chamber method is defined as:

TRP=  PavgC1-R Pref

Where Pref is the reference power transfer function for the fixed measurement antenna, R is the reflection coefficient for the fixed measurement antenna and C is the path loss in the cables connecting the measurement receiver to fixed measurement antenna. These parameters are calculated from the calibration measurement and are further discussed in calibration section. Pavg is the average power measured by the fixed measurement antenna and can be calculated using the following expression:

Pavg=  m=1MS21,.m2M

Where S21,m is sample number m of the complex transfer function measured with the fixed measurement antenna and M is the total number of samples measured.

### 5.2.2	Definition of the Total Radiated Sensitivity (TRS) for RC method

The calculation of TRS is based on searching for the lowest power received by the UE for a discrete number of field combinations in the chamber. The downlink power received by the UE at each discrete field combination that provides a BER (or BLER) which is better than the specified target BER/BLER level shall be averaged with other such measurements using different field combinations. By calibrating the average power transfer function, an absolute value of the TRS can be obtained when the linear values of all downlink power levels have been averaged.

The TRS with Reverberation Chamber method is defined as:

TRS=(Pref 1-R C)1M m=1M1PBSSm-1

Where Pref is the reference power transfer function for the fixed measurement antenna, R is the reflection coefficient for the fixed measurement antenna and C is the path loss in the cables connecting the measurement receiver to fixed measurement antenna. The parameters are calculated from the calibration measurement and are further discussed in calibration section. PBSSm is the output power from the base station emulator when it is adjusted to give the specified digital error rate or throughput from the DUT for mode-stirring sample m.

<<  Unchanged sections not included  >>

# 8	Reverberation Chamber test methodology

## 8.1	General

This clause defines the test method with a Reverberation Chamber (RC) system as an alternate method for FR1 TRP and TRS testing.

## 8.2	Test setup

A reverberation chamber (RC) is an electrically large shielded metal enclosure that employs one or several "stirring" methods to randomize the fields, such as moving paddles, turntables, etc. In this way, a large number of uncorrelated samples is obtained. The volume in the room where the field is well-stirred is the working volume. Here the E-field, averaged over an entire stirring cycle, is independent of the location in the room, i.e., the field is spatially uniform.

Figure 8.2-1: Test setup diagram

Figure 8.2-2: S-parameter test setup diagram

## 8.3	Chamber Characterization

### 8.3.1	S-parameters and power transfer functions

Measurements of S-parameters in the reverberation chamber and derivations of power transfer functions are explained in detail in the calibration procedure.

### 8.3.2	Chamber loading for coherence bandwidth

The reverberation chamber can be loaded to control the power delay profile or coherence bandwidth in the chamber. However, the reverberation chamber should not be loaded to such an extent that the mode statistics in the chamber are impaired. It is important to keep the same loss profile in the chamber during calibration, measurement and test, in order not to change the average power transfer function between these two cases. Examples of lossy objects are blocks of RF absorber, head and hand phantoms. Configuration should be calculated to achieve a minimum of [FFS]

#### 8.3.2.1	Coherence bandwidth calculation

<< For further development >>

### 8.3.3	Chamber spatial uniformity

<< For further development >>

## 8.4	Calibration procedure

This procedure to measure S-parameters through the reverberation chamber must be performed for each measurement setup in which the loading of the chamber has been changed. Examples of objects that change the loading are phantoms, absorbers, antennas, etc. The calibration procedure must be repeated for each frequency range as defined above. Therefore, it is advantageous if the vector network analyzer can be configured to a frequency sweep covering the defined frequencies, so that all frequencies of interest can be measured with a minimal number of measurements runs. Place all objects into the RC which will be used during TRP or TRS measurements, including a head phantom, hand phantom and fixture for the UE. This ensures that the loss in the chamber, which determines the average power transfer level, is the same during both calibration and test measurements, e.g., TRP and TRS measurements.

Figure 8.4-1 S-parameter measurement setup in the RC, using a vector network analyzer.

- Place the calibration antenna in the chamber. The calibration antenna should be mounted on a low-loss dielectric fixture (i.e., loss tangent less than 0.05), to avoid effects from the fixture itself which may affect the UE’s radiation efficiency and mismatch factor. The calibration antenna must be placed in the chamber in such a way that it is far enough from any reflective and/or absorbing material such as walls, mode-stirrers, head phantom, or other object. This is to ensure that the environment for the calibration antenna (taken over the complete stirring sequence) resembles a free space environment. “Far enough away” depends on the type of calibration antenna used. For low gain nearly omni-directional antennas like dipoles, it is normally sufficient to ensure that this spacing is larger than 0.5 wavelengths from reflective objects and 0.7 wavelengths from absorbing objects at the lowest operating frequency. More directive calibration antennas should be oriented towards the center of the chamber. The calibration antenna should remain in the chamber during the TRP/TRS measurements.

- Calibrate the vector network analyzer with a full 2-port calibration in such a way that the vector complex S-parameters between the ports of the fixed measurement antenna and the calibration antenna can be accurately measured. Preferably, the vector network analyzer is set to perform a frequency sweep at each stirrer position. This will enable calibration of several frequency points during the same stirring sequence, thereby reducing calibration time. This will also enable frequency stirring, i.e., averaging the measured power transfer function over a small frequency bandwidth around each measured frequency point (moving frequency window). This will increase accuracy at the expense of (lower) frequency resolution.

- Place the calibration antenna in the chamber and connect it through the cables. The same procedure should be done for measurement antenna, if necessary. Measure the S-parameters for each fixed measurement antenna. The number of stirrer positions i.e. the number of S-parameter samples at each frequency point, should be chosen in such a way that it is large enough to yield an acceptable statistical contribution to the total measurement uncertainty. To this aim, at least 250 uncorrelated samples shall be used as defined in TR 37.941 clause 7.8.1 [8].

### 8.4.1	Calculation of the chamber reference transfer function

From the S-parameters obtained in the calibration measurement, the chamber reference transfer function for fixed antenna n can be calculated. The chamber reference transfer function for a fixed measurement antenna can be calculated as

Pref =1Mm=1 MS21,m2(1-R)(1-S22,ref2)⋅ 1eref

Where M is total number of stirrer positions, m is S-parameter sample number i.e., sampled at stirrer position m, the operator ⋅ denotes the absolute value, the operator  .   acting on parameter S22,ref denotes the average of complex parameters S22,ref over mode stirring sequence, S22,ref is the reflection coefficient for calibration antenna, S21,m is sample number m of the transfer function for the measurement antenna, eref is the radiation efficiency of the calibration antenna, and R is the reflection coefficient for the fixed antenna and can be calculated as

R= 1Mm=1MS11,m2=S112

Where S11,m  is sample number m of the reflection coefficient for a fixed antenna and S11 is the average of reflection coefficient.

Note that the radiation efficiency of the fixed antenna is not corrected for, because it will be the same both during calibration and measurements. Therefore, the fixed antenna’s radiation efficiency will not affect the final results.

## 8.5	TRP Test procedure

### 8.5.1	TRP for SA and EN-DC

#### 8.5.1.1	Test conditions

The TRP of the DUT is measured by sampling the S-parameters through a mode-stirring sequence and using these parameters to estimate the average power measured by any fixed measurement antenna. This data is used to determine the TRP values measured in the reverberation chamber. The measurement is performed with a step-mode stirring sequence and a frequency stirrer can be used.

#### 8.5.1.2	UE configurations

For devices containing multiple Tx antennas, the Tx Antenna Switching (TAS) function should be OFF, and the TRP should be measured for each Tx antenna individually. The antenna with better TRP is identified as the primary antenna, and the corresponding TRP result will be used to determine the pass/fail compliance. Otherwise, the primary antenna should be selected based on manufacturer declaration.

For Standalone, the NR System Simulator (SS) and DUT shall be configured per TS 38.521-1 [5], section 6.2.1 (UE maximum output power) using the default settings specified in TS 38.521-1 [5] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4.3.

For EN-DC, the SS and DUT shall be configured per TS 38.521-3 [6], Section 6.2B.1 (UE Maximum Output Power for EN-DC) using the default settings specified in TS 38.521-3 [6] and TS 38.508 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4.3. The UL output power of LTE carrier should be set as a constant power of 10dBm, while measuring NR at maximum output power, i.e., with fixed p-MaxEUTRA-r15=10 dBm, and p-NR-FR1 not configured.

#### 8.5.1.3	Test procedure

For TRP measurement, the evaluations shall be performed at maximum transmit power. The measurement procedure includes the following steps:

- Place the DUT inside the test zone.

- Connect the DUT through the cables/chamber to the instrument and ensure the DUT is transmitting at its maximum power.

- Measure the power at each measurement point, and calculate the TRP by averaging over the sampling sequence. At least [150] uncorrelated measurement samples shall be used. The loss of the entire transmission path must be compensated.

The TRP value is calculated using the TRP equation in the definition outlined in 5.2.1.

## 8.6	TRS Test procedure

### 8.6.1   TRS for SA and EN-DC

#### 8.6.1.1   Test conditions

The TRS of the DUT is based on searching for the lowest power received by UE for a discrete number of field combinations in the chamber. The power received by the UE at each discrete field combination that provides a BER (or BLER) which is better than the specified target BER/BLER level shall be averaged with other such measurements using different field combinations. The TRS of the DUT is estimated by measuring the transfer function for each fixed measurement antenna, which gives the BER threshold.

#### 8.6.1.2   UE configurations

For Standalone, the NR System Simulator (SS) and DUT shall be configured per section 7.3.2 (Reference sensitivity power level) of TS 38.521-1 [5] using the defaults specified in TS 38.521-1 [5] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4.

For EN-DC, the EN-DC SS and DUT shall be configured per section 7.3B.2 (Reference Sensitivity for EN-DC) of TS 38.521-3 [6], using the defaults specified in TS 38.521-3 [6] and TS 38.508 [7], as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4. The UL power configuration for LTE and NR is 50%-50% power splitting, i.e.,

- For PC3, p-MaxEUTRA-r15=20 dBm, and p-NR-FR1= 20dBm;

- For PC2, p-MaxEUTRA-r15=23 dBm, and p-NR-FR1= 23dBm.

#### 8.6.1.3   Test procedure

For TRS measurement, the evaluations shall be performed at maximum transmit power.

The measurement procedure includes the following steps:

- Place the DUT inside the test zone.

- Connect the DUT through the cables/chamber to the instrument and ensure the DUT is transmitting at its maximum power.

- Follow steps 1 through 4 in section 7.3.2.4.2 of TS 38.521-1 [5], with the following exception: determine each EIS, i.e., by adjusting the downlink signal level until the minimum power level at which the throughput exceeds or equal to 95% of the maximum throughput of the specified RMC, at each sampling point. The downlink power step size shall be no more than 0.5 dB when the RF power level is near the NR sensitivity level. At least [100] uncorrelated measurement sampling points shall be used.

- The TRS value is calculated using the equation outlined in 5.2.2

## 8.6	Test Volume

The reverberation chamber shall have a volume large enough to support the number of modes needed for the stated accuracy at the lowest operating frequency. If the UE/MS is moved around in the chamber during the measurement, the volume of the reverberation chamber can be reduced. In addition to the physical chamber size considerations, the test volume of the chamber must be large enough to place a reference antenna/test object at a minimum distance of one-half wavelength, i.e., λ/2 from the chamber walls, mode stirrers, and other electromagnetic reflective objects within the chamber, and 0.7 wavelengths from absorbing objects at the lowest test frequency.