3GPP TSG-RAN4 Meeting #115	R4-2505843

# Malta, Malta, 19th - 23rd May, 2025

Agenda item:		7.14.2

Source:				Keysight Technologies UK Ltd

Title:					On Inadequate Range Length MU

Document for:		Approval

# Introduction

This contribution introduces the next steps to determine the MU element for inadequate measurement distance.

# Min Measurement Distance for 50cm QZ and Motivation for Inadequate Measurement Distance MU

In the past, the minimum measurement distances for 20cm [1] and 30cm [2][3] DFF systems were defined as the maximum of the following three limits

-	The phase uncertainty limit: RQZ+2Drad2/

-	The amplitude uncertainty limit: 3DQZ

-	The reactive Near-Field limit: RQZ+2

where RQZ is defined as the radius of the quiet zone, i.e., RQZ=DQZ/2, and Drad is the diameter of the effective radiating aperture. The minimum range length calculations for DQZ={20, 30}cm quiet zone size OTA test systems shall assume that Drad is {20, 30}cm below 1GHz and decrease linearly from {20, 30}cm to 5cm from 1GHz to 7.125GHz, respectively.

For the 50cm QC, the minimum measurement distance/range length is tabulated in Table 1 while the three different limits including the maximum are plotted in Figure 1. The distribution of the effective radiating aperture is shown in Figure 2. The values in Table 1 match those introduced in [4].

Table 1: Minimum Range Length for OTA systems with 50cm quiet zone size.

| f [GHz] | Drad [m] | RQZ+2Drad²/λ | 3DQZ = 6RQZ | RQZ+2λ | max(RQZ+2λ,3DQZ,RQZ+2Drad²/λ) |
| --- | --- | --- | --- | --- | --- |
| 0.41 | 0.50 | 0.93 | 1.5 | 1.71 | 1.71 |
| 0.6 | 0.50 | 1.25 | 1.5 | 1.25 | 1.50 |
| 0.7 | 0.50 | 1.42 | 1.5 | 1.11 | 1.50 |
| 0.8 | 0.50 | 1.58 | 1.5 | 1.00 | 1.58 |
| 1 | 0.50 | 1.92 | 1.5 | 0.85 | 1.92 |
| 1.2 | 0.49 | 2.14 | 1.5 | 0.75 | 2.14 |
| 1.4 | 0.47 | 2.32 | 1.5 | 0.68 | 2.32 |
| 1.6 | 0.46 | 2.47 | 1.5 | 0.62 | 2.47 |
| 1.8 | 0.44 | 2.59 | 1.5 | 0.58 | 2.59 |
| 2 | 0.43 | 2.68 | 1.5 | 0.55 | 2.68 |
| 2.2 | 0.41 | 2.74 | 1.5 | 0.52 | 2.74 |
| 2.4 | 0.40 | 2.78 | 1.5 | 0.50 | 2.78 |
| 2.6 | 0.38 | 2.79 | 1.5 | 0.48 | 2.79 |
| 2.8 | 0.37 | 2.78 | 1.5 | 0.46 | 2.78 |
| 3 | 0.35 | 2.74 | 1.5 | 0.45 | 2.74 |
| 4 | 0.28 | 2.34 | 1.5 | 0.40 | 2.34 |
| 5 | 0.21 | 1.67 | 1.5 | 0.37 | 1.67 |
| 6 | 0.13 | 0.95 | 1.5 | 0.35 | 1.50 |
| 7 | 0.06 | 0.41 | 1.5 | 0.34 | 1.50 |
| 7.125 | 0.05 | 0.37 | 1.5 | 0.33 | 1.50 |

Figure 1: Range length limits for 50cm QZ

Figure 2: Effective radiating aperture, Drad, for 50cm QZ

Clearly, these analyses yield some very large and rather unpractical length lengths for 50 cm QZ systems, especially between 1 GHz and 4.5 GHz.

Observation 1: Testing large devices in a 50cm QZ requires very large and rather unpractical range lengths, especially between 1 and 4.5 GHz.

Consequently, it should be investigated whether smaller range lengths could be considered for large devices tested in a 50 cm QZ with the introduction of a measurement uncertainty for inadequate measurement distance which is what the editor’s note of [4] suggests

Editor’s Note: Performing measurements in NR FR1 TRP-TRS OTA systems with 50cm quiet zone size at shorter range lengths than those defined in Table 7.7.1-2 with an additional measurement uncertainty for inadequate measurement distance is FFS.

# Inadequate Measurement Distance MU Analyses

The initial approach to determine the effect of reducing the min. measurement distance was to look into the impact the difference in path loss has on an omnidirectional antenna pattern in addition to the limited/finite beamwidth of the measurement probe, i.e., for relatively small range lengths, the offset radiator is not always in the boresight direction of the probe but at an offset direction which has lower gain than along boresight. For this analysis, the path losses from the probe antenna

- to the centre of the QZ zone, to which OTA systems are generally calibrated to and

- to the known offset location radiator offset

were determined. The measurement probe was assumed with a rotationally symmetrical 40° HPBW and a gain variation as a function of offset angles as shown in Figure 3, i.e.,

Grelative=-12Offset AngleHPBW2

For each measurement grid point, the angular offset from boresight between the measurement probe and the offset radiator was determined and taken into account.

Figure 3: Relative measurement probe gain as a function of offset angle

TRPs were calculated using a classical TRP measurement grid with Dq=Df=15° and the Clenshaw-Curtis quadrature. Figure 4 illustrates the impact the path loss differences and the finite measurement probe beamwidth of 40° HPBW have on a radiator with omnidirectional pattern displaced in the y direction for various offsets and range lengths.

- Top left: displacement by 0.15m (RQZ) in y with a range length of 1.15m (min range length for 30cm QZ with ≥600MHz)  TRP impact of -0.25dB

- Top right: displacement by 0.25m (RQZ) in y with a range length of 2.8m (min range length for 50cm QZ) TRP impact of -0.12dB

- Bottom left: displacement by 0.25m in y with a range length of 1.15m (min range length for 30cm QZ) TRP impact of -0.70dB

- Bottom right: displacement by 0.25m in y with a range length of 1.5m  TRP impact of -0.41dB

Figure 4: FSPL and finite measurement probe beamwidth impact due to offset omnidirectional radiator

Unfortunately, this analysis is somewhat inconclusive as the TRP impacts presented above are frequency independent and represent a worst-case only and thus do not support the assessment of a measurement uncertainty for inadequate measurement distance.

The next set of analyses assumed that the offsets were distributed uniformly within a 30 cm or 50 cm spherical volume. The simulated offsets are visualized in Figure 5.

Figure 5: Simulated offsets distributed evenly within the 30 cm (left) or 50 cm (right) quiet zone

For each simulated offset, the TRP error was calculated based on the FSPL differences, i.e., the path loss differences were not compensated since the antenna offsets for OTA measurements are generally unknown, and the finite beamwidth effect of the probe and recorded. The histogram of all TRP errors is then plotted and shown for the previously investigated maximum offsets/QZ radii and range lengths in Figure 6. As expected, the maximum TRP error in Figure 6 matches the previously reported maximum TRP error based on the maximum radiator offset, see Figure 4. It should also be noted that the distribution is single sided, i.e., the TRP cannot be overestimated using the effects investigated. This is why the RMS values in addition to the standard deviations are presented.

Figure 6: Histogram of TRP Errors for various maximum offsets and range lengths

The latest set of results show that based on this very simple assumption of point radiators with omnidirectional patterns and measurement probes with finite beamwidth of 40° HPBW:

- A range length of 1.15m for a 30cm QZ incurs a small MU for inadequate measurement distance (even though it was assumed that the 1.15m range length is sufficient for a 30cm QZ)

- A range length of 2.8m for a 50cm QZ incurs an almost insignificant MU for inadequate measurement distance

- Range lengths of 1.15m and 1.5m for a 50cm QZ incur a noticeable MU for inadequate measurement distance

Observation 2: A range length of 1.15m for a 30cm QZ incurs a small MU for inadequate measurement distance (even though it was assumed that the 1.15m range length is sufficient for a 30cm QZ)

Observation 3: A range length of 2.8m for a 50cm QZ incurs an almost insignificant MU for inadequate measurement distance

Observation 4: Range lengths of 1.15m and 1.5m for a 50cm QZ incur a noticeable MU for inadequate measurement distance

Needless to say that these analyses are based on very simplistic assumptions, e.g., point sources with omnidirectional patterns, no frequency dependence of patterns or probe HPBW, symmetrical beamwidth assumptions of the measurement probe, far-field assumptions, etc. However, some very interesting and potentially relevant observations can be made based on the statistical analyses presented here and adding some realism could allow an assessment of the MU for inadequate measurement distance. Extensive numerical and/or empirical analyses could also be performed for different form factors, frequency ranges, range lengths, etc. to assess the MU due to inadequate measurement distance more accurately; however, such simulations and/or measurements will take time and require volunteers.

It is proposed to determine whether to base the inadequate measurement distance MU for large devices tested with 50 cm QZ on measurements and/or EM simulations at various frequencies and for various form factors or whether to continue with rather simplistic statistical analyses such as those presented in this contribution but by adding a little bit more realism, e.g., directive patterns with frequency dependence, frequency-dependent measurement probe beamwidths, etc.

Proposal 1: Determine whether to base the inadequate measurement distance MU for large devices tested with 50 cm QZ on measurements and/or EM simulations at various frequencies and for various form factors or whether to continue with rather simplistic statistical analyses such as those presented in this contribution with more realism, e.g., frequency dependendent patterns, frequency dependent HPBWs, etc.

A couple of notes based on feedback received from offline discussions:

- The simulated offsets were distributed within a spherical volume and could be expanded to cover a cylindrical volume. While some ripple testing is performed over a spherical volume (theta-axis ripple test), other testing is performed over a cylindrical volume (phi-axis ripple test).

- The limited/finite beamwidth of the measurement probe is, to some degree, already included in the ripple test and some aspects could potentially be accounted for twice. However, the SSD calculation of multiple 2D scans was meant to address the multipath effects of the support structure on the positioner, which caused significant deviations to the pattern (deep nulls either side of the pole) but had relatively small impact on the TRP.

Observation 5: The effect of the limited/finite beamwidth of the measurement probe might already included in the ripple test

The next section extends some of the existing analyses to the ripple test.

# SSD MU Analyses from Ripple Test

The ripple test for SISO OTA test cases is defined in Clauses 7.4 and B.2.8 of [5] and Clauses 7.6 and B.2.8 of [6], specifically the test procedure and the analysis of the measured data. The phi-axis test is investigated more closely next. Here, for the standard 30 cm QZ, three different cuts are defined

For the phi-axis ripple test, each reference antenna is oriented with its axis parallel to the phi axis at a total of three positions, offset 15 cm perpendicular to the phi axis with 0 cm and ±15 cm offsets parallel to the phi axis. At each position, the phi axis is rotated 360° to record the ripple. Each position is labeled by its radial and axial offset from the center position, (R, Z). See Figure 7.6.2-2 for additional information. Figure 7.6.2-2: Phi-axis test geometry

Three similar cuts are defined for the larger 50 cm QZ

For the notebook sized test volume, three additional test positions are added to the phi-axis test in order to cover the larger test cylinder. These positions are offset 25 cm perpendicular to the phi axis with -15, 0, and +21 cm offsets parallel to the phi axis, as shown in Figure 7.6.2-3. Figure 7.6.2-3: Phi-axis test geometry for the notebook sized test volume

Similar analyses as in the previous section were performed for each of the phi-axis cuts and the SSDs were calculated for each cut j

sjpk=1N-1k-0N-1pkp-12

with pk = 10PkdB10 and p as the average of all pk values.

However, some important aspects such as the assessment of the MU from the ripple test results and the allowance to compensate the path loss variation when the antenna is offset from the centre of QZ is not specified/defined in [5], [6]; on the other hand, some of those details are defined in [3], [7] where the ripple test was defined originally.

Observation 6: Some important aspects of the ripple test for FR1 are not sufficiently defined in 3GPP

The remaining analyses take some of the missing aspects from CTIA into account [3], [7], e.g., MU calculations. The standard uncertainty contribution due to the ripple test is given by the maximum of all the 𝑠(𝑝𝑘) values for all offsets 𝑢(𝑥) = 10 log(1 + 𝑚𝑎𝑥(𝑠𝑗(𝑝𝑘))). For this investigation of the SSD MU, the path loss due to offset was ignored as the ripple test per CTIA [3], [7] allows the compensation of the FSPL (using the law of cosine to account for the path loss variation when the antenna is offset from the centre). Additionally, the unlike the analyses in the previous section which assumed an omnidirectional antenna, the analyses in this section assume dipole like patterns as the phi-axis ripple test assume dipoles and loops with “donut shape” antenna patterns. Here, the offset of the radiator in (x,y,z) is taken into account to determine the actual patterns perceived by the test antenna, e.g., by re-assessing the coordinate system specifically the q' coordinate for the dipole pattern, see Figure 7.

Figure 7: Coordinate system re-assessment for offset radiator

The histograms of the pk(dB) values for all three cuts and angles are shown for the two different QZs and the three different range lengths considered are shown in Figure 8. Just like the distributions in the previous section, they are single sided for the pk(dB) values.

Figure 8: Histogram of pk(dB) values for all cuts and angles

The statistics of the TRP errors assuming omnidirectional radiators and the ripple test statistics assuming a dipole like pattern are compared in Table 2. For the TRP errors, the unknown spatial offsets from the centre of QZ and the angular offsets from the probe antenna were taken into account, see Figure 6, while for the SSD MU, only the angular offsets from the probe antenna were taken into account with the FSPL compensated.

Table 2: Comparison of TRP Error and Ripple Test SSD MUs

| QZ Zone Radius [cm] | Range Length [m] | TRP Error (FSPL and finite/limited HPBW) |  | Ripple Test Statistics of pk for dipole pattern (FSPL compensated) [dB] |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Std. Dev [dB] | RMS [dB] | SSD u(x) [dB] | Std. Dev. [dB] | RMS [dB] |
| 15 | 1.15 | 0.07 | 0.17 | 0.17 | 0.29 | 0.61 |
| 25 | 2.8 | 0.03 | 0.08 | 0.07 | 0.10 | 0.21 |
| 25 | 1.5 | 0.11 | 0.27 | 0.27 | 0.35 | 0.72 |
| 25 | 1.15 | 0.17 | 0.46 | 0.50 | 0.61 | 1.25 |

Observation 7: The ripple test SSD MUs are greater than the TRP error standard deviations.

Observation 8: The standard deviation of the pk(dB) from all three cuts is slightly larger than the MUs from the SSD.

Observation 9: The RMS of the ripple test errors is significantly larger than the RMS of the TRP errors.

Given the significantly larger uncertainties with similar analyses observed with the ripple test coordinates when compared to the test coordinates placed uniformly within the QZ, it seems impossible to determine and compensate the duplicate effects. Additionally, given the allowance to account for the path loss variation when the antenna is offset from the centre of the QZ in the ripple test procedure even for the Ripple Test MU for DUT measurements where the path loss variation cannot be compensated, it is proposed to include the finite/limited HPBW effect of the measurement probe in the MU analysis for the inadequate measurement distance MU without trying to account for duplication from the ripple test.

Proposal 2: Consider including the finite/limited HPBW effect of the measurement probe in the MU analysis for the inadequate measurement distance MU without trying to account for duplication from the ripple test.

# Proposal/Action Requested

Observation 1: Testing large devices in a 50cm QZ requires very large and rather unpractical range lengths, especially between 1 and 4.5 GHz.

Observation 2: A range length of 1.15m for a 30cm QZ incurs a small MU for inadequate measurement distance (even though it was assumed that the 1.15m range length is sufficient for a 30cm QZ)

Observation 3: A range length of 2.8m for a 50cm QZ incurs an almost insignificant MU for inadequate measurement distance

Observation 4: Range lengths of 1.15m and 1.5m for a 50cm QZ incur a noticeable MU for inadequate measurement distance

Observation 5: The effect of the limited/finite beamwidth of the measurement probe might already included in the ripple test

Observation 6: Some important aspects of the ripple test for FR1 are not sufficiently defined in 3GPP

Observation 7: The ripple test SSD MUs are greater than the TRP error standard deviations.

Observation 8: The standard deviation of the pk(dB) from all three cuts is slightly larger than the MUs from the SSD.

Observation 9: The RMS of the ripple test errors is significantly larger than the RMS of the TRP errors.

Proposal 1: Determine whether to base the inadequate measurement distance MU for large devices tested with 50 cm QZ on measurements and/or EM simulations at various frequencies and for various form factors or whether to continue with rather simplistic statistical analyses such as those presented in this contribution with more realism, e.g., frequency dependendent patterns, frequency dependent HPBWs, etc.

Proposal 2: Consider including the finite/limited HPBW effect of the measurement probe in the MU analysis for the inadequate measurement distance MU without trying to account for duplication from the ripple test.

# References

- TR 38.827, Study on radiated metrics and test methodology for the verification of multi-antenna reception performance of NR User Equipment (UE)

- TS 38.161, User Equipment (UE) TRP (Total Radiated Power) and TRS (Total Radiated Sensitivity) requirements; Range 1 Standalone and Range 1 Interworking operation with other radios

- CTIA 01.73, Supporting Procedures

- R4-2504758, DraftCR to add min measurement distance for 50cm QZ, Keysight Technologies UK Ltd, Huawei, Ericsson, Rohde&Schwarz, 3GPP TSG-RAN4 Meeting #114-bis, April 2025

- TR 38.834, Total Radiated Power (TRP) and Total Radiated Sensitivity (TRS) test methodology

- TR 38.870, Enhanced Over-the-Air (OTA) test methods for NR FR1 Total Radiated Power (TRP) and Total Radiated Sensitivity (TRS)

- CTIA 01.70, Measurement Uncertainty