3GPP TSG-RAN WG4 meeting #79	R4-163750

Nanjing, China, 23 - 27 May, 2016

Agenda Item:	6.2.2.2 OTA testing

Source: 	Sumitomo Electric

Title:  	TP for TR37.842: Adding uncertainty value to the EIRP measurement with Indoor Anechoic Chamber

Document for:	Approval

- 1. Introduction

During the RAN4#78 in Malta, the description, procedure, and uncertainty contribution factors for EIRP measurement with Indoor Anechoic Chamber was presented and approved [1]. According to the agreed conformance test framework in sub-clause 10.1 of the TR [2], the next step is to determine the uncertainty budget of the method in order to finally determine an overall acceptable test tolerance for the EIRP measurement.

Companies proposed the uncertainty budget in the last meeting RAN4#78bis [3] [4], but no agreement was achieved. During the online and offline discussion, it was identified that it is better to align the calibration setup of the Indoor Anechoic Chamber method with other test methods, using network analyzer rather than signal generator and spectrum analyzer, and revisit the uncertainty budget. An updated text for “indoor anechoic chamber” method is therefore proposed for this meeting in [5] where the test procedure and uncertainty contribution factors are revisited following the network analyzer calibration method.

This contribution proposes a TP to add the uncertainty value to the EIRP measurement with Indoor Anechoic Chamber following the updated uncertainty contribution factors given in [5].

- 2. Discussion

The tables below are the proposed values (dB) from SEI on the uncertainty of each contribution factor of the EIRP measurement with Indoor Anechoic Chamber. The final expanded uncertainty (1.96σ - confidence interval of 95%) can be determined from the proposed values as 1.14dB for f≦3.0GHz and 1.27dB for 3.0GHz<f≦4.2GHz.

Table 1-1 Proposed uncertainty value (f≦3.0GHz)

| UID | Description of uncertainty contribution | Uncertainty distribution | Maximum value of the uncertainty |  | Divisor | Standard uncertainty |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 1, DUT measurement |  |  |  |  |  |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | Rectangular | 0.03 | 3 |  |  | 0.02 |
| 2 | Pointing misalignment between the AAS BS and the receiving antenna. | Rectangular | Horizontal: 0.31 Vertical: 0.21 | 3 |  |  | 0.22 |
| 3 | Quality of quiet zone | Gaussian | Not applicable | Not applicable |  |  | 0.10 |
| 4 | Polarization mismatch between the AAS BS and the receiving antenna | Rectangular | 0.01 | 3 |  |  | 0.01 |
| 5 | Mutual coupling between the AAS BS and the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.00 |
| 6 | Phase curvature | Gaussian | Not applicable | Not applicable |  |  | 0.05 |
| 7 | Uncertainty of the measurement equipment | Gaussian | Not applicable | Not applicable |  |  | 0.32 |
| 8 | Impedance mismatch in the receiving chain | U | 0.20 | 2 |  |  | 0.14 |
| 9 | Random uncertainty | Gaussian | Not applicable | Not applicable |  |  | 0.20 |
| Stage 2, Calibration measurement |  |  |  |  |  |  |  |
| 10 | Impedance mismatch between the receiving antenna and the network analyzer | U | 0.05 | 2 |  |  | 0.04 |
| 11 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | Rectangular | 0.01 | 3 |  |  | 0.01 |
| 12 | Impedance mismatch between the reference antenna and the network analyzer. | U | 0.05 | 2 |  |  | 0.04 |
| 13 | Quality of quiet zone | Gaussian | Not applicable | Not applicable |  |  | 0.10 |
| 14 | Polarization mismatch for reference antenna | Rectangular | 0.01 | 3 |  |  | 0.01 |
| 15 | Mutual coupling between the reference antenna and the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.00 |
| 16 | Phase curvature | Gaussian | Not applicable | Not applicable |  |  | 0.05 |
| 17 | Uncertainty of the network analyzer | Gaussian | Not applicable | Not applicable |  |  | 0.10 |
| 18 | Influence of the reference antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | Rectangular | 0.05 | 3 |  |  | 0.03 |
| 19 | Reference antenna feed cable loss measurement uncertainty | Gaussian | Not applicable | Not applicable |  |  | 0.06 |
| 20 | Influence of the receiving antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | Rectangular | 0.05 | 3 |  |  | 0.03 |
| 21 | Uncertainty of the absolute gain of the reference antenna | Rectangular | 0.50 | 3 |  |  | 0.29 |
| 22 | Uncertainty of the absolute gain of the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.00 |
| Combined standard uncertainty (1 σ) |  |  |  |  |  |  | 0.58 |
| Expanded uncertainty (1.96σ – confidence interval of 95%) |  |  |  |  |  |  | 1.14 |

Table 1-2 Proposed uncertainty value (3.0GHz<f≦4.2GHz)

| UID | Description of uncertainty contribution | Uncertainty distribution | Maximum value of the uncertainty |  | Divisor | Standard uncertainty |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 1, DUT measurement |  |  |  |  |  |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | Rectangular | 0.03 | 3 |  |  | 0.02 |
| 2 | Pointing misalignment between the AAS BS and the receiving antenna. | Rectangular | Horizontal: 0.31 Vertical: 0.21 | 3 |  |  | 0.22 |
| 3 | Quality of quiet zone | Gaussian | Not applicable | Not applicable |  |  | 0.10 |
| 4 | Polarization mismatch between the AAS BS and the receiving antenna | Rectangular | 0.01 | 3 |  |  | 0.01 |
| 5 | Mutual coupling between the AAS BS and the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.00 |
| 6 | Phase curvature | Gaussian | Not applicable | Not applicable |  |  | 0.05 |
| 7 | Uncertainty of the measurement equipment | Gaussian | Not applicable | Not applicable |  |  | 0.45 |
| 8 | Impedance mismatch in the receiving chain | U | 0.33 | 2 |  |  | 0.23 |
| 9 | Random uncertainty | Gaussian | Not applicable | Not applicable |  |  | 0.20 |
| Stage 2, Calibration measurement |  |  |  |  |  |  |  |
| 10 | Impedance mismatch between the receiving antenna and the network analyzer | U | 0.05 | 2 |  |  | 0.01 |
| 11 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | Rectangular | 0.01 | 3 |  |  | 0.04 |
| 12 | Impedance mismatch between the reference antenna and the network analyzer. | U | 0.05 | 2 |  |  | 0.10 |
| 13 | Quality of quiet zone | Gaussian | Not applicable | Not applicable |  |  | 0.01 |
| 14 | Polarization mismatch for reference antenna | Rectangular | 0.01 | 3 |  |  | 0.00 |
| 15 | Mutual coupling between the reference antenna and the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.05 |
| 16 | Phase curvature | Gaussian | Not applicable | Not applicable |  |  | 0.11 |
| 17 | Uncertainty of the network analyzer | Gaussian | Not applicable | Not applicable |  |  | 0.03 |
| 18 | Influence of the reference antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | Rectangular | 0.05 | 3 |  |  | 0.06 |
| 19 | Reference antenna feed cable loss measurement uncertainty | Gaussian | Not applicable | Not applicable |  |  | 0.03 |
| 20 | Influence of the receiving antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | Rectangular | 0.05 | 3 |  |  | 0.17 |
| 21 | Uncertainty of the absolute gain of the reference antenna | Rectangular | 0.30 | 3 |  |  | 0.00 |
| 22 | Uncertainty of the absolute gain of the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.01 |
| Combined standard uncertainty (1 σ) |  |  |  |  |  |  | 0.65 |
| Expanded uncertainty (1.96σ – confidence interval of 95%) |  |  |  |  |  |  | 1.27 |

- 3. Conclusion

This contribution proposed the uncertainty values to the EIRP measurement with Indoor Anechoic Chamber. The final expanded uncertainty (1.96σ - confidence interval of 95%) can be determined from the proposed values as 1.14dB for f≦3.0GHz and 1.27dB for 3.0GHz<f≦4.2GHz.

- 4. Text Proposal:

TR37.842 v1.11.0

--------------Start of text proposal-------------

###### 10.3.1.1.1.4 Uncertainty assessment

Table 10.3.1.1.1.4-1.1: EIRP measurement uncertainty with the Indoor Anechoic Chamber

(f≦3.0GHz)

| UID | Description of uncertainty contribution | Uncertainty distribution | Maximum value of the uncertainty |  | Divisor | Standard uncertainty |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 1, DUT measurement |  |  |  |  |  |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | Rectangular | 0.03 | 3 |  |  | 0.02 |
| 2 | Pointing misalignment between the AAS BS and the receiving antenna. | Rectangular | Horizontal: 0.31 Vertical: 0.21 | 3 |  |  | 0.22 |
| 3 | Quality of quiet zone | Gaussian | Not applicable | Not applicable |  |  | 0.10 |
| 4 | Polarization mismatch between the AAS BS and the receiving antenna | Rectangular | 0.01 | 3 |  |  | 0.01 |
| 5 | Mutual coupling between the AAS BS and the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.00 |
| 6 | Phase curvature | Gaussian | Not applicable | Not applicable |  |  | 0.05 |
| 7 | Uncertainty of the measurement equipment | Gaussian | Not applicable | Not applicable |  |  | 0.32 |
| 8 | Impedance mismatch in the receiving chain | U | 0.20 | 2 |  |  | 0.14 |
| 9 | Random uncertainty | Gaussian | Not applicable | Not applicable |  |  | 0.20 |
| Stage 2, Calibration measurement |  |  |  |  |  |  |  |
| 10 | Impedance mismatch between the receiving antenna and the network analyzer | U | 0.05 | 2 |  |  | 0.04 |
| 11 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | Rectangular | 0.01 | 3 |  |  | 0.01 |
| 12 | Impedance mismatch between the reference antenna and the network analyzer. | U | 0.05 | 2 |  |  | 0.04 |
| 13 | Quality of quiet zone | Gaussian | Not applicable | Not applicable |  |  | 0.10 |
| 14 | Polarization mismatch for reference antenna | Rectangular | 0.01 | 3 |  |  | 0.01 |
| 15 | Mutual coupling between the reference antenna and the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.00 |
| 16 | Phase curvature | Gaussian | Not applicable | Not applicable |  |  | 0.05 |
| 17 | Uncertainty of the network analyzer | Gaussian | Not applicable | Not applicable |  |  | 0.10 |
| 18 | Influence of the reference antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | Rectangular | 0.05 | 3 |  |  | 0.03 |
| 19 | Reference antenna feed cable loss measurement uncertainty | Gaussian | Not applicable | Not applicable |  |  | 0.06 |
| 20 | Influence of the receiving antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | Rectangular | 0.05 | 3 |  |  | 0.03 |
| 21 | Uncertainty of the absolute gain of the reference antenna | Rectangular | 0.50 | 3 |  |  | 0.29 |
| 22 | Uncertainty of the absolute gain of the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.00 |
| Combined standard uncertainty (1 σ) |  |  |  |  |  |  | 0.58 |
| Expanded uncertainty (1.96σ – confidence interval of 95%) |  |  |  |  |  |  | 1.14 |

Table 10.3.1.1.1.4-1.2: EIRP measurement uncertainty with the Indoor Anechoic Chamber

(3.0GHz<f≦4.2GHz)

| UID | Description of uncertainty contribution | Uncertainty distribution | Maximum value of the uncertainty |  | Divisor | Standard uncertainty |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 1, DUT measurement |  |  |  |  |  |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | Rectangular | 0.03 | 3 |  |  | 0.02 |
| 2 | Pointing misalignment between the AAS BS and the receiving antenna. | Rectangular | Horizontal: 0.31 Vertical: 0.21 | 3 |  |  | 0.22 |
| 3 | Quality of quiet zone | Gaussian | Not applicable | Not applicable |  |  | 0.10 |
| 4 | Polarization mismatch between the AAS BS and the receiving antenna | Rectangular | 0.01 | 3 |  |  | 0.01 |
| 5 | Mutual coupling between the AAS BS and the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.00 |
| 6 | Phase curvature | Gaussian | Not applicable | Not applicable |  |  | 0.05 |
| 7 | Uncertainty of the measurement equipment | Gaussian | Not applicable | Not applicable |  |  | 0.45 |
| 8 | Impedance mismatch in the receiving chain | U | 0.33 | 2 |  |  | 0.23 |
| 9 | Random uncertainty | Gaussian | Not applicable | Not applicable |  |  | 0.20 |
| Stage 2, Calibration measurement |  |  |  |  |  |  |  |
| 10 | Impedance mismatch between the receiving antenna and the network analyzer | U | 0.05 | 2 |  |  | 0.01 |
| 11 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | Rectangular | 0.01 | 3 |  |  | 0.04 |
| 12 | Impedance mismatch between the reference antenna and the network analyzer. | U | 0.05 | 2 |  |  | 0.10 |
| 13 | Quality of quiet zone | Gaussian | Not applicable | Not applicable |  |  | 0.01 |
| 14 | Polarization mismatch for reference antenna | Rectangular | 0.01 | 3 |  |  | 0.00 |
| 15 | Mutual coupling between the reference antenna and the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.05 |
| 16 | Phase curvature | Gaussian | Not applicable | Not applicable |  |  | 0.11 |
| 17 | Uncertainty of the network analyzer | Gaussian | Not applicable | Not applicable |  |  | 0.03 |
| 18 | Influence of the reference antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | Rectangular | 0.05 | 3 |  |  | 0.06 |
| 19 | Reference antenna feed cable loss measurement uncertainty | Gaussian | Not applicable | Not applicable |  |  | 0.03 |
| 20 | Influence of the receiving antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | Rectangular | 0.05 | 3 |  |  | 0.17 |
| 21 | Uncertainty of the absolute gain of the reference antenna | Rectangular | 0.30 | 3 |  |  | 0.00 |
| 22 | Uncertainty of the absolute gain of the receiving antenna | Rectangular | 0.00 | 3 |  |  | 0.01 |
| Combined standard uncertainty (1 σ) |  |  |  |  |  |  | 0.65 |
| Expanded uncertainty (1.96σ – confidence interval of 95%) |  |  |  |  |  |  | 1.27 |

--------------End of text proposal-------------

References

[1] R4-161372, TP for TR 37.842: Adding EIRP test methods in Section 10, Ericsson, SEI, NTT DOCOMO, NEC, Kathrein, MVG.

[2] TR37.842 v1.11.0

[3] R4-161631, TP for TR37.842: Adding uncertainty value to the EIRP measurement with Indoor Anechoic Chamber, Sumitomo Elec. Industries, NTT DOCOMO.

[4] R4-162496, TR 37.842: Indoor Anechoic Chamber EIRP testing uncertainty value, NEC.

[5] R4-163751, TP for TR 37.842: Indoor Anechoic Chamber EIRP testing procedure, Sumitomo Electric, NTT DOCOMO, NEC.