3GPP TSG-RAN WG4 meeting #79	R4-163751

Nanjing, China, 23 - 27 May, 2016

Source:	Sumitomo Electric, NTT DOCOMO, NEC

Title:	TP for TR 37.842: Indoor Anechoic Chamber EIRP testing procedure

Agenda item:	6.2.2.2 OTA testing

Document for:	Approval

Introduction

In the RAN4#78 in Malta, a combined TP for adding EIRP test methods in TR37.842 was agreed [1].

In the agreed TP, four methods to test EIRP were introduced. They were “indoor anechoic chamber”, “compact antenna test range”, “1D compact range”, and “near field test range”.

In this contribution, we provide updated text for “indoor anechoic chamber” method. The major change is related to the calibration stage. In the current TR [2], signal generator and spectrum analyzer is used at the calibration stage for the indoor anechoic chamber far field method. In the last RAN4#78bis held in San Jose del Cabo, after offline discussion with other companies, it was agreed that it is better to use the network analyzer for calibration in order to be aligned with other test methods. Thus, in the TP we change the test procedure and uncertainty contribution factors following the network analyzer calibration method. Some notations are also improved for easier reading.

References

[1] R4-161372, TP for TR 37.842: Adding EIRP test methods in Section 10, Ericsson, SEI, NTT DoCoMo, NEC, Kathrein, MVG.

[2] TR37.842, v1.11.0

Text Proposal

---------------------------------- Start of the TP---------------------------------------------------

10.3.1 Radiated transmit power

10.3.1.1 Test methods

##### 10.3.1.1.1 Indoor Anechoic Chamber

<Title to be updated once the test method is identified>

###### 10.3.1.1.1.1 Description

This method measures the EIRP in an anechoic chamber with the separation between the manufacturer declared coordinate system reference point of the AAS BS and the phase center of the receiving antenna of no less than 2D2/λ, where D is the largest dimension of the antenna of AAS BS and λ is the wavelength. The measurement system setup is as depicted in Fig. 10.3.1.1.1.1-1.

Fig. 10.3.1.1.1.1-1 Measurement system setup.

###### 10.3.1.1.1.2 Procedure

The test consists of two stages: the calibration and the measurement. The test procedure is as follows:

Fig. 10.3.1.1.1.2-1 Calibration system setup.

Calibration:

  - Connect the reference antenna and the receiving antenna to the RF out port and RF in port of the network analyzer, respectively, as shown in Fig. 10.3.1.1.1.2-1.

  - Install the reference antenna with its beam peak direction and the height of its phase centre aligned with the receiving antenna.

  - Set the center frequency of the network analyzer to the carrier centre frequency of the tested signal for EIRP measurement of AAS BS and measure LFEIRP, E→D, which is equivalent to 20log|S21| (dB) obtained by the network analyzer.

LFEIRP, E→D: Pathloss between E and D in Fig. 10.3.1.1.1.2-1.

  - Measure the cable loss, LFEIRP, E→F,between the reference antenna connector and the network analyzer connector.

LFEIRP, E→F: Cable loss between E and F in Fig. 10.3.1.1.2-1

  - Calculate the calibration value between A and D with the following formula:

LEIRP_cal, A→D = LFEIRP, E→D  + GREF_ANT_EIRP, A→F –LFEIRP, E→F

LEIRP_cal, A→D:  Calibration value between  A and D in Fig. 10.3.1.1.1.2-1.

GREF_ANT_EIRP, A→F: Antenna gain of the reference antenna

Measurement:

- Uninstall the reference antenna and install the AAS BS with the manufacturer declared coordinate system reference point in the same place as the phase centre of the reference antenna. The manufacturer declared coordinate system orientation of the AAS is set to be aligned with the testing system.

- Set the AAS BS to generate the tested beam with the beam peak direction intended to be the same as the testing direction.

- Rotate the AAS BS to make the testing direction align with the direction of the receiving antenna.

- Set the AAS BS to transmit the test signal at the maximum power according to E-TM1.1.

- Measure the mean power for each carrier arriving at the measurement equipment connector, denoted by PR_AAS_EIRP, D.

PR_AAS_EIRP, D: Measured mean power for each carrier at the measurement equipment connector at D in Fig. 10.3.1.1.1.1-1

- Calculate the EIRP with the following formula:

EIRP = PR_AAS_EIRP, D- LEIRP_cal, A→D

- Repeat above step 7) ~ 11) per conformance test beam direction.

###### 10.3.1.1.1.3 Uncertainty budget format

Table 10.3.1.1.1.3 -1: Indoor Anechoic Chamber Uncertainty contributions

for AAS BS EIRP measurement

| UID | Description of uncertainty contribution | Details in paragraph |
| --- | --- | --- |
| Stage 1, DUT measurement |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | B1-1 |
| 2 | Pointing misalignment between the AAS BS and the receiving antenna. | B1-2 |
| 3 | Quality of quiet zone | B1-3 |
| 4 | Polarization mismatch between the AAS BS and the receiving antenna | B1-4 |
| 5 | Mutual coupling between the AAS BS and the receiving antenna | B1-5 |
| 6 | Phase curvature | B1-6 |
| 7 | Uncertainty of the measurement equipment | B1-7 |
| 8 | Impedance mismatch in the receiving chain | B1-8 |
| 9 | Random uncertainty | B1-9 |
| Stage 2, Calibration measurement |  |  |
| 10 | Impedance mismatch between the receiving antenna and the network analyzer | B1-10 |
| 11 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | B1-11 |
| 12 | Impedance mismatch between the reference antenna and the network analyzer. | B1-12 |
| 13 | Quality of quiet zone | B1-3 |
| 14 | Polarization mismatch for reference antenna | B1-4 |
| 15 | Mutual coupling between the reference antenna and the receiving antenna | B1-5 |
| 16 | Phase curvature | B1-6 |
| 17 | Uncertainty of the network analyzer | B1-13 |
| 18 | Influence of the reference antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | B1-14 |
| 19 | Reference antenna feed cable loss measurement uncertainty | B1-15 |
| 20 | Influence of the receiving antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | B1-16 |
| 21 | Uncertainty of the absolute gain of the reference antenna | B1-17 |
| 22 | Uncertainty of the absolute gain of the receiving antenna | B1-18 |

10.3.1.1.1.4 Uncertainty assessment

<Texts to be added>

---------------------------------- Unchanged part omitted-----------------------------------

Annex B Radiated transmit power measurement error contribution descriptions

B.1 Indoor Anechoic Chamber

B1-1 Positioning misalignment between the AAS BS and the reference antenna

This contribution originates from the misalignment of the manufacturer declared coordinate system reference point of the AAS BS and the phase centre of the reference antenna. The uncertainty makes the space propagation loss between the AAS BS and the receiving antenna at the DUT measurement stage different from the space propagation loss between the reference antenna and the receiving antenna at the calibration stage.

B1-2 Pointing misalignment between the AAS BS and the receiving antenna.

This contribution originates from the misalignment of the testing direction and the beam peak direction of the receiving antenna due to imperfect rotation operation. The pointing misalignment may happen in both azimuth and vertical directions and the effect of the misalignment depends highly on the beamwidth of the beam under test. The same level of misalignment results in a larger measurement error for a narrower beam.

B1-3 Quality of quiet zone

This contribution originates from a reflectivity level of an anechoic chamber. The reflectivity level is determined from the average standard deviation of the electric field in the quiet zone. By repeating a free space VSWR measurement in 15-degree grid in elevation and azimuth, 264 standard deviation values in both polarizations are determined. From these values an average standard deviation of electric field in the quiet zone can be calculated from the equation:

where

is the number of angular intervals in elevation,

is the number of angular intervals in azimuth and

is elevation of single measurement .

If an efficiency calibration with omnidirectional calibration antenna is performed, the effect of reflectivity level decreases in Stage 2 and  may be divided by factor 2. This is due to correcting impact of data averaging in this type of calibration. Efficiency calibration done with sampling step ≤ 30°, can be considered to have at least four independent samples.  may be divided by factor 2 also in stage 1 for the same reason.

It’s likely that asymmetry of the field probe will have a very small impact on this measurement uncertainty contributor, however, an upper bound to probe symmetry should be considered.

B1-4 Polarization mismatch between the AAS BS/reference antenna and the receiving antenna

This contribution originates from the misaligned polarization between the AAS BS/reference antenna and the receiving antenna.

B1-5 Mutual coupling between the AAS BS/reference antenna and the receiving antenna

This contribution originates from mutual coupling between the AAS BS/reference antenna and the receiving antenna. Mutual coupling degrades not just the antenna efficiency, i. e. the EIRP value, but it can alter the antenna’s radiation pattern as well.  For indoor anechoic chamber, usually the spacing between the AAS BS/reference antenna and the receiving antennas is large enough so that the level of mutual coupling might be negligible.

B1-6 Phase curvature

This contribution originates from the finite far-field measurement distance, which causes phase curvature across the antenna of AAS BS/reference antenna.

B1-7 Uncertainty of the measurement equipment

This contribution originates from limited absolute level accuracy and non-linearity of the measurement equipment. The measurement equipment such as a Base Station Simulator (BSS), spectrum analyzer (SA), or power meter (PM) measures the received signal level in EIRP tests either as an absolute level or as a relative level. The uncertainty value will be indicated in the manufacturer’s data sheet in logs.

B1-8 Impedance mismatch in the receiving chain

This contribution originates from multiple reflections between the receiving antenna and the measurement equipment. The multiple reflections can produce an overall reflection that depends not only upon the individual reflections of each part but their reflective interactions as well. The combination loss by the overall reflection can be higher or lower than individual loss by multiple reflections. The combination loss is called the mismatch error and leads to the measurement uncertainty.

B1-9 Random uncertainty

The random uncertainty characterizes the undefined and miscellaneous effects which cannot be forecasted. One can estimate this type of uncertainty with a repeatability test by making a series of repeated measurement with a reference DUT without changing anything in the measurement set-up.

B1-10 Impedance mismatch between the receiving antenna and the network analyzer

This contribution originates from multiple reflections between the receiving antenna and the network analyzer. After appropriate calibration, the network analyzer may not introduce impedance mismatch error, but the error still happens between the receiving antenna feed cable and the receiving antenna.

B1-11 Positioning and pointing misalignment between the reference antenna and the receiving antenna

This contribution originates from reference antenna alignment and pointing error.  In this measurement if the maximum gain directions of the reference antenna and the receiving antenna are aligned each other, this contribution can be considered negligible and therefore set to zero.

B1-12 Impedance mismatch between the reference antenna and the network analyzer

This contribution originates from multiple reflections between the reference antenna and the network analyzer. After appropriate calibration, the network analyzer may not introduce impedance mismatch error, but the error still happens between the reference antenna feed cable and the reference antenna.

B1-13 Uncertainty of the network analyzer

This contribution originates from all uncertainties involved in the S21 measurement (including drift and frequency flatness) with a network analyzer. The uncertainty value will be indicated in the manufacturer’s data sheet in logs.

B1-14 Influence of the reference antenna feed cable

In the calibration stage, the influence of the calibration antenna feed cable shall be assessed by measurements. A measurement for calibration may be repeated with a reasonably differing routing of the feed cable. Largest difference among the results is entered to the uncertainty budget with a rectangular distribution.

B1-15 Reference antenna feed cable loss measurement uncertainty

Before performing the calibration, the reference antenna feed cable loss have to be measured. The measurement can be done with a network analyzer to measure its S21 and uncertainty is introduced.

B1-1 6Influence of the receiving antenna feed cable

If the probe antenna is directional (i.e. peak gain >+5dBi e.g. horn, LPDA, etc.) and the same probe antenna cable configuration is used for both stages, the uncertainty is considered systematic and constant  0.00dB value.

In other cases a technical study should be done.

B1-17 Uncertainty of the absolute gain of the reference antenna

This contribution originates from the gain/efficiency uncertainty of calibration antenna that is employed in Stage 2 only.

This uncertainty value will be indicated in the manufacturer’s data in logs with a rectangular distribution, unless otherwise informed.

B1-18 Uncertainty of the absolute gain of the receiving antenna

The uncertainty appears in both stages and it is thus considered systematic and constant  0.00dB value.

---------------------------------- End of the TP---------------------------------------------------