3GPP TSG- RAN WG4 Meeting #87 	R4-1806588

# Busan, Korea, 21 – 25 May 2018

| CR-Form-v11.2 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 37.843 | CR | CRNum | rev | - | Current version: | 15.0.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network | X | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | Draft CR to TR 37.843: AAS OTA sensitivity conformance testing with Indoor Anechoic Chamber (10.3.2.2) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Nokia, Nokia Shanghai Bell |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R4 |  |  |  |  |  |  |  |  |  |
| Work item code: | AASenh_BS_LTE_UTRA-Perf |  |  |  |  |  |  | Date: |  | 2018-05-14 |
| Category: | B |  |  |  |  |  |  | Release: |  | Rel-15 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) Rel-12 (Release 12) Rel-13 (Release 13) Rel-14 (Release 14) Rel-15 (Release 15) Rel-16 (Release 16) |  |
| Reason for change: |  | The measurement uncertainty budget and value for Rel-13 AAS BS OTA sensitivity conformance testing with Indoor Anechoic Chamber have been recorded in TR 37.842, and they can be reused in TR 37.843 for Rel-15 eAAS BS. |  |  |  |  |  |  |  |  |
| Summary of change: |  | Record the measurement uncertainty budget and value for Rel-15 eAAS BS OTA sensitivity conformance testing with Indoor Anechoic Chamber in TR 37.843, according to those in TR 37.842. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | The measurement uncertainty budget and value for Rel-15 eAAS BS OTA sensitivity conformance testing with Indoor Anechoic Chamber in TR 37.843 remain missing. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | 10.3.2.2 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  | TS/TR ... CR ... |  |  |  |
| affected: |  |  | X | Test specifications |  |  | TS/TR ... CR ... |  |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  | TS/TR ... CR ... |  |  |  |
| Other comments: |  |  |  |  |  |  |  |  |  |  |

<Start of change>

#### 10.3.2.2	In-door anechoic chamber

##### 10.3.2.2.1	General

This method measures the OTA sensitivity in an anechoic chamber with the separation between the manufacturer declared coordinate system reference point of the AAS BS and the phase centre of the transmitting antenna of no less than 2D2/λ, where D is the largest dimension of the antenna of AAS BS and λ is the wavelength. The measurement system setup is as depicted in figure 10.3.2.2.1-1.

Figure 10.3.2.2.1-1: Indoor Anechoic Chamber measurement system setup

##### 10.3.2.2.2	Calibration

Figure 10.3.2.2.2-1: Indoor Anechoic Chamber calibration system setup

1)  Connect the reference antenna and the transmitting antenna to RF in port and RF out port of the network analyzer, respectively, as shown in figure 10.3.2.2.2-1.

2)	Install the reference antenna with its beam peak direction and the height of its phase centre aligned with the transmitting antenna.

3)	Set the centre frequency of the network analyzer to the carrier centre frequency of the tested signal for EIS measurement of AAS BS and measure LFRX_E→D, which is equivalent to 20log|S21| (dB) obtained by the network analyzer:

-	LFRX_ E→D: Pathloss between E and D in figure 10.3.2.2.2-1.

4)	Measure the cable loss, LFRX_ E→F, between the reference antenna connector and the network analyzer connector:

-	LFRX_ E→F: Cable loss between E and F in figure 10.3.2.2.2-1.

5)	Calculate the calibration value between A and D with the following formula:

-	LRX_cal_A→D = LFRX_ E→D + GREF_ANT_EIS, A→F - LFRX_ E→F.

-	LRX_cal_A→D:  Calibration value between  A and D in figure 10.3.2.2.2-1.

-     GREF_ANT_EIS, A→F: Antenna gain of the reference antenna.

##### 10.3.2.2.3 	Procedure

1)	Uninstall the reference antenna and install the AAS with its manufacturer declared coordinate system reference point in the same place as the phase centre of the reference antenna. The manufacturer declared coordinate system orientation of the AAS is set to be aligned with testing system.

2)	Set the AAS BS to comply with the RoAoA desired by the requirement covering conformance testing receiving direction.

3)	Rotate the AAS BS to make the conformance testing at receiving direction aligned with the beam peak direction of the reference antenna at the calibration stage.

4)	For FDD AAS BS start BS transmission at the required condition.

5)	Set the test signal mean power at the RF signal source generator on the required level plus LRX_cal_A→D and the reference measurement channel.

6)	Measure the throughput according to annex E of the 3GPP TS 36.141 [22].

7)	Repeat the above steps 1)~6) per conformance testing direction.

For the OTA sensitivity,

The required level in step 5) is the minimum EIS level

##### 10.3.2.2.4 	MU assessment

###### 10.3.2.2.4.1 	MU Budget

Table 10.3.2.2.4.1-1: Indoor Anechoic Chamber uncertainty contributions
for AAS BS EIS measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | C1-1 |
| 2 | Pointing misalignment between the AAS BS and the transmitting antenna. | C1-2 |
| 3 | Quality of quiet zone | C1-3 |
| 4 | Polarization mismatch between the AAS BS and the transmitting antenna | C1-4 |
| 5 | Mutual coupling between the AAS BS and the transmitting antenna | C1-5 |
| 6 | Phase curvature | C1-6 |
| 7 | Uncertainty of the RF signal source generator | E |
| 8 | Impedance mismatch in the transmitting chain | C1-8 |
| 9 | Random uncertainty | C1-9 |
| Stage 1: Calibration measurement |  |  |
| 10 | Impedance mismatch between the transmitting antenna and the network analyzer | C1-10 |
| 11 | Positioning and pointing misalignment between the reference antenna and the transmitting antenna | C1-11 |
| 12 | Impedance mismatch between the reference antenna and network analyzer. | C1-12 |
| 13 | Quality of quiet zone | C1-3 |
| 14 | Polarization mismatch for reference antenna | C1-4 |
| 15 | Mutual coupling between the reference antenna and the transmitting antenna | C1-5 |
| 16 | Phase curvature | C1-6 |
| 17 | Uncertainty of the network analyzer | E |
| 18 | Influence of the reference antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | C1-14 |
| 19 | Reference antenna feed cable loss measurement uncertainty | C1-15 |
| 20 | Influence of the transmitting antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | C1-16 |
| 21 | Uncertainty of the absolute gain of the reference antenna | E |
| 22 | Uncertainty of the absolute gain of the transmitting antenna | C1-18 |

###### 10.3.2.2.4.2 	MU Value

# Table 10.3.2.2.4.2-1: Indoor Anechoic Chamber uncertainty assessment for EIS measurement

| UID | Uncertainty source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.02 | 0.02 |
| 2 | Pointing misalignment between the AAS BS and the transmitting antenna | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 | 0.17 |
| 3 | Quality of quiet zone | 0.1 | 0.1 | Gaussian | 1.00 | 1 | 0.10 | 0.10 |
| 4 | Polarization mismatch between the AAS BS and the transmitting antenna | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 |
| 5 | Mutual coupling between the AAS BS and the transmitting antenna | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 6 | Phase curvature | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 | 0.05 |
| 7 | Uncertainty of the RF signal generator | 0.46 | 0.46 | Gaussian | 1.00 | 1 | 0.46 | 0.46 |
| 8 | Impedance mismatch in the transmitting chain | 0.14 | 0.23 | U-shaped | 1.41 | 1 | 0.10 | 0.16 |
| 9 | Random uncertainty | 0.1 | 0.1 | Rectangular | 1.73 | 1 | 0.06 | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 10 | Impedance mismatch between the transmitting antenna and the network analyzer | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 | 0.04 |
| 11 | Positioning and pointing misalignment between the reference antenna and the transmitting antenna | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 |
| 12 | Impedance mismatch between the reference antenna and network analyzer | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 | 0.04 |
| 13 | Quality of quiet zone | 0.1 | 0.1 | Gaussian | 1.00 | 1 | 0.10 | 0.10 |
| 14 | Polarization mismatch for reference antenna | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 | 0.01 |
| 15 | Mutual coupling between the reference antenna and the transmitting antenna | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 16 | Phase curvature | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 | 0.05 |
| 17 | Uncertainty of the network analyzer | 0.13 | 0.20 | Gaussian | 1.00 | 1 | 0.13 | 0.20 |
| 18 | Influence of the reference antenna feed cable | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 |
| 19 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | Gaussian | 1.00 | 1 | 0.06 | 0.06 |
| 20 | Influence of the transmitting antenna feed cable | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 |
| 21 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.25 | Rectangular | 1.73 | 1 | 0.29 | 0.25 |
| 22 | Uncertainty of the absolute gain of the transmitting antenna | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.62 | 0.64 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.22 | 1.25 |

<End of change>