## Slide 1: WF on Maximum Test System Uncertainty Evolution for FR2
Keysight Technologies, Rohde &Schwarz, Anritsu, Samsung, MVG, ETS-Lindgren, PCTEST Engineering Lab, Ericsson
3GPP TSG-RAN5 Meeting #82 R5-192820 Athens, Greece, 25 Feb - 1 Mar 2019

## Slide 2: Definition: Measurement Uncertainty
Measurement Uncertainty is “non-negative parameter characterizing the dispersion of the quantity values being attributed to a measurand, based on the information used” (per JCGM 200:2012)
For the power-based EIRP, EIS, TRP, TRS test cases 3GPP and CTIA follow the “Guide to the Expression of Uncertainty in Measurement” published by the International Organization for Standardization (ISO) Geneva, Switzerland 1995
For every test case, a detailed and exhaustive MU assessment needs to be performed

## Slide 3: Sample
Sample MU Table for Test Case X.Y.
UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB]
Stage 2: DUT measurement |  |  |  |  | 
1 | MU Element 1 | MU1 | {rect, u-shape, normal, std} | {√3,√2,2,1} | StD1
2 | MU Element 2 | MU2 | {rect, u-shape, normal, std} | {√3,√2,2,1} | StD2
... |  |  | {rect, u-shape, normal, std} | {√3,√2,2,1} | 
i | MU Element i | MUi | {rect, u-shape, normal, std} | {√3,√2,2,1} | StDi
Stage 1: Calibration measurement |  |  |  |  | 
i+1 | MU Element i+1 | MUi+1 | {rect, u-shape, normal, std} | {√3,√2,2,1} | StDi+1
... |  |  | {rect, u-shape, normal, std} | {√3,√2,2,1} | 
N-1 | MU Element N-1 | MUN-1 | {rect, u-shape, normal, std} | {√3,√2,2,1} | StDN-1
 | Systematic uncertainties |  |  |  | Value
N | MU Element N |  |  |  | SUN
Total measurement uncertainty |  |  |  |  | Value
Total Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 
Total Expanded uncertainty =
Expanded uncertainty

## Slide 4: Definition: MU Element
For each test case, MU elements are identified that can have an impact on the measured quantity
For each MU element, an uncertainty value is determined based on statistical methods of theoretical analyses and measurements.
  proper documentation and evidence must be presented to be included for each MU element, e.g., experimental data, simulation results, datasheets, research papers, etc.
Generally, for OTA tests the MU elements are tabulated in informative sections of Technical Specifications (TS) or in Technical Reports (TR)
  Tabulated separately for the DUT measurement stage and for the calibration measurement stage

## Slide 5: Definition: Probability Distribution&Divisor
For each MU element, the probability distribution needs to be determined which is based on the method the example value has been determined
Each distribution has its own distribution divisor
  Rectangular  Divisor: √3
  U-shaped  Divisor: √2
  Normal  Divisor: 2
  Standard  Divisor: 1

## Slide 6: Definition: Standard Uncertainty s
For each element, the standard uncertainty, s, is determined by dividing the uncertainty value by the probability distribution divisor

## Slide 7: Confidence Interval
The x% confidence interval is a range of values that you can be x% certain contains the true mean of the population
  For RF measurements, a 95% confidence interval is typical

## Slide 8: (Total) Expanded Uncertainty
Once all MU elements have been defined, the individual standard uncertainties are RSS’ed, i.e., square root of the squared sum of each standard uncertainty.
The expanded uncertainty without systematic errors is the result of the total standard uncertainty multiplied by the confidence level factor of K=1.96 for 95% confidence level
Any systematic errors are added to the expanded uncertainty to yield the total expanded uncertainty
Total Expanded uncertainty =
Expanded uncertainty

## Slide 9: MU Threshold (TR38.903 and TR38.810)
The MU Threshold, MUThreshold, is used to determine a permitted test methodology
  If the expanded MU in TR38.903 for a given test methodology is below the MU threshold, it is a permitted test methodology

## Slide 10: Maximum Test System Uncertainty
The Maximum Test System Uncertainty, MTSU, is used
  To determine the Test Tolerance (TT) that test requirements are relaxed by, i.e., 0≤TT≤MTSU
  To determine whether a test system installed at a test lab and whose MUTS has been assessed can be used for certification purposes, i.e., when MUTS≤MTSU, the test system can be used for certification purposes
Reducing the Maximum Test System Uncertainty over time could result in authorized labs to become unauthorized
Maximum Test System Uncertainties for conducted and OTA test cases have traditionally not evolved over time
The Maximum Test System Uncertainty for FR2 OTA test cases will be captured in Annex F of TS38.521-2 (SA) and TS38.521-3 (NSA)

## Slide 11: Maximum Test System Uncertainty Framework
Currently, there is no framework in place on how to determine the Maximum Test System Uncertainty for FR2 OTA test cases
Test Specifications TS37.544 and TS34.114 are studied to determine how Maximum Test System Uncertainty has been defined as a function of the Expanded MU for OTA test cases

## Slide 12: Maximum Test System Uncertainty for OTA TCs
The vast majority of OTA test cases defined ratios of 1 and greater
Spec | Test Case | RAT | Phantom | Expanded Uncertainty [dB] | Maximum Test System Uncertainty [dB] MTSU | Ratio of MTSU to Expanded Uncertainty (95%-tile) | Comments
37.544 | TRP | UTRA | Head only | 1.75 | 1.9 | 1.09 | Note 1
37.544 | TRP | UTRA | Head and hand | 2.1 | 1.9 | 0.90 | Note 2
37.544 | TRP (LEE) | E-UTRA | ground plane | 1.75 | 1.9 | 1.09 | 
37.544 | TRS | UTRA | Head only | 2.16 | 2.3 | 1.06 | Note 1
37.544 | TRS | UTRA | Head and hand | 2.64 | 2.3 | 0.87 | Note 2
37.544 | TRS (LEE) | E-UTRA | ground plane | 2.16 | 2.3 | 1.06 | 
37.544 | TRMS | E-UTRA | free space | 2.65 | 2.65 | 1.00 | 
34.114 | TRP | UMTS | Head only | 1.86 | 1.9 | 1.02 | 
34.114 | TRP | GSM | Head only | 1.86 | 1.9 | 1.02 | 
34.114 | TRP | TD-SCDMA | Head only | 1.86 | 1.9 | 1.02 | 
34.114 | TRP (LME) | UMTS | ground plane | 1.75 | 1.9 | 1.09 | 
34.114 | TRP (LME) | GSM | ground plane | 1.75 | 1.9 | 1.09 | 
34.114 | TRP (LME) | TD-SCDMA | ground plane | 1.75 | 1.9 | 1.09 | 
34.114 | TRP (LEE) | UMTS | ground plane | 1.75 | 1.9 | 1.09 | 
34.114 | TRP (LEE) | GSM | ground plane | 1.75 | 1.9 | 1.09 | 
34.114 | TRP (LEE) | TD-SCDMA | ground plane | 1.75 | 1.9 | 1.09 | 
34.114 | TRS | UMTS | Head only | 2.29 | 2.3 | 1.00 | 
34.114 | TRS | GSM | Head only | 2.29 | 2.3 | 1.00 | 
34.114 | TRS | TD-SCDMA | Head only | 2.29 | 2.3 | 1.00 | 
34.114 | TRS (LME) | UMTS | ground plane | 2.16 | 2.3 | 1.06 | 
34.114 | TRS (LME) | GSM | ground plane | 2.16 | 2.3 | 1.06 | 
34.114 | TRS (LME) | TD-SCDMA | ground plane | 2.16 | 2.3 | 1.06 | 
34.114 | TRS (LEE) | UMTS | ground plane | 2.16 | 2.3 | 1.06 | 
34.114 | TRS (LEE) | GSM | ground plane | 2.16 | 2.3 | 1.06 | 
34.114 | TRS (LEE) | TD-SCDMA | ground plane | 2.16 | 2.3 | 1.06 | 
Note 1: introduced in TS37.544 V0.6.1 (2015-11) together with expanded MUs for BH and BHH |  |  |  |  |  |  | 
Note 2: Maximum Test System Uncertainty introduced in TS37.544 V14.2.0 (2017-06), added by Intel with R5-173368. BHH Maximum Test System Uncertainty communicated in LS from RAN4 to RAN5 in R4-1704067: "RAN4 assumed the same test tolerance and measurement uncertainty levels as earlier for developing both LTE tablet TRP and TRS requirement and for developing UMTS BHH handheld TRP and TRS requirements, as reported in table below" |  |  |  |  |  |  | 

## Slide 13: Evolution of MU vs MTSU
TE Vendors are committed to improving the MU of test systems over time by evolution in technology, HW, and SW
Reducing the Maximum Test System Uncertainty, MTSU, over time could have significant consequences for labs with existing test systems by getting unauthorized. Potentially significant investments might be required to get authorized again

## Slide 14: Evolution of Total Expanded Uncertainty
For the last 2 years, TE and TS vendors have been working on MU assessments in RAN4 and later in RAN5.
Analyses were based on
  state of the art equipment (highest performance) and components
  optimized algorithms and procedures (measurement grids, calibrations, etc.)
Group/TR | TRP | EIRP | EIS
RAN4 (TR38.810) | 5.13 | 5.99 | 6.49
RAN5 (TR38.903) | 4.29-4.69 | 4.38-4.78 | 5.02

## Slide 15: MTSU: Release Dependent vs Independent
The applicability of the MTSU on Releases can be handled in two ways: release dependent vs release independent
For example, assume the MTSU for MOP EIRP TC is determined to be 5.0dB and entered in F.1.2 of TS38.521-2 (Release 15). Since there no review of MTSU during Release 16, the MTSU for MOP EIRP TC remains 5.0dB in F.1.2 of TS38.521-2 (Release 16). Furthermore, assume that TE vendors find that the MTSU can be reduced to 4.0dB during the first review period at the introduction of TS38.521-2 (Release 17). This evolved MTSU for EIRP TC is updated in F.1.2 of TS38.521-2 (Release 17) only.
  In release 17 time frame, when MTSU is considered release independent, an MTSU of 4dB will be applicable for Release 15 through 17 UEs for the MOP EIRP TCs
  In release 17 time frame, when MTSU is considered release dependent, an MTSU of 5dB will be applicable for Release 15 through 16 UEs and an MTSU of 4dB will be applicable for Release 17 (and beyond) UEs for MOP EIRP TCs
The need for updating and re-validating existing test cases with reduced TT (based on evolved MTSU) needs to be discussed separately

## Slide 16: Way Forward
Proposal 1: For UE RF testing, The Maximum Test System Uncertainty for FR2 in TS38.521-2 and TS38.521-3 Annex F.1.2 & F.1.3 shall correspond to Total Expanded Uncertainty*1.00.
Proposal 2: For UE Demodulation and RRM testing, the Maximum Test System Uncertainty for FR2 in TS38.521-4 and TS38.533 shall correspond to Total Expanded Uncertainty*1.00.
Proposal 3: Based on industry input, RAN5 shall review the Maximum Test System Uncertainty for FR2 defined in TS38.521-2 and TS38.521-3 with the goal to reduce them on a case-by-case basis.

## Slide 17: Way Forward
Proposal 4: To help facilitate any reduction, RAN5 agrees to start a study item for optimization of MU for FR2
  To allow for sufficient time in technology and process improvements with the intention to optimize the evolved MTSU rollouts in terms of device certification, the SI shall be created no earlier than RAN#90 (December 2020). The proposed target completion date (TCD) of the study item would be determined during the SI proposal stage with the goal to limit the duration of the SI to one year maximum.
  The study item would determine the achievable FR2 Maximum Test System Uncertainty improvement and would define the percentage improvement that would be required in order to justify any update to the Maximum Test System Uncertainty value.
  The SI proposal will define a goal for the MTSU improvement percentage(s)
  If required, the scope of the SI could be extended also to UE Demodulation and RRM based on industry input.
Proposal 5: The evolved FR2 MUs are release independent
Proposal 6: Enforcement dates of the revised FR2 Maximum Test System Uncertainty will be dictated by industry needs and be part of the outcome of the proposed SI

## Slide 18: References
NIS 81, “The Treatment of Uncertainty in EMC Measurements,” Ed. 1, NAMAS Executive, National Physical Laboratory, Teddington, Middlesex, TW11 0LW, England, 1994.
NIST, “Guidelines for evaluating and expressing the uncertainty of NIST measurement results”, Technical Note 1297 (TN 1297), United States Department of Commerce Technology Administration, National Institute of Standards and Technology, Gaithersburg, MD, 1994.
ISO/IEC 17025:2005: General requirements for the competence of testing and calibration laboratories.
JCGM 101:2008, Evaluation of measurement data — Supplement 1 to the “Guide to the expression of uncertainty in measurement” — Propagation of distributions using a Monte Carlo method
JCGM 200:2012, International vocabulary of metrology – Basic and general concepts and associated terms (VIM)