3GPP TSG RAN WG5 #83 Meeting	R5-194627

Reno, USA, 13th – 17th May 2019

Source:		Rohde & Schwarz

Title:	On testability of Rx spurious emission test case

Agenda Item:		5.3.5.17

Document for:	Information

# Introduction

In this contribution, we shed light on the testability of the Rx spurious emission test case.

# Discussion

The minimum conformance requirement for Rx spurious is specified in TS 38.101-2 [1] as follows

7.9	Spurious emissions

The spurious emissions power is the power of emissions generated or amplified in a receiver. The spurious emissions power level is measured as TRP.

The power of any narrow band CW spurious emission shall not exceed the maximum level specified in Table 7.9-1. The requirement is verified in beam locked mode with the test metric of TRP (Link=TX beam peak direction).

Table 7.9-1: General receiver spurious emission requirements

| Frequency range | Measurement bandwidth | Maximum level | NOTE |
| --- | --- | --- | --- |
| 30MHz f < 1GHz | 100 kHz | -57 dBm | 1 |
| 1GHz f 2nd harmonic of the upper frequency edge of the DL operating band in GHz | 1 MHz | -47 dBm |  |
| NOTE 1: Unused PDCCH resources are padded with resource element groups with power level given by PDCCH_RA/RB as defined in Annex C.3.1. |  |  |  |

Currently band n260 has the largest frequency resulting in a 2nd harmonic of 80 GHz and we conduct our analysis at this frequency. In this contribution, we focus on a CATR system with a focal length of 75cm since for already an antenna aperture of 5cm a DFF system obeying far-field condition has a 5 dB larger free space path loss. And for a DUT without declaration of antenna configuration and a size of 15cm, the path loss for DFF is worse by 24dB compared to an IFF system at 80 GHz.

For a typical CATR with a focal distance of 75 cm, the free space path loss at 80 GHz is roughly 68 dB. Assuming a simplified setup, where the feed antenna is connected via a short cable to a state of the art spectrum analyzer results in a SNR of -42.6 dB. Here, an optimistically short cable of 1m length with a loss of 11.3 dB has been assumed.

An alternate setup is the down-conversion of the signal directly after the measurement antenna. In this way, the above mentioned cable loss is avoided. However, the conversion loss of the mixer of approximately 28 dB has to be taken into account. Since the noise figure of the spectrum analyzer at the down-converted frequencies is lower than at 80 GHz, a SNR of -39.8 dB is obtained with an improvement of only 2.8 dB compared to the simplified setup.

Applying an additional non-bypass-able LNA between measurement antenna and mixer improves the SNR to a value of -14.8 dB. However, the associated impact of noise still results in a measurement uncertainty of at least 15 dB where all the other MU contributors have not been considered yet.

It has to be noted that in the above examples the ideal case of a setup dedicated only for spurious emissions in the frequency range of 60 GHz to 80 GHz has been considered. In the case of combining the setup with other spurious emission frequency ranges or an in-band measurement setup further degradation of the SNR is to be expected.

Observation 1: The SNR at 80 GHz for Rx Spurious Emissions is approximately -15 dB in an idealized case.

Our conclusion of the above analysis is that receiver spurious emissions cannot be tested without relaxation in a fully automated test system with state of the art test system components.

Observation 2: Rx spurious emissions cannot be tested without relaxation.

# Assumptions

| ID | Description | Assumption |
| --- | --- | --- |
| #1 | Frequency ranges under consideration | 6 – 80 GHz |
| #2 | Size of QZ for IFF | 30 cm |
| #3 | UE power class | PC3 |
| #5 | Temperature range of the test equipment | 20°C – 35°C |
| #6 | Channel bandwidth | 400 MHz |

# Conclusion

In summary, the testability of Rx spurious emissions has been analyzed.

Observation 1: The SNR at 80 GHz for Rx Spurious Emissions is approximately -15 dB in an idealized case

Observation 2: Rx spurious emissions cannot be tested without relaxation.

# References

- TS 38.101-2, V15.4.0, 3GPP TSG RAN WG4, December 2018

- TR 38.810, V16.1.0, 3GPP TSG RAN WG4, December 2018