3GPP TSG-RAN5 Meeting #89-e		R5-206154
Electronic meeting, 9 November - 20 November 2020

Agenda item:	5.3.2.17

Source:	Keysight Technologies

Title:	On DFF MU Tables in 38.903

Document for:	Discussion and Endorsement

# Introduction

This contribution describes and proposes Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power for DFF test methodology.

# Discussion

Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power is already is already concluded in TR 38.903 Annex E.3.1.3 for IFF test setup. Following MU elements are considered in IFF test setup:

| UID |  | Description of uncertainty contribution |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 |  | Positioning misalignment |  |  |  |
| 2 |  | Measure distance uncertainty |  |  |  |
| 3 |  | Quality of Quiet Zone |  |  |  |
| 4 |  | Mismatch |  |  |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  |  |  |
| 6 |  | gNB emulator uncertainty |  |  |  |
| 7 |  | Phase curvature |  |  |  |
| 8 |  | Amplifier uncertainties |  |  |  |
| 9 |  | Random uncertainty |  |  |  |
| 10 |  | Influence of the XPD |  |  |  |
| 11 |  | Insertion Loss Variation |  |  |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |
| 13 |  | Multiple measurement antenna uncertainty |  |  |  |
| 14 |  | DUT repositioning |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 |  | Mismatch |  |  |  |
| 16 |  | Amplifier Uncertainties |  |  |  |
| 17 |  | Misalignment of positioning System |  |  |  |
| 18 |  | Uncertainty of the Network Analyzer |  |  |  |
| 19 |  | Uncertainty of the absolute gain of the calibration antenna |  |  |  |
| 20 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |
| 21 |  | Phase centre offset of calibration antenna |  |  |  |
| 22 |  | Quality of quiet zone for calibration process |  |  |  |
| 23 |  | Standing wave between reference calibration antenna and measurement antenna |  |  |  |
| 24 |  | Influence of the calibration antenna feed cable |  |  |  |
| 25 |  | Insertion Loss Variation |  |  |  |
| Systematic uncertainties |  |  |  |  |  |
| 26 |  | Systematic error related to beam peak search |  |  |  |

Table 1: Uncertainty contributions for DL AWGN absolute power or wanted DL signal absolute power for IFF test setup

Analysing one by one each of IFF MU elements, two MU elements groups can be identified:

- Group 1: MU elements independent of test methodology, highlighted in green in table 2.

| UID |  | Description of uncertainty contribution |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 |  | Positioning misalignment |  |  |  |
| 2 |  | Measure distance uncertainty |  |  |  |
| 3 |  | Quality of Quiet Zone |  |  |  |
| 4 |  | Mismatch |  |  |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  |  |  |
| 6 |  | gNB emulator uncertainty |  |  |  |
| 7 |  | Phase curvature |  |  |  |
| 8 |  | Amplifier uncertainties |  |  |  |
| 9 |  | Random uncertainty |  |  |  |
| 10 |  | Influence of the XPD |  |  |  |
| 11 |  | Insertion Loss Variation |  |  |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |
| 13 |  | Multiple measurement antenna uncertainty |  |  |  |
| 14 |  | DUT repositioning |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 |  | Mismatch |  |  |  |
| 16 |  | Amplifier Uncertainties |  |  |  |
| 17 |  | Misalignment of positioning System |  |  |  |
| 18 |  | Uncertainty of the Network Analyzer |  |  |  |
| 19 |  | Uncertainty of the absolute gain of the calibration antenna |  |  |  |
| 20 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |
| 21 |  | Phase centre offset of calibration antenna |  |  |  |
| 22 |  | Quality of quiet zone for calibration process |  |  |  |
| 23 |  | Standing wave between reference calibration antenna and measurement antenna |  |  |  |
| 24 |  | Influence of the calibration antenna feed cable |  |  |  |
| 25 |  | Insertion Loss Variation |  |  |  |
| Systematic uncertainties |  |  |  |  |  |
| 26 |  | Systematic error related to beam peak search |  |  |  |

Table 2: MU elements independent of far-field test methodology

Definition of MU elements highlighted in green in Table 2 are not linked with specific OTA test methodology. Hence, these MU elements can be considered common across all permitted far-field test methodologies, i.e., IFF and DFF.

For MU elements belonging to Group 1, the proposal is to re-use values from IFF test setup defined in TR 38.903 Table E.3.1.3-2: Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power.

Proposal 1: For MU elements belonging to Group 1, the proposal is to re-use values from IFF test setup defined in TR 38.903 Table E.3.1.3-2 for DFF test setup.

- Group 2: MU elements dependent of test methodology, highlighted in yellow in table 3

| UID |  | Description of uncertainty contribution |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 |  | Positioning misalignment |  |  |  |
| 2 |  | Measure distance uncertainty |  |  |  |
| 3 |  | Quality of Quiet Zone |  |  |  |
| 4 |  | Mismatch |  |  |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  |  |  |
| 6 |  | gNB emulator uncertainty |  |  |  |
| 7 |  | Phase curvature |  |  |  |
| 8 |  | Amplifier uncertainties |  |  |  |
| 9 |  | Random uncertainty |  |  |  |
| 10 |  | Influence of the XPD |  |  |  |
| 11 |  | Insertion Loss Variation |  |  |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |
| 13 |  | Multiple measurement antenna uncertainty |  |  |  |
| 14 |  | DUT repositioning |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 |  | Mismatch |  |  |  |
| 16 |  | Amplifier Uncertainties |  |  |  |
| 17 |  | Misalignment of positioning System |  |  |  |
| 18 |  | Uncertainty of the Network Analyzer |  |  |  |
| 19 |  | Uncertainty of the absolute gain of the calibration antenna |  |  |  |
| 20 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |
| 21 |  | Phase centre offset of calibration antenna |  |  |  |
| 22 |  | Quality of quiet zone for calibration process |  |  |  |
| 23 |  | Standing wave between reference calibration antenna and measurement antenna |  |  |  |
| 24 |  | Influence of the calibration antenna feed cable |  |  |  |
| 25 |  | Insertion Loss Variation |  |  |  |
| Systematic uncertainties |  |  |  |  |  |
| 26 |  | Systematic error related to beam peak search |  |  |  |

Table 3: MU elements dependent of far-field test methodology

Definition of MU elements highlighted in yellow in Table 3 have dependencies with OTA test methodology and these ones needs to be evaluated per each OTA test methodology.

For MU elements belonging to Group 2, specific MU value needs to be analysed for DFF test setup.

Measurement distance uncertainty

TBD

Quality of Quite Zone

In [2], QoQZ for DFF is proposed. Hence, this value will be used for DL AWGN absolute power or wanted DL signal absolute power for DFF.

Phase curvature

TBD

Influence of the XPD

In [2], Influence of the XPD is proposed. Hence, this value will be used for DL AWGN absolute power or wanted DL signal absolute power for DFF.

Hence, Table 4 is the proposal for Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) to be added in TR 38.903 for DFF test setup

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | TBD | Rectangular | 1.73 | TBD |
| 3 | Quality of Quiet Zone (NOTE 4) | TBD | Actual | 1.00 | TBD |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | TBD | U-shaped | 1.41 | TBD |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.03 | U-shaped | 1.41 | 0.02 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 16 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 17 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 19 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 20 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 21 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 22 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 23 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 24 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 25 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 26 | Systematic error related to beam peak search |  |  |  | 0.5 |
| Total measurement uncertainty |  |  |  |  | Value |
| DL AWGN absolute power expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | TBD |
| NOTE 1: The analysis was done only for the case of operating in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in Annex D.2 of TR 38.810 [13] for Quiet Zone size less or equal to 30 cm. NOTE 5: The values in this table have been derived for DL powers above and equal to REFSENS. The values might need to be revisited for power levels below REFSENS |  |  |  |  |  |

Table 4: Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for DFF

Proposal 2: To add Table 4 in TR 38.903 Annex E.3.1.1 as Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for DFF test setup

# Conclusion

The following observations and proposals were made in this contribution:

Proposal 1: For MU elements belonging to Group 1, the proposal is to re-use values from IFF test setup defined in TR 38.903 Table E.3.1.3-2 for DFF test setup.

Proposal 2: To add Table 4 in TR 38.903 Annex E.3.1.1 as Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for DFF test setup

# References

- TR 38.903, Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases, V16.5.0 (2020-09)

- R5-206075, On DFF QoQZ MU, Keysight Technologies, 3GPP TSG-RAN5 Meeting #89-e, November 2020