3GPP TSG RAN WG5 Meeting #96-e 	R5-225646

Electronic Meeting, 15th – 26th August 2022

Source:		Rohde & Schwarz

Title:	Discussion on QoQZ and XPD MUs for 40cm

Agenda Item:		5.4.17

Document for:	Discussion and Endorsement

# Introduction

Following the discussion during last meeting regarding 40cm QZ size and the impact on MU, we present here our results and proposals.

# QoQZ measurement results

The results presented here are based on an EIRP and TRP QoQZ measurement campaign with 238 measurements points (assuming re-positioning approach) from the combination of reference points (P1 to P7), reference antenna orientations (and) and reference antenna polarization (pol). No ETC enclosure was used for this measurement campaign. The reference points are as defined in clause O.2.5 of TS 38.521-2 [1] with the following coordinates:

Table 2-1: AUT reference points coordinates for 40cm QZ

| Position | x | y | z |
| --- | --- | --- | --- |
| P1 | 0 | 0 | 0 |
| P2 | 20 cm | 0 | 0 |
| P3 | - 20 cm | 0 | 0 |
| P4 | 0 | 20 cm | 0 |
| P5 | 0 | - 20 cm | 0 |
| P6 | 0 | 0 | 20 cm |
| P7 | 0 | 0 | - 15 cm |

The following table show the EIRP and TRP standard deviation calculated from full campaign measurement results for 40cm QZ (i.e. 238 points, P1 to P7, all antenna orientations, 2 polarizations):

Table 2-2: EIRP QoQZ MU for 40cm QZ

|  |  |  |  |  | EIRP Std Deviation [dB] / Frequency |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Set of points | # meas. | γpol [º] | α [º] | β [º] | 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| P1 to P7 | 238 | all | {0} | {0,45,90,270,315} | 0.55 | 0.71 | 0.75 | 0.98 |
| P1 | 34 | all | {0} | {0,45,90,270,315} | 0.14 | 0.24 | 0.25 | 0.20 |

Table 2-3: TRP QoQZ MU for 40cm QZ

|  |  |  |  |  | TRP Std Deviation [dB] / Frequency |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Set of points | # meas. | γpol [º] | α [º] | β [º] | 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| P1 to P7 | 238 | all | {0} | {0,45,90,270,315} | 0.44 | 0.63 | 0.68 | 0.70 |
| P1 | 34 | all | {0} | {0,45,90,270,315} | 0.15 | 0.22 | 0.23 | 0.21 |

The measurement results for 30cm QZ for the same chamber without the 40cm QZ upgrade are also provided in the following tables as a reference. The same conditions described above (i.e. 238 measurement points assuming re-positioning and no ETC enclosure) were used while the reference positions are adapted accordingly to 30cm QZ:

Table 2-4: AUT reference points coordinates for 30cm QZ

| Position | x | y | z |
| --- | --- | --- | --- |
| P1 | 0 | 0 | 0 |
| P2 | 15 cm | 0 | 0 |
| P3 | - 15 cm | 0 | 0 |
| P4 | 0 | 15 cm | 0 |
| P5 | 0 | - 15 cm | 0 |
| P6 | 0 | 0 | 15 cm |
| P7 | 0 | 0 | - 15 cm |

Table 2-5: EIRP QoQZ MU for 30cm QZ

|  |  |  |  |  | EIRP Std Deviation [dB] / Frequency |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Set of points | # meas. | γpol [º] | α [º] | β [º] | 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| P1 to P7 | 238 | all | all | {0,45,90,270,315} | 0.49 | 0.57 | 0.57 | 0.63 |
| P1 | 34 | all | all | {0,45,90,270,315} | 0.14 | 0.13 | 0.11 | 0.23 |

Table 2-6: TRP QoQZ MU for 30cm QZ

|  |  |  |  |  | TRP Std Deviation [dB] / Frequency |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Set of points | # meas. | γpol [º] | α [º] | β [º] | 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| P1 to P7 | 238 | all | all | {0,45,90,270,315} | 0.42 | 0.46 | 0.48 | 0.51 |
| P1 | 34 | all | all | {0,45,90,270,315} | 0.10 | 0.12 | 0.05 | 0.10 |

As it can be seen from the comparison between the results for a 30cm QZ chamber and the upgrade to 40cm, there is an expected 0.3dB increase in the EIRP QoQZ MU, which confirms the values presented in [7].

Observation 1: the QZ extension of current test systems from 30cm to 40cm require an increase of the EIRP Quiet Zone uncertainty contribution by 0.3dB.

Table 2-7 presents a summarized view of the QoQZ MU target value per frequency used in TR 38.903 [2], with the additional value for 44.3GHz corresponding to n259 as endorsed in [6]:

Table 2-7: EIRP and TRP QoQZ MU target value for 30cm QZ

|  | Std Deviation [dB] / Frequency |  |  |  |
| --- | --- | --- | --- | --- |
| Stage | 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| Stage 2 (P1 to P7) | 0.60 | 0.60 | 0.60 | 0.70 |
| Stage 1 (P1) | 0.40 | 0.40 | 0.40 | 0.50 |

Following the results presented above, taking the worst case found for 40.8GHz as a reference, the values in Table 2-8 are proposed as QoQZ MU for EIRP and TRP for 40cm QZ:

Table 2-8: Proposed EIRP and TRP QoQZ MU target value for 40cm QZ

|  | Std Deviation [dB] / Frequency |  |  |  |
| --- | --- | --- | --- | --- |
| Stage | 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| Stage 2 (P1 to P7) | 0.75 | 0.75 | 0.75 | TBD |
| Stage 1 (P1) | 0.40 | 0.40 | 0.40 | 0.50 |

Proposal 1: Relax the Stage 2 QoQZ MU for 40cm QZ by 0.15dB as shown in Table 2-8, i.e. 0.75dB for frequencies ≤ 40.8GHz for NTC. Other frequencies and ETC are FFS.

# XPD measurement results

In previous analyses [5][7], but also confirmed with measurements in [3][4], a degradation of the cross polar discrimination (XPD) is expected on systems upgraded to 40cm QZ.

Table 3-1 presents a summarized view of the Influence of XPD MU target value per frequency used in TR 38.903 [2],

Table 3-1: Influence of XPD MU target value for 30cm QZ

| Influence of XPD MU [dB] / Frequency |  |  |  |
| --- | --- | --- | --- |
| 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| 0.01 | 0.01 | 0.01 | Not defined |

Using the results from the QoQZ measurement campaign for described in previous section, a similar degradation is observed for 40cm, and therefore the following values are proposed:

Table 3-2: Proposed Influence of XPD MU target value for 40cm QZ

| Influence of XPD MU [dB] / Frequency |  |  |  |
| --- | --- | --- | --- |
| 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| 0.03 | 0.03 | 0.03 | 0.03 |

Proposal 2: Relax the Influence of XPD MU for 40cm QZ by 0.02dB as shown in Table 3-2, i.e. 0.03dB for frequencies ≤ 44.3GHz. Other frequencies are FFS.

# Impact on overall MTSU

As presented before [5], the impact on the overall MU for 40cm QZ, taking MOP and REFSENS for PC3 ≤40.8GHz as example, is evaluated here for information using the proposed values above.

Table 4-1: Summary example of impact on EIRP MTSU for 40cm QZ

|  | Expanded Uncertainty (1.96σ - confidence interval of 95 %) [dB] / Frequency |  |
| --- | --- | --- |
|  | 23.45GHz <= f <= 32.125GHz | 32.125GHz < f <= 40.8GHz |
| 30cm QZ, PC3 UEs, NTC | 4.89 | 5.09 |
| 40cm QZ, PC3 UEs, NTC | 4.98 | 5.18 |
| Increase 30cm to 40cm QZ | 0.09 | 0.09 |

Table 4-2: Summary example of impact on TRP MTSU for 40cm QZ

|  | Expanded Uncertainty (1.96σ - confidence interval of 95 %) [dB] / Frequency |  |
| --- | --- | --- |
|  | 23.45GHz <= f <= 32.125GHz | 32.125GHz < f <= 40.8GHz |
| 30cm QZ, PC3 UEs, NTC | 4.42 | 4.62 |
| 40cm QZ, PC3 UEs, NTC | 4.51 | 4.71 |
| Increase 30cm to 40cm QZ | 0.09 | 0.09 |

Table 4-3: Summary example of impact on EIS MTSU for 40cm QZ

|  | Expanded Uncertainty (1.96σ - confidence interval of 95 %) [dB] / Frequency |
| --- | --- |
|  | 23.45GHz <= f <= 40.8GHz |
| 30cm QZ, PC3 UEs, NTC | 5.19 |
| 40cm QZ, PC3 UEs, NTC | 5.27 |
| Increase 30cm to 40cm QZ | 0.08 |

# Conclusion

In this contribution, the following observations and proposals were made:

Observation 1: the QZ extension of current test systems from 30cm to 40cm require an increase of the EIRP Quiet Zone uncertainty contribution by 0.3dB.

Proposal 1: Relax the Stage 2 QoQZ MU for 40cm QZ by 0.15dB as shown in Table 2-8, i.e. 0.75dB for frequencies ≤ 40.8GHz for NTC. Other frequencies and ETC are FFS.

Proposal 2: Relax the Influence of XPD MU for 40cm QZ by 0.02dB as shown in Table 3-2, i.e. 0.03dB for frequencies ≤ 44.3GHz. Other frequencies are FFS.

# References

- 3GPP TS 38.521-2, v16.12.0 (2022-06)

- 3GPP TR 38.903, v16.12.0 (2022-06)

- R5-223613, “MU discussion on 40 cm Quiet Zone”, Anritsu, 3GPP TSG RAN WG5 Meeting #95-e, May 2022

- R5-223614, “On QoQZ for 40cm QZ”, Keysight Technologies, 3GPP TSG RAN WG5 Meeting #95-e, May 2022

- R5-213193, “Discussion on the size of Quiet Zone above 30cm”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #90-e, May 2021

- R5-211656, “On the MU for n259”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #90-e, February 2021

- R5-211235, “Discussion on the size of Quiet Zone above 30cm”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #90-e, February 2021

- R5-206130, “Discussion on size of Quiet Zone above 30cm”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #89-e, November 2020

- R5-204200, “Discussion on size of Quiet Zone above 30cm”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #88-e, August 2020

- R5-202082, “Discussion on size of Quiet Zone above 30cm”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #87-e, May 2020