3GPP TSG-RAN5 Meeting #98	R5-230221

# Athens, Greece, 27th February – 3rd March 2023

| CR-Form-v12.2 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.903 | CR | 0467 | rev | - | Current version: | 17.0.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | Update of the uncertainty of the network analyzer |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | ROHDE & SCHWARZ |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | TEI15_Test, 5GS_NR_LTE-UEConTest |  |  |  |  |  | Date: |  |  | 2023-02-10 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-17 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-16 (Release 16) Rel-17 (Release 17) Rel-18 (Release 18) Rel-19 (Release 19) |  |
| Reason for change: |  | The uncertainty of the network analyzer of PC3 and that of PC1 do not match for the spurious emissions test cases although the uncertainty is independent of the power class. |  |  |  |  |  |  |  |  |
| Summary of change: |  | The uncertainty of the network analyzer of PC3 has been aligned to that of PC1 for the spurious emissions test cases. The offset value for coarse TRP measurement grids has been updated accordingly. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | The uncertainty assessment will remain inconsistent. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | B.2.2.14, B.18, B.18.2, B.25, B.25.2 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  | x |  | Other core specifications |  |  |  | TS 38.521-2 CR 0880 |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | Associated discussion paper in R5-230220. Associated CR R5-230222 updates the MTSU in TS 38.521-2. |  |  |  |  |  |  |  |  |
| This CR's revision history: |  |  |  |  |  |  |  |  |  |  |

< Unchanged sections omitted >

### B.2.2.14	Uncertainty of the Network Analyzer

See B.2.1.14.

The uncertainty value of uncertainty of the network analyzer is estimated as below table and used across clause B.

Table B.2.2.14-1: Uncertainty value for uncertainty of the network analyser for IFF

| Power class | Test case | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| PC1 | Default (6GHz to 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
|  | SE (40.8GHz to 66GHz) | 1.70 | Normal | 2.00 | 0.85 |
|  | SE (66GHz to 80GHz) | 1.70 | Normal | 2.00 | 0.85 |
| PC3 | Default | 0.73 | Normal | 2.00 | 0.37 |
|  | Minimum output power, OFF power (EIRP, TRP), ACLR | 1.50 | Normal | 2.00 | 0.75 |
|  | SE (6GHz to 12.75GHz) | 1.50 | Normal | 2.00 | 0.75 |
|  | SE (12.75GHz to 23.45GHz) | 1.50 | Normal | 2.00 | 0.75 |
|  | SE (23.45GHz to 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
|  | SE (40.8GHz to 66GHz) | 1.70 | Normal | 2.00 | 0.85 |
|  | SE (66GHz to 80GHz) | 1.70 | Normal | 2.00 | 0.85 |

< Unchanged sections omitted >

# B.18	Spurious emissions

Editor’s Note:

-	MU value analysis and offset value analysis for PC1, 2 and 4 are not complete.

-	MU value analysis for various test setups in clause B.18.x is not complete for above 80 GHz.

-	Offset value analysis is not complete as it is derived from MU value analysis for above 80 GHz.

Test procedure of general spurious emission comprises 2 stages: coarse TRP measurement and fine TRP measurement BW. Coarse TRP measurement is introduced to reduce the measurement time by applying sparser grids and/or wider measurement BW than fine TRP measurement while having offset dB more stringent test requirement in order not to cause additional misjudgement risk. For the frequency ranges for which coarse TRP measurement does not PASS, the measurement is continued with fine TRP measurement procedure.

Tables B.18-1,  B.18-1a, B.18-1b summarizes the MU threshold for fine TRP measurements for General spurious emissions, spurious emission band UE co-existence and additional spurious emission, respectively. The origin MU values for fine TRP measurement for different test setups can be found in following subclauses.

Table B.18-1: MU threshold for TRP measurement for general spurious emission

| Power Class | Frequency | In-band BW | In-band Power (NOTE2) | Threshold MU value [dB] (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | 6 GHz <= f <=12.75 GHz | BW <= 400MHz | P = Max Output Power | 5.29 |
|  | 12.75 GHz <= f <= 23.45 GHz |  |  | 5.25 |
|  | 23.45 GHz <= f <= 40.8 GHz |  |  | 5.41 |
|  | 40.8 GHz <= f <= 66 GHz |  |  | 7.42 |
|  | 66 GHz <= f <= 80 GHz |  |  | 7.72 |
| PC1 | 6 GHz <= f <=12.75 GHz | BW <= 400MHz | P = Max Output Power | [5.28] |
|  | 12.75 GHz <= f <= 23.45 GHz |  |  | [5.91] |
|  | 23.45 GHz <= f <= 40.8 GHz |  |  | [6.07] |
|  | 40.8 GHz <= f <= 66 GHz |  |  | [8.09] |
|  | 66 GHz <= f <= 80 GHz |  |  | FFS |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.18.2-3 to Table B.18.2-11 for PC3 UEs and in Table B.18.2-12 to Table B.18.2-16 for PC1 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |

Table B.18-1a: MU threshold for TRP measurement for spurious emission band UE co-existence

| Power Class | Frequency | In-band BW | In-band Power (NOTE2) | Threshold MU value [dB] (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | n257, n260, n261 | BW <= 400MHz | P = Max Output Power | 6.00 |
|  | 23.6 GHz < f <= 24.0 GHz |  |  | 6.00 |
|  | 36 GHz <= f <= 37 GHz |  |  | 6.00 |
|  | 57 GHz <= f <= 66 GHz |  |  | 8.01 |
| PC1 | n257, n261 | BW <= 400MHz | P = Max Output Power | FFS |
|  | n260 |  |  | FFS |
|  | 23.6 GHz < f <= 24.0 GHz |  |  | [7.32] |
|  | 36 GHz <= f <= 37 GHz |  |  | FFS |
|  | 57 GHz <= f <= 66 GHz |  |  | FFS |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.18.2-3 to Table B.18.2-11 for PC3 UEs and in Table B.18.2-12 to Table B.18.2-16 for PC1 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |

Table B.18-1b: MU threshold for TRP measurement for additional spurious emission

| Power Class | Frequency | In-band BW | In-band Power (NOTE2) | Threshold MU value [dB] (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | 6 GHz <= f <=12.75 GHz NS_202 | BW <= 400MHz | P = Max Output Power | 5.29 |
|  | 12.75 GHz <= f <= 23.45 GHz NS_202 |  |  | 5.84 |
|  | 23.45 GHz <= f <= 40.8 GHz NS_202, NS_203 |  |  | 6.00 |
|  | 40.8 GHz <= f <= 2nd harmonic of the upper frequency edge of the UL operating band NS_202 |  |  | 8.01 |
| PC1 | 6 GHz <= f <=12.75 GHz NS_202 | BW <= 400MHz | P = Max Output Power | FFS |
|  | 12.75 GHz <= f <= 23.45 GHz NS_202 |  |  | FFS |
|  | 23.45 GHz <= f <= 40.8 GHz NS_202, NS_203 |  |  | FFS |
|  | 40.8 GHz <= f <= 2nd harmonic of the upper frequency edge of the UL operating band NS_202 |  |  | FFS |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.18.2-3 to Table B.18.2-11 for PC3 UEs and in Table B.18.2-12 to Table B.18-2.16 for PC1 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |

Table B.18-2 provides valid coarse TRP measurement grids and corresponding offset dB value that may be used for UE general spurious emission test case. The offset value is derived as 95%-tile TRP measurement uncertainty including the effect from uncertainty due to Coarse TRP measurement grid, excluding influence of noise.

Table B.18-2: Coarse TRP measurement grids and offset values for UE Tx spurious emission

| Power Class | Coarse TRP measurement grid | Frequency | Min Number of measurement points on the grid | Influence of coarse TRP measurement grid (dB) | Systematic error due to coarse TRP calculation/quadrature (dB) | Offset value (dB) |
| --- | --- | --- | --- | --- | --- | --- |
| PC3 | Constant density grid (charged particle based) | 6 GHz <= f <=12.75 GHz | 35 | 0.94 | 0.09 | 5.27 |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | 5.23 |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | 5.38 |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | 7.31 |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | 7.61 |
|  | Constant step size grid | 6 GHz <= f <=12.75 GHz | 62 | 0.97 | 0.2 | 5.40 |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | 5.36 |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | 5.52 |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | 7.43 |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | 7.73 |
| PC1 | Constant density grid (charged particle based) | 6 GHz <= f <=12.75 GHz | FFS | FFS | FFS | FFS |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | FFS |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | FFS |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | FFS |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | FFS |
|  | Constant step size grid | 6 GHz <= f <=12.75 GHz | FFS | FFS | FFS | FFS |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | FFS |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | FFS |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | FFS |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | FFS |
| NOTE 1: Based on Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.18.2-3 to Table B.18.2-11 for PC3 UEs and in Table B.18.2-12 to Table B.18.2-16 for PC1 UEs, replacing “Influence of TRP measurement grid” and “Systematic error due to TRP calculation/quadrature” by the values for coarse TRP grid, and excluding “Influence of noise”. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |  |  |

< Unchanged sections omitted >

## B.18.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.18.2-1.

Table B.18.2-1: Uncertainty contributions for TRP measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | # B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of TRP measurement grid | B.2.2.22 |
| 14 | Influence of beam peak search grid | B.2.2.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 16 | DUT repositioning | B.2.2.26 |
| 17 | Misalignment of DUT due to change of DUT orientation | B.2.2.31 |
| Stage 1: Calibration measurement |  |  |
| 18 | Mismatch | B.2.2.4 |
| 19 | Amplifier Uncertainties | B.2.2.8 |
| 20 | Misalignment of positioning System | B.2.2.13 |
| 21 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 24 | Phase centre offset of calibration antenna | B.2.2.18 |
| 25 | Quality of quiet zone for calibration process | B.2.2.19 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 27 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 28 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 29 | Systematic error due to TRP calculation/quadrature | B.2.2.24 |
| 30 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet zone size ≤ 30 cm, f = {6 GHz to 80 GHz}, P = Maximum output power.

-	The uncertainty assessment for TRP is provided from Table B.18.2-2 to Table B.18.2-11 for PC3 UEs and from Table B.18.2-12 to Table B.18.2-16 for PC1 UEs.

Table B.18.2-2: Void

Table B.18.2-3: Uncertainty assessment for TRP measurement (f=6 GHz to 12.75GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.00 | Normal | 2.00 | 1.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (6 GHz < f <= 12.75 GHz) [dB] (a) |  |  |  | 4.88 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | General spurious emissions Influence of noise (c1) (6 GHz < f <= 12.75 GHz) |  |  |  | 0.41 |
| 31 | Additional spurious emissions Influence of noise (c2) NS_202 (6 GHz < f <= 12.75 GHz) |  |  |  | 0.41 |
| 32 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (6 GHz < f <= 12.75 GHz) |  |  |  |  | 5.29 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c2) [dB] NS_202 (6 GHz < f <= 12.75 GHz) |  |  |  |  | 5.29 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.18.2-4: Void

Table B.18.2-5: Uncertainty assessment for TRP measurement (f=12.75 GHz to 23.45 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (12.75 GHz < f <= 23.45 GHz) [dB] (a) |  |  |  | 4.84 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | General spurious emissions Influence of noise (c1) (12.75 GHz < f <= 23.45 GHz) |  |  |  | 0.41 |
| 31 | Additional spurious emissions Influence of noise (c2) NS_202 (12.75 GHz < f <= 23.45 GHz) |  |  |  | 1.0 |
| 32 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (12.75 GHz < f <= 23.45 GHz) |  |  |  |  | 5.25 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c2) [dB] NS_202 (12.75 GHz < f <= 23.45 GHz) |  |  |  |  | 5.84 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.18.2-6: Void

< Unchanged sections omitted >

# B.25	Receiver spurious emissions

Editor’s Note:

-	MU value analysis and offset value analysis for PC1, 2 and 4 are not complete.

-	MU value analysis for various test setups in subsection B.25.x is not complete for above 80 GHz for PC3

-	Offset value analysis is not complete as it is derived from MU value analysis for above 80 GHz for PC3

Test procedure of general spurious emission comprises 2 stages: coarse TRP measurement and fine TRP measurement BW. Coarse TRP measurement is introduced to reduce the measurement time by applying sparser grids and/or wider measurement BW than fine TRP measurement while having offset dB more stringent test requirement in order not to cause additional misjudgement risk. For the frequency ranges for which coarse TRP measurement does not PASS, the measurement is continued with fine TRP measurement procedure.

Table B.25-1 summarizes the MU threshold for fine TRP measurements for General spurious emissions. The origin MU values for fine TRP measurement for different test setups can be found in following subclauses.

Table B.25-1: MU threshold for TRP measurement for Rx spurious emission

| Power Class | Frequency | In-band BW | In-band Power (NOTE2) | Threshold MU value [dB] (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | 6 GHz <= f <=12.75 GHz | BW <= 400MHz | P = Max Output Power | 5.64 |
|  | 12.75 GHz <= f <= 23.45 GHz |  |  | 5.60 |
|  | 23.45 GHz <= f <= 40.8 GHz |  |  | 6.11 |
|  | 40.8 GHz <= f <= 66 GHz |  |  | 7.65 |
|  | 66 GHz <= f <= 80 GHz |  |  | 7.95 |
| PC1 | 6 GHz <= f <=12.75 GHz | BW <= 400MHz | P = Max Output Power | [5.63] |
|  | 12.75 GHz <= f <= 23.45 GHz |  |  | [5.59] |
|  | 23.45 GHz <= f <= 40.8 GHz |  |  | [6.10] |
|  | 40.8 GHz <= f <= 66 GHz |  |  | [7.64] |
|  | 66 GHz <= f <= 80 GHz |  |  | FFS |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.25.2-3 to Table B.25.2-11 for PC3 UEs and in Table B.25.2-12 to Table B.25.2-16 for PC1 UEs. |  |  |  |  |

Table B.25-2 provides valid coarse TRP measurement grids and corresponding offset dB value that may be used for UE general spurious emission test case.  The offset value is derived as 95%-tile TRP measurement uncertainty including the effect from uncertainty due to Coarse TRP measurement grid, excluding influence of noise.

Table B.25-2: Coarse TRP measurement grids and offset values for UE Rx spurious emission

| Power Class | Coarse TRP measurement grid | Frequency | Min Number of measurement points on the grid | Influence of coarse TRP measurement grid (dB) | Systematic error due to coarse TRP calculation/quadrature (dB) | Offset value (dB) |
| --- | --- | --- | --- | --- | --- | --- |
| PC3 | Constant density grid (charged particle based) | 6 GHz <= f <=12.75 GHz | 35 | 0.94 | 0.09 | 5.38 |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | 5.35 |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | 5.49 |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | 7.31 |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | 7.61 |
|  | Constant step size grid | 6 GHz <= f <=12.75 GHz | 62 | 0.97 | 0.2 | 5.51 |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | 5.48 |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | 5.62 |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | 7.43 |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | 7.73 |
| PC1 | Constant density grid (charged particle based) | 6 GHz <= f <=12.75 GHz | FFS | FFS | FFS | FFS |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | FFS |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | FFS |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | FFS |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | FFS |
|  | Constant step size grid | 6 GHz <= f <=12.75 GHz | FFS | FFS | FFS | FFS |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | FFS |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | FFS |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | FFS |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | FFS |
| NOTE 1: Based on Total TRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.25.2-3 to Table B.25.2-11, replacing “Influence of TRP measurement grid” and “Systematic error due to TRP calculation/quadrature” by the values for coarse TRP grid, and excluding “Influence of noise”. |  |  |  |  |  |  |

< Unchanged sections omitted >

## B.25.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.25.2-1.

Table B.25.2-1: Uncertainty contributions for TRP measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of TRP measurement grid | B.2.2.22 |
| 14 | Influence of beam peak search grid | B.2.2.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 16 | DUT repositioning | B.2.2.26 |
| 17 | Misalignment of DUT due to change of DUT orientation | B.2.2.31 |
| Stage 1: Calibration measurement |  |  |
| 18 | Mismatch | B.2.2.4 |
| 19 | Amplifier Uncertainties | B.2.2.8 |
| 20 | Misalignment of positioning System | B.2.2.13 |
| 21 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 24 | Phase centre offset of calibration antenna | B.2.2.18 |
| 25 | Quality of quiet zone for calibration process | B.2.2.19 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 27 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 28 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 29 | Systematic error due to TRP calculation/quadrature | B.2.2.24 |
| 30 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {6 GHz to 80 GHz}, P =Receiver Spurious Core Requirement Level + Relaxation(For n257, 10.2dB for 6GHz  f < 20GHz, 17.2 dB for 20GHz f < 40GHz, 33.1dB for 40GHz  f  2nd harmonic)

-	The uncertainty assessment for TRP is provided from Table B.25.2-2 to Table B.25.2-11 for PC3 UEs and from Table B.25.2-12 to Table B.25.2-16 for PC1 UEs.

Table B.25.2-2: Void

Table B.25.2-3: Uncertainty assessment for TRP measurement (f=6 GHz to 12.75 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 4 | Mismatch | 1.60 | Actual | 1.00 | 1.60 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.00 | Normal | 2.00 | 1.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (6 GHz < f <= 12.75 GHz) [dB] (a) |  |  |  | 5.00 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | Influence of noise (6 GHz < f <= 12.75 GHz) (c) |  |  |  | 0.64 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | 5.64 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.25.2-4: void

Table B.25.2-5: Uncertainty assessment for TRP measurement (f=12.75 GHz to 23.45 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 4 | Mismatch | 1.60 | Actual | 1.00 | 1.60 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (12.75 GHz < f <= 23.45 GHz) [dB] (a) |  |  |  | 4.96 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | Influence of noise (12.75 GHz < f <= 23.45 GHz) (c) |  |  |  | 0.64 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | 5.60 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

< End of changes >