3GPP TSG RAN5 Meeting #98	 	R5-231304

Athens, Greece, 27th February – 3rd March 2023

Source:		Rohde & Schwarz

Title:	On the measurements of phase continuity requirements for FR1

Agenda Item:		5.3.18.8

Document for:	Discussion and Endorsement

# Introduction

At RAN5 meeting #97-e, RAN5 has received an LS from RAN4 on the measurement of phase continuity requirements for DMRS bundling [1]. At the last meeting this has been discussed for FR2 [2]. In this contribution, we continue the discussion for FR1 and propose a response to the questions raised by RAN4.

# Discussion

At RAN5 meeting #97-e, RAN4 liaised with RAN5 asking for information on the measurement accuracy associated with phase continuity requirements. RAN5 created the action point AP#96e.22 to gather feedback on the LS:

| Action ID | sWG | Action | Responsible | Relevant Tdoc | Deadline | Status |
| --- | --- | --- | --- | --- | --- | --- |
| AP#96e.22 | RF | Perform Phase continuity measurement analysis for DMRS bundling with Global In-channel Tx test to respond to RAN4 on the requested action for Freq error measurement feasibility | KEYS, Anritsu, R&S | R5-223917 (LS from RAN4) R5-227959 R5-227773 | RAN5#98 | Open |

## Phase continuity requirements

RAN4 has added new requirements for phase continuity in TS 38.101-1 [3] and TS 38.101-2 [4].

Since the action point is open for FR1, we focus on this frequency range in the following:

#### 6.4.2.5	Phase continuity requirements for DMRS bundling

For bands that UE indicates the support of DMRS bundling, when the UE is configured with DMRS bundling, the maximum allowable difference between the measured phase value in any slot p-1 and slot p, or slot 0 and any slot p for each antenna connector shall satisfy the requirements as listed in Table 6.4.2.5-1 for the measurement conditions defined in Table 6.4.2.5-2, within a measurement time window limited by the UE capability of maximum duration for DMRS bundling [maxDurationDMRS-Bundling-r17], and defined for each frequency band separately. The phase value for each slot is measured as shown in Annex F.9. These requirements apply to PUCCH and PUSCH transmissions with DFT-s-OFDM and CP-OFDM waveforms.

Table 6.4.2.5-1: Maximum allowable phase difference for DMRS bundling

| UL channel | Modulation order | Phase difference between any slot p-1 and slot p (NOTE 2) | Phase difference between slot 0 and any slot p (NOTE 3) |
| --- | --- | --- | --- |
| PUSCH | Pi/2 BPSK, QPSK | [25] degrees | [30] degrees |
| PUCCH | Pi/2 BPSK, BPSK, QPSK |  |  |
| NOTE 1: The UE capability of the length of maximum duration refers to the maximum time duration during which UE is able to meet the phase continuity requirements, assuming no phase consistency violating events defined in TS 38.214 in between. NOTE 2: This requirement applies for FDD and TDD bands, for supported DMRS bundling configurations ≤ 8 slots. NOTE 3: This requirement applies only for FDD bands, for supported DMRS bundling configurations of 16 slots. |  |  |  |

The above requirements are applicable when all the following conditions are met within the measurement time window:

-	RB allocation in terms of length and frequency position does not change, and intra-slot and inter-slot frequency hopping is not activated.

-	Modulation order does not change.

-	No network commanded TA takes effect.

-	The TPMI precoder does not change.

-	There is no change in UE transmission power level, and no change in the level of P-MPR applied by the UE.

-	UE is not scheduled with uplink transmission of other physical channel/signal in-between the PUSCH or PUCCH transmissions.

-	For TDD, no downlink slot(s) or downlink symbol(s) or flexible symbol(s) with/without DL monitoring occasion configured in-between the PUSCH or PUCCH transmissions.

Table 6.4.2.5-2: Measurement conditions for the maximum allowable phase difference

| Parameter | Unit | Level |
| --- | --- | --- |
| UE Output Power | dBm | PCMAX,f,c in clause 6.2.4, P-MPR = 0 |
| UE downlink received power |  | Not change |
| Operating conditions |  | Normal conditions |
| Transmission bandwidth |  | Confined within FUL_low + [4] MHz and FUL_high – [4] MHz |
| DL signal frequency |  | Not change before and during the measurement window |
| DL signal timing |  | Maintained constant before and during the measurement window |
| UL slots for testing |  | Tested on consecutive UL slots |
| PUSCH waveform for testing |  | DFT-s-OFDM |

In addition the active UL slot, TDD pattern and RMC have been specified in Annex A.2.1 and A.2.3 of TS 38.101-1 [3] :

## A.2.1	General

…

The active uplink slots for additional TDD configuration specified in Table A.2.1-2, with the TDD pattern defined in Table A.2.1-3, is used for the phase continuity tests for DMRS bundling with 15 kHz SCS.

Table A.2.1-2: TDD active uplink slots for additional TDD configuration

| SCS | Active Uplink slots |
| --- | --- |
| 15 kHz | 8, 9 |

Table A.2.1-3: TDD pattern for additional TDD configuration

| Parameter |  | Value |
| --- | --- | --- |
|  |  | SCS 15 kHz (µ0) |
| TDD Slot Configuration pattern (Note 1) |  | 7DS2U |
| Special Slot Configuration (Note 2) |  | 6D+4G+4U |
| referenceSubcarrierSpacing |  | 15 kHz |
| UL-DL configuration | dl-UL-TransmissionPeriodicity | 10 ms |
|  | nrofDownlinkSlots | 7 |
|  | nrofDownlinkSymbols | 6 |
|  | nrofUplinkSlot | 2 |
|  | nrofUplinkSymbols | 4 |
| NOTE 1: D denotes a slot with all DL symbols; S denotes a slot with a mix of DL, UL and guard symbols; U denotes a slot with all UL symbols. The field is for information. NOTE 2: D, G, U denote DL, guard and UL symbols, respectively. The field is for information. |  |  |

## A.2.3	Reference measurement channels for TDD

…

The TDD UL RMCs for the phase continuity tests for DMRS bundling are defined in clause A.2.2 with the active UL slots and TDD pattern specified in tables A.2.1-2 and A.2.1-3.

Annex F.9 in TS 38.101-1 [3] treats the phase offset measurement for DMRS bundling.

# F.9	Phase offset measurement for DMRS bundling

## F.9.1	Measurement point

The measurement point for phase offset measurement is defined in Figure F.9.1-1.

Figure F.9.1-1: Measurement point for phase offset for DMRS bundling

## F.9.2	Symbols used

Phase offset is determined based on DMRS REs (3 DMRS symbols per slot) with the option to use data symbols.

## F.9.3	Modified test signal

[editor notes: updates based on LS reply from RAN5]

## F.9.4	Phase offset measurement

The phase offset measurement is based on the phase response of the Tx chain  as derived based on Annex F.4.

The phase difference ∆φf for each subcarrier between a reference timeslot tref and the measurement timeslot tm is then calculated as defined below:

∆φf= φtm,f-φ(tref,f)

The phase offset between the reference and measurement timeslots are then calculated as the maximum over the results for all subcarriers as shown below:

PhaseOffset=maxf(∆φf)

## On Common Frequency Offset

First part of the LS considers the common frequency offset correction and its handling (within one bundling time). The UE requirement for frequency error of ±0.1 ppm will also apply for the DMRS bundling feature. The questions from RAN4 associated to the accuracy of frequency correction by TE are:

Q1-1: What would be the impact of the residual frequency error after CFO correction by TE on the testability of the phase offset measurement for UE complying with +/- 0.1 ppm frequency error requirement, particularly for the upper range of FR2-1 and lowest allowed SCS?

The upper frequency limit of FR1 is specified as 7.125 GHz. The test requirement for phase continuity in FR1 is measured in two adjacent slots with a requirement of [25] degrees for TDD and FDD bands, and in slot 0 and any slot p for FDD bands and DMRS bundling configuration of 16 slots. Currently, the highest UL frequency of FDD bands is 2750 MHz. The longest slot time is associated to an SCS of 15 kHz and its duration is 1 ms. An UE fulfilling the frequency error requirement of ±0.1 ppm can change its phase up to 256 deg for adjacent slots (assuming a carrier frequency of 7125 MHz) and up to 1584 deg between slot 0 and any slot p (assuming 2750 MHz as carrier frequency). However, this is a factor of approximately 10 or 53 larger than the RAN4 requirement for phase continuity between adjacent slots and between slot 0 and any slot p, respectively.

Observation 1: Due to the ±0.1 ppm requirement, the UE can change its phase by several 100 deg.

Observation 2: The phase continuity requirement is a stricter requirement than the UE frequency error requirement.

The highest FDD band (n7) has a maximum UL frequency of 2750 MHz which is below 3 GHz so that the MTSU for frequency error of ±15 Hz applies and the phase difference can be ±86 deg for the measurement between slot 0 and any slot p. This exceeds the minimum conformance requirement of [30] deg by a factor of 2.9 which blocks the measurement.

Considering the MTSU for frequency error of ±36 Hz for frequencies between 3 GHz and 5.925 GHz, the phase difference can be ±13 deg for the adjacent slot measurement. This value is 52% of the minimum conformance requirement.

We note that the MTSU for frequency error has not yet been discussed and specified for frequencies above 5.925 GHz. Therefore, we expect that the phase difference caused by the test equipment can be larger than ±13 deg for the TDD bands n102, n96 and n104.

Observation 3: In the case of FDD, the MTSU for frequency error results in an approximately 3 times larger phase offset than the targeted RAN4 requirement for UE phase continuity which makes a direct measurement not reasonable.

Observation 4: In the case of TDD, the MTSU for frequency error results in an MU which is 52% of targeted RAN4 requirement for UE phase continuity. This high level of inaccuracy can result in failing conformant UEs and the requirement needs to be altered by a significant test tolerance.

Since the accumulated phase error can be quite large, we propose to correct the mean phase for each slot where the mean is with respect to the different sub carriers (SCs).

Base stations are correcting the signal in the best way. Therefore, it is reasonable for the phase continuity measurement to correct the mean phase per slot.

Proposal 1: For FR1, correct the mean phase for each slot and apply the phase continuity requirement on the worst case phase difference between the adjacent slots with respect to the sub carriers.

The mean phase error correction can be defined as

φ't,f= φt,f-1ni=1nφt,fi,

where φt,f is the phase response of the TX chain and fi are the sub carriers.

After the mean phase error correction, the processing as described in Annex F.9.4 of TS 38.101-1 [3] can be performed, where we suggest to apply absolute value as highlighted in yellow below:

The phase difference ∆φf for each subcarrier between a reference timeslot tref and the measurement timeslot tm is then calculated as defined below:

∆φf= φ'tm,f-φ'(tref,f)

The phase offset between the reference and measurement timeslots are then calculated as the maximum over the results for all subcarriers as shown below:

PhaseOffset=maxf(|∆φf|)

The phase continuity requirement is applied on the residual phase difference for each SC between the two slots after mean phase error correction. Among all these differences the maximum absolute value is taken as the result and compared to the phase continuity requirement.

A further question raised by RAN4 is:

Q1-2: Can Frequency error be corrected in such way that it has minimal or no impact on measurement uncertainty for phase discontinuity requirements? If answer is yes, in which options below could achieve this:

- CFO correction in each slot should be applied taking into account a slot starting from a reference time slot (e.g. slot 0 or any slot p-1) to measured time slot (slot p) to enable a continuous correction to avoid any potential phase step introduced by CFO correction.

- CFO in each slot is corrected with its individual frequency error

In order to minimize the impact of frequency error we propose to correct the mean phase in each slot. Thus, option 2 in combination with mean phase correction results in no impact of frequency error on the phase continuity measurement.

Observation 5: Mean phase error correction per slot results in no impact of frequency error on the phase continuity measurement.

Q1-3: Does RAN5 see any other testing issue apart from issues raised by RAN4 above?

It is also required to correct the timing of the signal. Thus, Annex F of RAN4 TS 38.101-1 [3], needs to be updated to consider mean phase correction and timing alignment.

Observation 6: It is not enough to correct the frequency offset, also the timing of the signal needs to be corrected.

Proposal 2: For FR1, correct also the timing of the signal.

## On improving test accuracy by repeated testing

Q2-1: The requirement of phase continuity requirement is applied within one DMRS bundling configuration, would the measurement uncertainty be improved as the result of repeating the test over several bundles and if so, what would be the recommended number of bundles and should it be dependent on the subcarrier spacing?

Whether repeated testing can improve the accuracy depends on the details of definition of the phase measurement. If a mean phase error correction per slot is not applied, repeated testing will reduce the impact of random fluctuations on the MU. However, we note that repeated testing is not required when a mean phase error is corrected per slot.

Observation 7: If mean phase error correction per slot is not applied, the accuracy can be improved by repeated testing. If mean phase error is corrected per slot, repeated testing is not needed.

Q2-2: In case multiple bundles will be tested, what is Ran5 view on how the test results in terms of phase offset from individual bundles be used to improve the test reliability and reduce measurement uncertainty (e.g., maximum over all bundles or RMS average over all bundles or others method)?

In the case of no mean phase error correction per slot, direct averaging of the results for individual subcarriers of the repetitions can be applied. No improvement of accuracy is possible when the results of individual subcarriers are processed by maximum or RMS average due to the non-linearity of these functions:

PhaseOffset=maxf(|mean bundles(∆φf,bundle|)

However, due to the issues of common frequency offset correction, we prefer to apply a mean phase error correction per slot.

Observation 8: Improvement by repeated testing requires averaging with no maximum or RMS average allowed to be placed before the averaging. Due to the issues of common frequency offset correction, a mean phase error correction per slot is preferred and repeated testing not needed.

# Assumptions

| ID | Description | Assumption |
| --- | --- | --- |
| #1 | Frequency ranges under consideration | FR1 |

# Conclusion

In summary, we have analyzed the measurement of phase continuity requirements.

Observation 1: Due to the ±0.1 ppm requirement, the UE can change its phase by several 100 deg.

Observation 2: The phase continuity requirement is a stricter requirement than the UE frequency error requirement.

Observation 3: In the case of FDD, the MTSU for frequency error results in an approximately 3 times larger phase offset than the targeted RAN4 requirement for UE phase continuity which makes a direct measurement not reasonable.

Observation 4: In the case of TDD, the MTSU for frequency error results in an MU which is 52% of targeted RAN4 requirement for UE phase continuity. This high level of inaccuracy can result in failing conformant UEs and the requirement needs to be altered by a significant test tolerance.

Proposal 1: For FR1, correct the mean phase for each slot and apply the phase continuity requirement on the worst case phase difference between the adjacent slots with respect to the sub carriers.

Observation 5: Mean phase error correction per slot results in no impact of frequency error on the phase continuity measurement.

Observation 6: It is not enough to correct the frequency offset, also the timing of the signal needs to be corrected.

Proposal 2: For FR1, correct also the timing of the signal.

Observation 7: If mean phase error correction per slot is not applied, the accuracy can be improved by repeated testing. If mean phase error is corrected per slot, repeated testing is not needed.

Observation 8: Improvement by repeated testing requires averaging with no maximum or RMS average allowed to be placed before the averaging. Due to the issues of common frequency offset correction, a mean phase error correction per slot is preferred and repeated testing not needed.

# References

- R5-223917, LS to RAN5 on measurement of phase continuity requirements for DMRS bundling, 3GPP TSG RAN WG4, WG5 Meeting #96-e, August 2022

- R5-227959, On the measurements of phase continuity requirements, Rohde & Schwarz, 3GPP TSG RAN WG 5 Meeting #97, Toulouse, November 2022

- TS 38.101-1, V17.8.0, 3GPP TSG RAN WG4, December 2022

- TS 38.101-2, V17.8.0, 3GPP TSG RAN WG4, December 2022