3GPP TSG RAN WG5 Meeting #98 	R5-231800

Athens, Greece, February 27th – March 3rd 2023

Source:		Bluetest

Title:	RC MU Analysis for NR FR1 TRP-TRS Enhancement (Rel-18)

Agenda Item:		5.3.31.2

Document for:	Discussion and Endorsement

# Introduction

In the November meeting, RAN5 endorsed an MU workplan for NR FR1 TRP-TRS Enhancement (Rel-18) [1] which included an initial assessment for the Reverberation Chamber (RC) method for TR FR1 TRP and TIS. This document proposes a draft text and discussion to begin this analysis.

## Approach

This initial draft of the MU analysis is primarily drawn from existing RC method MU tables for TRP and TRS such as Annex A in TR 25.914 [2], specifically tables A.3a, A.3b, A.3c, A.7 and A.8.  Also, the existing RC MU from Section 11.2.5 in TR 37.941 [3] for base station conformance testing, specifically tables 11.2.5.3-1, 11.3.5.3-1. In addition, significant effort is made to align as many terms as possible with the AC reference method to facilitate method harmonization efforts per the RAN4 WF on this subject [4]. The AC MU from Section B.2 in TR 38.834, specifically tables B.3-1 and B.4-1 was used for this initial draft with additional consideration of TR 38.810 [5] section B.1.1.2.

## Common MU Definitions

We anticipate many, but not all, measurement uncertainty terms for the RC method will be shared with the AC reference method. As the document structure for TR 38.870 [6] and measurement definitions for the AC method are developed these can be referenced for a planned RC summary table.

## RC Specific MU Definitions

In this initial draft we are introducing descriptions only for items that the RC method takes a different approach.

## Ongoing Development of Terms

Quality of Spatial Uniformity – Uncertainty term will be dependent on the finalization of the procedure currently under development in RAN4 [4]

Effect of coherence bandwidth – whether this term is required or is built into the spatial uniformity requirement is for further study under agreed work in [4].

# References

- R5-225924, “MU workplan for NR FR1 TRP-TRS Enhancement (Rel-18)”, Rhode & Schwarz, vivo, 3GPP TSG RAN5 Meeting #97-e, November 2022

- 3GPP TR 25.914 V17 2022-03

- 3GPP TR 37.941 V17 2022-03

- R4-2220610, “WF for Rel-18 TRP TRS WI”, vivo, RAN#105 meeting, Nov 2022

- 3GPP TR 38.810 V16.6.0 2020-10

- 3GPP TR 38.870 V0.1.0 2022-12

# Text Proposal to TR 38.870

For TR 38.870 [6] Annex B we propose to add:

<<  Start of text proposal >>

# B.3	MU contribution descriptions for Alternative method

<Editor’s note: Detailed structure of the subclause is TBD.

MU assessment for 1Tx, 2Tx and CA (if MU is different compared with 1Tx case), including different scenarios, e.g. Free space, hand only, head and hand, forearm >

## B.3.1	Additional Power Loss in EUT Chassis

When the EUT is small and do not add noticeable loss to the chamber, the calibration procedure outlined in section 8.4, is performed without the EUT present in the chamber. The possible difference in average chamber transmission level between the EUT measurement and the reference measurement must in this case be considered in the uncertainty evaluation. The uncertainty value for this contribution can be tested empirically by choosing a unit within a set of samples which is considered to incur the highest amount of loss (normally the largest unit), and measure the average transmission loss in the chamber with and without the test unit present in the chamber. The difference between the two cases shall be used in the uncertainty calculation and the distribution should be assumed to be rectangular.

Alternatively, a fixed value of 0.2 dB with a rectangular distribution can be used in the uncertainty calculations.

## B.3.2   Random Uncertainty

The random uncertainty characterizes the undefined and miscellaneous effects which cannot be forecasted. One can estimate this type of uncertainty with a repeatability test by making a series of repeated measurement with a reference antenna without changing anything in the measurement set-up.

## B.3.3   Quality of Spatial Uniformity

The standard deviation over calibration antenna positions and rotations of the transfer function, chamber reference or chamber loss S212. This measures the lack of spatial uniformity over the test zone in the reverberation chamber which is also referred as lack of isotropy, statistical ripple, or uniformity of the transfer function.

Mean value estimation of transfer function: this contribution originates from the error of the estimated mean related to the use of a finite number N of samples in the measurement of the transfer function. The mean value estimation is calculated as σ/N where σ is the standard deviation of the series of measured values.

# B.4	MU Assessment for TRP

<Editor’s note: Detailed structure of the subclause is TBD.

MU assessment for both Reference method and alternative method, under different test cases>

### B.4.1	MU Assessment for TRP (reference method)

TBD

### B.4.2	MU Assessment for TRP (alternative method)

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гpower meter <0.05 Гfixed measurement antenna connection <0.16 | 0.07 | U-shaped | 1.41 | 1 | [0.05] |
| 2 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 3 | Influence of the fixed measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 4 | Measurement Receiver: uncertainty of tshe absolute level | Power Meter | 0.06 | Rectangular | 1.73 | 1 | [0.03] |
| 5 | Uniformity of transfer function | Statistics of chamber | 0.5 | Gaussian | 1 | 1 | [0.5] |
| 6 | Additional power loss in EUT chassis | The EUT not present in the chamber during calibration measurement | 0.1 | Rectangular | 1.73 | 1 | [0.06] |
| 7 | DUT Tx-power drift | Drift | 0.2 | Rectangular | 1.73 | 1 | [0.12] |
| 8 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | 0.32 | Rectangular | 1.73 | 1 | [0.18] |
| 9 | Repeatability | Using the same setup and stirring sequence | 0.4 | Rectangular | 1.73 | 1 | [0.23] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 10 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers whole NA setup | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 11 | Mismatch of receiver chain | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 12 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 13 | Mismatch in the connection of calibration antenna | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 14 | Influence of the feed cable of the calibration antenna | Gain calibration with a dipole | 0.3 | Rectangular | 1.73 | 1 | [0.17] |
| 15 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 16 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 17 | Uniformity of transfer function | Statistics of chamber | Gaussian | 1.00 | 1 | 1 | [0.5] |
| Combined standard uncertainty |  |  |  |  |  |  | 0.7441 |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | 1.4585 |

Table B.4.2-1: Example of uncertainty budget for TRP measurement, alternative test method

# B.5	MU Assessment for TRS

<Editor’s note: Detailed structure of the subclause is TBD.

MU assessment for both Reference method and alternative method, under different test cases >

### B.5.1	MU Assessment for TRS (reference method)

<Editor’s note: Detailed structure MU assessment for Reference >

### B.5.2	MU Assessment for TRS (alternative method)

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | ГBSS <0.13 Г antenna connection <0.03 | 0.02 | Gaussian | 1 | 1 | [0.02] |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 3 | Influence of the fixed measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 5 | Base station simulator: uncertainty of the absolute level |  | 1 | Rectangular | 1.73 | 1 | [0.58] |
| 6 | BER measurement: output level step resolution | Step 0.1dB | 0.05 | Rectangular | 1.73 | 1 | [0.03] |
| 7 | Statistical uncertainty of the BER measurement | BER target 10%±2% , 20000 tested bits , N=60 | 0.12 | Gaussian | 1 | 1 | [0.12] |
| 8 | TRS data rate normalization | 4 reference points measured | 0.12 | Gaussian | 1 | 1 | [0.12] |
| 9 | Uniformity of transfer function | Statistics of chamber | 0.5 | Gaussian | 1 | 1 | [0.5] |
| 10 | Additional power loss in EUT chassis | The EUT not present in the chamber during calibration measurement | 0.1 | Rectangular | 1.73 | 1 | [0.06] |
| 11 | DUT sensitivity drift | Drift measurement | 0.2 | Rectangular | 1.73 | 1 | [0.12] |
| 12 | Uncertainty related to the use of phantoms | Standard SAM with standard tissue simulant | 0 | Rectangular | 1.73 | 1 | 0 |
| 13 | Repeatability | Using the same setup and stirring sequence | 0.4 | Rectangular | 1.73 | 1 | [0.23] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 14 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers NA setup | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 15 | Mismatch of transmitter chain | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0 |
| 16 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 17 | Mismatch in the connection of calibration antenna | Taken in to account in NA setup uncertainty | 0 | Rectangular | 1.73 | 1 | 0 |
| 18 | Influence of the feed cable of the calibration antenna | Gain calibration with dipole | 0.3 | Rectangular | 1.73 | 1 | [0.17] |
| 19 | Influence of the fixed measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 21 | Uniformity of transfer function | Statistics of chamber | 0.5 | Gaussian | 1 | 1 | [0.5] |
| Combined standard uncertainty |  |  |  |  |  |  | 1.0220 |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | 2.0031 |

B.5.2-1: Example of uncertainty budget for TRS measurement, alternative test method

# Conclusion

Proposal 1: The text submission should be considered for the development of the MU sections in TR 37.870