3GPP TSG RAN WG5 Meeting #99 	R5-232586

Incheon, South Korea, 22 – 26 May 2023

Source:		Rohde & Schwarz

Title:	AC MU Analysis for NR FR1 TRP-TRS (Rel.17)

Agenda Item:		5.3.31.2

Document for:	Discussion and Endorsement

# Introduction

Following the closure of the work item led by RAN4 for NR FR1 TRP-TRS, a new RAN5 work item was approved to define the corresponding NR FR1 TRP-TRS test specification TS 38.561 [1]. The aim of this contribution is to progress on the MU assessment work for TS 38.561 [1].

# MU contributors in TR 38.834

As approved during last meeting, Annex A.4 of TS 38.561 [1] contains the MU calculation principles, the list of MU contributors, the description for each MU contribution and preliminary uncertainty budgets for hand only (browsing mode), copied from the Annex B of TR 38.884 [2] which was updated from TS 37.544 [3] based on RAN5 work [4][5].

Given the limited timeline to finalize the MU work for TR 38.884 [2], the effort was focused on most critical updates to MU descriptions and overall clean-up, but there was not a detailed analysis of each of the values used to derive the example expanded uncertainty.

Following the agreed MU assessment framework [6] for TS 38.561 [1], Table 2-1 and Table 2-2 present the list of uncertainty contributors, for TRP and TRS respectively, extracted from TS 38.561 [1] with an additional status column to reflect the items that:

- The description was updated from TS 37.544.

- The example uncertainty value was updated from TS 37.544.

- The uncertainty value need to or could be revised.

Table 2-1: Uncertainty contributions in TRP measurement for anechoic chamber method (Table A.4.3.1-1 of TS 38.561 [1])

| UID | Description of uncertainty contribution | Details in clause | Status |
| --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |
| 1 | Mismatch of receiver chain | A.4.2.1 | b). The value was updated to show the right probability distribution. c). The assumptions to derive the value (i.e. 0dB cable attenuation) does not seem realistic. |
| 2 | Insertion loss of receiver chain | A.4.2.2 | - |
| 3 | Influence of the measurement antenna cable | A.4.2.3 | - |
| 4 | Measurement Receiver: uncertainty of the absolute level | A.4.2.4 | c). The value could be updated to reflect the performance of state-of-the-art equipment |
| 5 | Measurement distance | A.4.2.7 | a) and b) to reflect the impact of the change in minimum range length definition |
| 6 | Quality of quiet zone | A.4.2.8 | a). c). The value could be updated to reflect realistic performance from labs with the revised definition. |
| 7 | DUT Tx-power drift | A.4.2.9 | c). The definition and value were not updated from TS 37.544, which refer to studies performed for 3G UEs. |
| 8 | Uncertainty related to the use of phantoms | A.4.2.11 | c). The definition and value could be updated to cover new hand phantoms (i.e. wide grip hand) |
| 9 | Coarse sampling grid | A.4.2.12 | c). The definition and value were not updated from TS 37.544, which refer to studies performed for 3G UEs. |
| 10 | Random uncertainty | A.4.2.13 | c). The definition and value were not updated from TS 37.544, which refer to old studies. |
| 11 | Frequency Response | A.4.2.14 | a). This term was new compared to TS 37.544 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |
| 12 | Uncertainty of network analyzer | A.4.2.15 | c). The value could be updated to reflect the performance of state-of-the-art equipment |
| 13 | Mismatch of receiver chain | A.4.2.1 | - |
| 14 | Insertion loss of receiver chain | A.4.2.2 | - |
| 15 | Mismatch in the connection of calibration antenna | A.4.2.1 | - |
| 16 | Influence of the calibration antenna feed cable | A.4.2.3 | c). The value could be updated to reflect realistic performance from labs. |
| 17 | Influence of the measurement antenna cable | A.4.2.3 | - |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | A.4.2.16 | c). The value could be updated to reflect the performance of state-of-the-art equipment |
| 19 | Measurement distance | A.4.2.7 | a) and b) to reflect the impact of the change in minimum range length definition |
| 20 | Quality of the Quiet Zone | A.4.2.8 | a). c). The value could be updated to reflect realistic performance from labs with the revised definition. |

Table 2-2: Uncertainty contributions in TRS measurement for anechoic chamber method (Table A.4.4.1-1 of TS 38.561 [1])

| UID | Description of uncertainty contribution | Details in clause | Status |
| --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |
| 1 | Mismatch of transmitter chain | A.4.2.1 | b). The value was updated to show the right probability distribution. c). The assumptions to derive the value (i.e. 0dB cable attenuation) does not seem realistic. |
| 2 | Insertion loss of transmitter chain | A.4.2.2 | - |
| 3 | Influence of the measurement antenna cable | A.4.2.3 | - |
| 4 | Communication Tester: uncertainty of the absolute output level | A.4.2.5 | c). The value could be updated to reflect the performance of state-of-the-art equipment |
| 5 | Sensitivity measurement: output level step resolution | A.4.2.6 | b) to reflect the minimum step used for sensitivity search. |
| 6 | Measurement distance | A.4.2.7 | a) and b) to reflect the impact of the change in minimum range length definition |
| 7 | Quality of quiet zone | A.4.2.8 | a). c). The value could be updated to reflect realistic performance from labs with the revised definition. |
| 8 | DUT sensitivity drift | A.4.2.10 | c). The definition and value were not updated from TS 37.544, which refer to studies performed for 3G UEs. |
| 9 | Uncertainty related to the use of phantoms | A.4.2.11 | c). The definition and value could be updated to cover new hand phantoms (i.e. wide grip hand) |
| 10 | Coarse sampling grid | A.4.2.12 | c). The definition and value were not updated from TS 37.544, which refer to studies performed for 3G UEs. |
| 11 | Random uncertainty | A.4.2.13 | c). The definition and value were not updated from TS 37.544, which refer to old studies. |
| 12 | Frequency Response | A.4.2.14 | a). This term was new compared to TS 37.544 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |
| 13 | Uncertainty of network analyzer | A.4.2.15 | c). The value could be updated to reflect the performance of state-of-the-art equipment |
| 14 | Mismatch of transmitter chain | A.4.2.1 | - |
| 15 | Insertion loss of transmitter chain | A.4.2.2 | - |
| 16 | Mismatch in the connection of calibration antenna | A.4.2.1 | - |
| 17 | Influence of the calibration antenna feed cable | A.4.2.3 | c). The value could be updated to reflect realistic performance from labs. |
| 18 | Influence of the measurement antenna cable | A.4.2.3 | - |
| 19 | Uncertainty of the absolute gain/radiation efficiency of the calibration antenna | A.4.2.16 | c). The value could be updated to reflect the performance of state-of-the-art equipment |
| 20 | Measurement distance | A.4.2.7 | a) and b) to reflect the impact of the change in minimum range length definition |
| 21 | Quality of quiet zone | A.4.2.8 | a). c). The value could be updated to reflect realistic performance from labs with the revised definition. |

For those marked with c), further review of several MU values in Stage 2 used in the tables is presented in next section.

# MU values analysis

## Mismatch of receiver/transmitter chain (A.4.2.1)

As shown in Table 2-1 and Table 2-2 above, the values were corrected to show the right probability distribution, but they are derived assuming:

- For TRP

  - 0dB cable attenuation.

  - Гpower meter < 0.05, reflection coefficient for power meter.

  - Гmeasurement antenna <0.16, reflection coefficient for the measurement antenna.

- For TRS

  - 0dB cable attenuation.

  - ГCommTester <0.13, reflection coefficient for the communication tester.

  - Г antenna connection <0.03, reflection coefficient for the measurement antenna.

Obviously, the 0dB cable attenuation is an unrealistic case but the remaining assumptions also require to be update to reflect existing test system implementations for NR FR1.

Proposal 1: all the assumptions for mismatch of receiver/transmitter chain shall be updated.

## Measurement Receiver: uncertainty of the absolute level (A.4.2.4)

A power meter is used as base assumption in the example tables, although frequency selective equipment (e.g. Spectrum Analyzers, Communication Testers) can be used as measurement receivers and are more relevant for dual connectivity measurements (i.e. EN-DC). Therefore, the uncertainty value considered for the example is too restrictive.

Considering a 0.2dB absolute level uncertainty, assuming actual distribution, is more realistic to cover a variety of receiver options.

Proposal 2: consider a 0.2dB absolute level uncertainty, with actual distribution, for the uncertainty source Measurement Receiver: uncertainty of the absolute level.

## Communication Tester: uncertainty of the absolute output level (A.4.2.5)

In this case, a 1dB uncertainty value is considered in the example tables assuming a rectangular distribution, although typical performance in the communication tester datasheet is expressed as expanded uncertainty. Therefore, it is proposed to change the probability distribution to actual and assume the 1dB uncertainty value is considered standard uncertainty.

Proposal 3: consider an actual distribution for the uncertainty source Communication Tester: uncertainty of the absolute output level.

Feedback from TE vendors is recommended to confirm example value reflects the performance of state-of-the-art equipment for NR FR1.

Observation 1: feedback from TE vendors is recommended to confirm if the example value for the uncertainty of the absolute output level for Communication Testers reflects the performance of state-of-the-art equipment for NR FR1.

## Quality of quiet zone (A.4.2.8)

As mentioned in Table 2-1 and Table 2-2 above, the description was updated from the one in TS 37.544 [3] although the example value was not updated. It is therefore recommended to get feedback from labs to reflect realistic performance with the revised definition.

Observation 2: feedback from labs is recommended to confirm if the example value for the Quality of quiet zone uncertainty reflects realistic performance.

## DUT Tx-power/sensitivity drift (A.4.2.9/A.4.2.10)

The definition and value were not updated from TS 37.544 [3], which refer to studies performed for 3G UEs and assume a 0.2dB value with rectangular distribution.

Observation 3: feedback from UE manufacturers is recommended to confirm if the example values for the DUT Tx-power/sensitivity drift uncertainty reflects realistic performance.

## Random uncertainty (A.4.2.13)

As mentioned in Table 2-1 and Table 2-2 above, the definition and example value were not updated from TS 37.544 [3] and reflect old studies. It is therefore recommended to get feedback from labs to update the example values, and/or the definition, to reflect realistic performance of existing devices.

Observation 4: feedback from labs is recommended to update the example value, and/or the definition, for the Random uncertainty to reflect realistic performance of existing devices.

# Conclusion

In this contribution, the following observations and proposals were made:

Observation 1: feedback from TE vendors is recommended to confirm if the example value for the uncertainty of the absolute output level for Communication Testers reflects the performance of state-of-the-art equipment for NR FR1.

Observation 2: feedback from labs is recommended to confirm if the example value for the Quality of quiet zone uncertainty reflects realistic performance.

Observation 3: feedback from UE manufacturers is recommended to confirm if the example values for the DUT Tx-power/sensitivity drift uncertainty reflects realistic performance.

Observation 4: feedback from labs is recommended to update the example value, and/or the definition, for the Random uncertainty to reflect realistic performance of existing devices.

Proposal 1: all the assumptions for mismatch of receiver/transmitter chain shall be updated.

Proposal 2: consider a 0.2dB absolute level uncertainty, with actual distribution, for the uncertainty source Measurement Receiver: uncertainty of the absolute level.

Proposal 3: consider an actual distribution for the uncertainty source Communication Tester: uncertainty of the absolute output level.

The proposals in this contribution are implemented in the associated Text Proposal [7].

# References

- 3GPP TS 38.561 v0.2.0 (March 2023)

- 3GPP TR 38.834 v17.2.0 (September 2022)

- 3GPP TS 37.544 v16.2.0 (September 2021)

- R5-221623, “TP to TR 38.834 for Annex B MU tables”, MVG Industries, Rohde & Schwarz, 3GPP TSG RAN5 Meeting #94-e, February 2022

- R5-221624, “TP to TR 38.834 on MU General section and descriptions”, Rohde & Schwarz, MVG Industries, 3GPP TSG RAN5 Meeting #94-e, February 2022

- R5-227881, “MU contributors and workplan for NR FR1 TRP-TRS (Rel.17)”, Rohde & Schwarz, 3GPP TSG RAN5 Meeting #97, November 2022

- R5-232587, “MU values for NR FR1 TRP-TRS”, Rohde & Schwarz, 3GPP TSG RAN5 Meeting #99, May 2023