3GPP TSG RAN WG5 Meeting #99	R5-232587

# Incheon, South Korea, 22 – 26 May 2023

| CR-Form-v12.2 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.561 | CR | N/A | rev | - | Current version: | 0.2.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | TP to TS 38.561 on MU contents |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | ROHDE & SCHWARZ |  |  |  |  |  |  |  |  |  |
| Source to TSG: | - |  |  |  |  |  |  |  |  |  |
| Work item code: | NR_FR1_TRP_TRS-UEConTest |  |  |  |  |  | Date: |  |  | 2023-02-18 |
| Category: | - |  |  |  |  |  | Release: |  |  | Rel-17 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-16 (Release 16) Rel-17 (Release 17) Rel-18 (Release 18) Rel-19 (Release 19) |  |
| Reason for change: |  | Update of Measurement Uncertainty example tables |  |  |  |  |  |  |  |  |
| Summary of change: |  | - Updated value for “Measurement Receiver: uncertainty of the absolute level” - Updated uncertainty distributions for “Measurement Receiver: uncertainty of the absolute level” and “Communication Tester: uncertainty of the absolute output level” |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | RAN5 test specification remains incomplete. |  |  |  |  |  |  |  |  |
| Clauses affected: |  |  |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  |  |  |  |  |  |  |  |  |  |
| This CR's revision history: |  |  |  |  |  |  |  |  |  |  |

< Unchanged Text Deleted >

< Beginning of Changes >

## A.4.3	Total Radiated Power (TRP)

#### A.4.3.1	Anechoic Chamber Method

The uncertainty contributions related to TRP for the Anechoic Chamber method are listed in Table A.4.3.1-1. Example uncertainty budget is presented in Table A.4.3.1-2.

Table A.4.3.1-1: Uncertainty contributions in TRP measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure A.3.1-1, Figure A.3.1-2) |  |  |
| 1 | Mismatch of receiver chain | A.4.2.1 |
| 2 | Insertion loss of receiver chain | A.4.2.2 |
| 3 | Influence of the measurement antenna cable | A.4.2.3 |
| 4 | Measurement Receiver: uncertainty of the absolute level | A.4.2.4 |
| 5 | Measurement distance | A.4.2.7 |
| 6 | Quality of quiet zone | A.4.2.8 |
| 7 | DUT Tx-power drift | A.4.2.9 |
| 8 | Uncertainty related to the use of phantoms | A.4.2.11 |
| 9 | Coarse sampling grid | A.4.2.12 |
| 10 | Random uncertainty | A.4.2.13 |
| 11 | Frequency Response | A.4.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure A.3.2-1) |  |  |
| 12 | Uncertainty of network analyzer | A.4.2.15 |
| 13 | Mismatch of receiver chain | A.4.2.1 |
| 14 | Insertion loss of receiver chain | A.4.2.2 |
| 15 | Mismatch in the connection of calibration antenna | A.4.2.1 |
| 16 | Influence of the calibration antenna feed cable | A.4.2.3 |
| 17 | Influence of the measurement antenna cable | A.4.2.3 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | A.4.2.16 |
| 19 | Measurement distance | A.4.2.7 |
| 20 | Quality of the Quiet Zone | A.4.2.8 |

Table A.4.3.1-2: Example of uncertainty budget for TRP hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гpower meter <0.05 Гmeasurement antenna <0.16 | [0.07] | U-shaped | 1.41 | 1 | [0.05] |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | [0] | Rectangular | 1.73 | 1 | [0.00] |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | [0] | Rectangular | 1.73 | 1 | [0.00] |
| 4 | Measurement Receiver: uncertainty of the absolute level | Power Meter | 0.2 | Actual | 1 | 1 | 0.2 |
| 5 | Measurement distance | d=1.6m, Δd=0.05m | [0.27] | Rectangular | 1.73 | 1 | [0.16] |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | [0.5] | Actual | 1 | 1 | [0.50] |
| 7 | DUT Tx-power drift | Drift | [0.2] | Rectangular | 1.73 | 1 | [0.12] |
| 8 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | [0.32] | Rectangular | 1.73 | 1 | [0.18] |
| 9 | Coarse sampling grid | Negligible 15° sampling grid | [0] | Actual | 1 | 1 | [0.00] |
| 10 | Random Uncertainty | Monoblock, clamshell and PDA design used for testing | [0.81] | Rectangular | 1.73 | 1 | [0.47] |
| 11 | Frequency Response | Average path loss corrected | [0] | Rectangular | 1.73 | 1 | [0.00] |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | Manufacturer’s uncertainty | [0.5] | Rectangular | 1.73 | 1 | [0.29] |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | [0] | U-shaped | 1.41 | 1 | [0.00] |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | [0] | Rectangular | 1.73 | 1 | [0.00] |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | [0] | U-shaped | 1.41 | 1 | [0.00] |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | [0.3] | Rectangular | 1.73 | 1 | [0.17] |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | [0] | Rectangular | 1.73 | 1 | [0.00] |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | [0.5] | Rectangular | 1.73 | 1 | [0.29] |
| 19 | Measurement distance | Dipole: aligned with phase center | [0] | Rectangular | 1.73 | 1 | [0.00] |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | [0.5] | Rectangular | 1.73 | 1 | [0.29] |
| Combined standard uncertainty |  |  |  |  |  |  | [0.93] |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | [1.82] |

## A.4.4	Total Radiated Sensitivity (TRS)

#### A.4.4.1	Anechoic Chamber Method

The uncertainty contributions related to TRS for the Anechoic Chamber method are listed in Table A.4.4.1-1. Example uncertainty budget is presented in Table A.4.4.1-2.

Table A.4.4.1-1: Uncertainty contributions in TRS measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure A.3.1-1, Figure A.3.1-2) |  |  |
| 1 | Mismatch of transmitter chain | A.4.2.1 |
| 2 | Insertion loss of transmitter chain | A.4.2.2 |
| 3 | Influence of the measurement antenna cable | A.4.2.3 |
| 4 | Communication Tester: uncertainty of the absolute output level | A.4.2.5 |
| 5 | Sensitivity measurement: output level step resolution | A.4.2.6 |
| 6 | Measurement distance | A.4.2.7 |
| 7 | Quality of quiet zone | A.4.2.8 |
| 8 | DUT sensitivity drift | A.4.2.10 |
| 9 | Uncertainty related to the use of phantoms | A.4.2.11 |
| 10 | Coarse sampling grid | A.4.2.12 |
| 11 | Random uncertainty | A.4.2.13 |
| 12 | Frequency Response | A.4.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure A.3.2-1) |  |  |
| 13 | Uncertainty of network analyzer | A.4.2.15 |
| 14 | Mismatch of transmitter chain | A.4.2.1 |
| 15 | Insertion loss of transmitter chain | A.4.2.2 |
| 16 | Mismatch in the connection of calibration antenna | A.4.2.1 |
| 17 | Influence of the calibration antenna feed cable | A.4.2.3 |
| 18 | Influence of the measurement antenna cable | A.4.2.3 |
| 19 | Uncertainty of the absolute gain/radiation efficiency of the calibration antenna | A.4.2.16 |
| 20 | Measurement distance | A.4.2.7 |
| 21 | Quality of quiet zone | A.4.2.8 |

Table A.4.4.1-2: Example of uncertainty budget for TRS hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | ГCommTester <0.13 Г antenna connection <0.03 | [0.07] | U-shaped | 1.41 | 1 | [0.05] |  |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | [0] | Rectangular | 1.73 | 1 | [0.00] |  |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | [0] | Rectangular | 1.73 | 1 | [0.00] |  |
| 4 | Communication Tester: uncertainty of the absolute output level | Manufacturer’s data sheet | [1] | Actual | 1 | 1 | [1.00] |  |
| 5 | Sensitivity measurement: output level step resolution | Step of 0.5 dB | [0.25] | Rectangular | 1.73 | 1 | [0.14] |  |
| 6 | Measurement distance | d=1.6m, Δd=0.05m | [0.27 | Rectangular | 1.73 | 1 | [0.16] |  |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | [0.5] | Actual | 1 | 1 | [0.5] |  |
| 8 | DUT sensitivity drift | Drift measurement | [0.2] | Rectangular | 1.73 | 1 | [0.12] |  |
| 9 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | [0.32] | Rectangular | 1.73 | 1 | [0.18] |  |
| 10 | Coarse sampling grid | 30° sampling grid | [0.15] | Actual | 1 | 1 | [0.15] |  |
| 11 | Random uncertainty | Monoblock, clamshell and PDA used for testing | [0.91] | Rectangular | 1.73 | 1 | [0.53] |  |
| 12 | Frequency Response | Included in the output level step resolution | [0] | Rectangular | 1.73 | 1 | [0.00] |  |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers NA setup | [0.5] | Rectangular | 1.73 | 1 | [0.29] |  |
| 14 | Mismatch of transmitter chain | Taken in to account in VNA setup uncertainty | [0] | U-shaped | 1.41 | 1 | [0.00] |  |
| 15 | Insertion loss of transmitter chain | Systematic with Stage 2 (=> cancels) | [0] | Rectangular | 1.73 | 1 | [0.00] |  |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | [0] | U-shaped | 1.41 | 1 | [0.00] |  |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with dipole | [0.3] | Rectangular | 1.73 | 1 | [0.17] |  |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | [0] | Rectangular | 1.73 | 1 | [0.00] |  |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | [0.5] | Rectangular | 1.73 | 1 | [0.29] |  |
| 20 | Measurement distance | Dipole: aligned with phase center | [0] | Rectangular | 1.73 | 1 | [0.00] |  |
| 21 | Quality of quiet zone | Peak-to-null ripple | [0.5] | Rectangular | 1.73 | 1 | [0.29] |  |
| Combined standard uncertainty |  |  |  |  |  |  | [1.39] |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | [2.72] |

< End of Changes >