3GPP TSG-RAN WG5 #107		R5-251852
Malta, ML, 19th May 2025 – 23rd May 2025

Agenda item:	5.3.6.9

Source:	Nokia

Title:	MU assessment of RRM FR2 for PC6 devices

Document for:	Discussion & Endorsement

# 1.	Introduction

At RAN5#97-e meeting, RAN WG5 endorsed Rel-17 NR_HST_FR2-UEConTest [1] WI which included the performing MU assessment for RF and RRM for PC6 devices. In RAN5#104 meeting, MU values for PC6 were agreed for DFF setup which can be leveraged to RRM FR2 PC6 MU for IFF and enhanced IFF.

This paper presents the results of our MU assessment and proposes to add the RRM PC6 MU tables based on RF decisions.

# 2.	Discussion

The uncertainty assessment for EIS measurement for PC6 UEs for f ≤ 32.125 GHz for RF REFSENS test case has been completed in B.19.2 of TR 38.903 [2].

Observation 1: Table B.19.2-2 includes uncertainty assessment for PC3 and PC6, but all defined frequencies don’t apply to PC6.

Proposal 1: Update the table title to clearly state the applicable frequencies for PC3 and PC6.

Table B.19.2-2: Uncertainty assessment for EIS measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) and PC6 (f=23.45GHz, 32.125GHz, Quiet Zone size ≤ 30 cm) UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 7) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |
| 3 | Quality of Quiet Zone (NOTE 7) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 4 | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level (23.45GHz <= f <= 40.8GHz) | 2.9 | Normal | 2.00 | 1.45 |
| 6 | gNB uncertainty on absolute level (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 6) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.00 (NOTE 4) 0.08 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.05 (NOTE 5) |
| 15 | Influence of spherical coverage grid (NOTE 4) | 0.12 | Actual | 1 | 0.12 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
| 19 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 7) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |
| 23 | Quality of quiet zone for calibration process (NOTE 7) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 27 | Systematic error related to beam peak search (NOTE 5) (PC3) |  |  |  | 0.5 |
| 27 | Systematic error related to beam peak search (NOTE 5) (PC6) |  |  |  | 0.7 |
| 28 | Systematic error related to EIS spherical coverage (NOTE 4) |  |  |  | DL power step size, 0.2 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIS Expanded uncertainty (23.45GHz <= f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] (PC3) |  |  |  |  | 5.36 |
| EIS Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] (PC3) |  |  |  |  | 6.34 |
| EIS Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] (PC6) |  |  |  |  | 5.56 |
| EIS Spherical coverage Expanded uncertainty (23.45GHz <= f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] (PC3) |  |  |  |  | 5.07 |
| EIS Spherical coverage Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] (PC3) |  |  |  |  | 6.04 |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: Void. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: This contributor shall only be considered for spherical EIS measurements. NOTE 5: This contributor shall only be considered for EIS measurements. NOTE 6: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 7: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. |  |  |  |  |  |

Observation 2: Clause E.3.1.3 of TR 38.903 [3] for IFF test setup doesn’t include the uncertainty assessment for DL AWGN absolute power for power class 6 devices.

Proposal 2: New table E.3.1.3-4 shall be added for IFF test setup for PC6 in Ref. [5].

E.3.1.3-4: Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, Quiet Zone size ≤ 30 cm) for PC6

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 16 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 17 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 19 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 20 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 21 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 22 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 23 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 24 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 25 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 26 | Systematic error related to beam peak search |  |  |  | 0.7 |
| Total measurement uncertainty |  |  |  |  | Value |
| DL AWGN absolute power expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.56 |
| NOTE 1: The analysis was done only for the case of operating in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in Annex D.2 of TR 38.810 [13] for Quiet Zone size less or equal to 30 cm. NOTE 5: The values in this table have been derived for DL powers above and equal to REFSENS. The values might need to be revisited for power levels below REFSENS |  |  |  |  |  |

Observation 3: Clause E.3.1.4 of TR 38.903 [3] for Enhanced IFF test setup doesn’t include the uncertainty assessment for DL AWGN absolute power for power class 6 devices.

Proposal 3: New table E.3.1.4-4 shall be added for Enhanced IFF test setup for PC6 in Ref. [5].

Table E.3.1.4-4: Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, Quiet Zone size ≤ 30 cm) for PC6

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 28 | Systematic error related to beam peak search |  |  |  | 0.7 |
| Total measurement uncertainty |  |  |  |  | Value |
| DL AWGN absolute power expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.62 |
| NOTE 1: The analysis was done only for the case of operating in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in Annex M of TR 38.521-2 [7] for Quiet Zone size less or equal to 30 cm. NOTE 5: The values in this table have been derived for DL powers above and equal to REFSENS. The values might need to be revisited for power levels below REFSENS |  |  |  |  |  |

Proposal 4: Apply a value of 5.65 dB for the MU threshold for DL AWGN absolute power for RRM FR2 for PC6.

Table E.3.1-1: MU threshold for DL AWGN absolute power for RRM FR2

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | As configured in the test case | 5.65 dB2 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.65 dB2 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | As configured in the test case | 5.65 dB3 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.65 dB3 |
| PC6 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | As configured in the test case | 5.65 dB4 |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table E.3.1.3-2 for PC3 UEs, Table E.3.1.3-3 for PC1 UEs and E.3.1.3-4 for PC6 UEs. NOTE 2: If the TT analysis for a specific test case based on this MU value results in an unsolvable conflict, making the test case untestable, even after the alternative solutions listed in clause A.4 have been considered for the test case in TS 38.133 [6] Annex A, the TT analysis shall be repeated using a lower MU value, taking into account the lower values defined in this clause. The test case will be applicable for the subset of the test systems meeting this reduced MU Threshold. NOTE 3: For simpler MU definition in TS 38.533 the PC3 Threshold MU value is reused for PC1. As the difference of PC1 MU of 5.63 dB compared to PC3 Threshold MU of 5.65 dB is insignificant. NOTE 4: For a simpler MU definition in TS 38.533, the PC3 Threshold MU value is reused for PC6. The difference between PC6 MU of 5.62 dB and PC3 Threshold MU of 5.65 dB is insignificant. |  |  |  |  |

The accompanying CR for the update of TR 38.903 [3] is contained in Ref. [5].

The accompanying CR for the update of TS 38.533 [4] is contained in Ref. [6].

# 3.	Conclusion

Observation 1: Table B.19.2-2 includes uncertainty assessment for PC3 and PC6, but all defined frequencies don’t apply to PC6.

Observation 2: Clause E.3.1.3 of TR 38.903 [3] for IFF test setup doesn’t include the uncertainty assessment for DL AWGN absolute power for power class 6 devices.

Observation 3: Clause E.3.1.4 of TR 38.903 [3] for Enhanced IFF test setup doesn’t include the uncertainty assessment for DL AWGN absolute power for power class 6 devices.

Proposal 1: Update the table title to clearly state the applicable frequencies for PC3 and PC6.

Proposal 2: New table E.3.1.3-4 shall be added for IFF test setup for PC6.

Proposal 3: New table E.3.1.4-4 shall be added for Enhanced IFF test setup for PC6.

Proposal 4: Apply a value of 5.65 dB for the MU threshold for DL AWGN absolute power for RRM FR2 for PC6.

# 4.	References

- R5-223323, “New WID on UE Conformance - NR support for high-speed train scenario in frequency range 2 (FR2)”, 3GPP TSG-RAN WG5 Meeting #95

- R5-244173, “PC6 FR2 MU definition in 38.903”, Keysight Technologies, 3GPP TSG-RAN WG5 Meeting #104

- TR 38.903 “Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases”, V18.6.0 (2025-03)

- TS 38.533 “User Equipment (UE) conformance specification; Radio Resource Management (RRM)”, V18.6.1 (2025-03)

- R5-251853, “Addition of MU for RRM FR2 in case of PC6”, 3GPP TSG-RAN WG5 Meeting #107

- R5-251854, “Addition of MU values for RRM FR2 PC6”, 3GPP TSG-RAN WG5 Meeting #107