3GPP TSG-RAN WG4 Meeting #100-e		                    R4-2113030

Online, 16th Aug – 27th Aug, 2021

Source: 	Keysight Technologies

Title: 	about BS conformance test FR2 Rx out of band test MU calculation
Agenda Item:	6.1.9.1

Document for:	Agreement

# Introduction

In previous RAN4#99e meeting, correction on FR2 Rx OOB MU calculation proposed and agreed, CR for TS38.141-2 also agreed. For this time, draft CR for TR37.941 is prepared to make correction on calculation equation for FR2 Rx OOB. This document describes what was corrected.

# Description

In previous RAN4#99e meeting, correction on FR2 Rx OOB MU calculation proposed and agreed, CR for TS38.141-2 also agreed. Regarding with other Receiver Test calculation and equations described in TR37.941, we verified there is no other such error (double count) exist.

In TR37.941 which captures equation for calculating total test system MU on FR2 Receiver Out of band Blocking test, unfortunately, there was error which Test Equipment MU was RSS summed twice. In more detail, as following from TR37.941, Rx OOB MU for FR2 was calculated with EIS, Modulated signal source (TestEquipment), OOB interferer (CW signal source), PA and Broadband noise. While MU for EIS includes wanted signal (modulated signal generator) uncertainty in value, another Modulated signal source (TestEquipment) value summed by RSS. In out of band blocking test, interferer is OOB interferer which also summed by RSS, so that this additional modulated signal source is not for interferer then it should be removed from calculation, double counted. Please note, out-of-band blocking test uses wanted signal and out of band interferer which is CW signal for FR2 testing.

Excerpt from TR37.941,

### 14.2.4	MU value derivation, FR2

It has been agreed that the MU for the out-of-band blocking requirement can be calculated as follows:

With

And

Substituting the variables above into the formula, the MU in FR2 for the out-of-band blocking requirement can be calculated as 4.1 dB, as shown in table 14.2.4-1 below.

Table 14.2.4-1: MU for out-of-band blocking

| Test System Uncertainty | Standard uncertainty ui (dB) IAC, CATR |
| --- | --- |
| MUEIS (Expanded uncertainty) | 2.4 |
| MUTestEquipment (Uncertainty of the RF signal generator) | 0.9 |
| MUOOBint (Additional uncertainty for the OOB interferer signal) | 1.1 |
| MUPA (Uncertainty due to use of PA) | 0.2 |
| Broadband noise effect (Impact of interferer broadband noise) | 0.4 |
| Combined standard uncertainty (1σ) | 2.09 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) | 4.10 |

End of excerpt text

The math should be modified as following (removing 2nd Test Equipment MU),

MUOOB=MUEIS2+MUOOBint2+MUPA2+ Broadband noise effect

MU calculation should be modified as following, (up to 43.5GHz) (removing 2nd Test Equipment MU)

Table 1, FR2 out of band blocking MU value

| Test System Uncertainty | Standard uncertainty ui (dB) IAC, CATR |
| --- | --- |
| MUEIS (Expanded uncertainty) | 2.4 |
| MUOOBint (Additional uncertainty for the OOB interferer signal) | 1.1 |
| MUPA (Uncertainty due to use of PA) | 0.2 |
| Broadband noise effect (Impact of interferer broadband noise) | 0.4 |
| Combined standard uncertainty (1σ) | 1.86 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) | 3.64 |

# References

[1] R4-2108501 CR to 38.141-2: BS conformance test, FR2 Rx OOB test MU value correction (4.1.2.3), Keysight Technologies

[2] R4-2113028 Draft CR to 37.941: BS OTA test, FR2 Rx OOB test MU value Math correction (14.2.4), Keysight Technologies