3GPP TSG-RAN WG4 Meeting #103-e	R4-2208321
E-meeting, May 09 – May 20, 2022

Agenda item:	9.1.2.1

Source:	CAICT

Title:	Summary of FR1 MIMO OTA lab alignment results

Document for:	Discussion

1	Introduction

The NR MIMO OTA WI is scheduled to finalize FR1 MIMO OTA lab alignment and reviewing the PAD measurement results during RAN4 #103e meeting in May 2022 [1].

Up to now, 5 labs have completed PAD measurement and submitted lab alignment results [2-6]. The last lab (Apple) has not received the PADs yet due to COVID, and thus is not able to submit the results currently.

This contribution presents a summary and analysis of the submitted PAD measurement results. This contribution can be updated based on discussion outcomes of RAN4 #103e meeting, and will be updated in the future to include Apple’s PAD measurement results.

2	Discussion

The reference values and pass/fail limits for FR1 MIMO OTA lab alignment have been discussed in prior RAN4 meetings, the agreements are as below:

Framework for FR1 MIMO OTA lab alignment activity [7]

  - The reference value of each PAD should be the average of the PAD measurement results submitted on or before 12:00 UTC 30th April 2022, based on the condition at least 3 labs’ results collected. Submission with measurement data after 12:00 UTC 30th April can be considered for lab alignment, but will not change the reference TRMS value.

Issue 2-1-2: Pass/fail limit for FR1 MIMO OTA lab alignment [8]

Agreements:

The pass/fail limit for lab PAD alignment: the maximum deviation of TRMS between each performance alignment lab and Reference Value

  - Starting point: +/- 0.75MU, i.e., +/- 2.25 dB for band <3GHz and +/- 2.55 dB for band >3GHz.

    - RAN4 aims to minimize the limit into reasonable range based on the collected lab alignment data

  - Further check is allowed after having some PAD test results next RAN4 meeting.

The summary and analysis of the lab alignment results are based on the following conditions:

- The reference value of each PAD at each band is the linear average (in dB) of the PAD measurement results submitted by the 5 labs in [2-6]

- The pass/fail limit for lab PAD alignment is +/- 0.75MU, i.e., +/- 2.25 dB for band <3GHz and +/- 2.55 dB for band >3GHz

Summary of submitted FR1 MIMO OTA lab alignment results are shown in Table 1.

Table 1. Summary of submitted FR1 MIMO OTA lab alignment results

| Device | Band | TRMS measurement result [dBm/30kHz] |  |  |  |  |  | Average approach | Reference value | Max-Min deviation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 |  |  |  |
| PAD_n41_1 | n41 | -96.43 | -97.61 | -98.20 | -97.45 | -96.88 | NA | Linear average | -97.31 | 1.77 |
| PAD_n41_2 | n41 | -99.30 | -97.80 | -100.02 | -99.96 | -99.62 | NA |  | -99.34 | 2.22 |
| PAD_n41_3 | n41 | -96.31 | -97.39 | -97.81 | -96.53 | -96.74 | NA |  | -96.96 | 1.50 |
| PAD_n78_1 | n78 | -96.02 | -96.54 | -96.44 | -96.10 | -96.53 | NA |  | -96.33 | 0.52 |
| PAD_n78_2 | n78 | -95.42 | -95.95 | -96.11 | -96.48 | -96.66 | NA |  | -96.12 | 1.24 |
| PAD_n78_3 | n78 | -99.06 | -97.42 | -99.53 | -99.08 | -99.54 | NA |  | -98.93 | 2.12 |
| Device | Band | TRMS offset [dBm/30kHz] |  |  |  |  |  | Pass/fail limit |  |  |
|  |  | Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 |  |  |  |
| PAD_n41_1 | n41 | 0.88 | -0.30 | -0.89 | -0.14 | 0.43 | NA | +/- 0.75 MU, i.e., +/- 2.25 dB |  |  |
| PAD_n41_2 | n41 | 0.04 | 1.54 | -0.68 | -0.62 | -0.28 | NA |  |  |  |
| PAD_n41_3 | n41 | 0.65 | -0.43 | -0.86 | 0.43 | 0.22 | NA |  |  |  |
| PAD_n78_1 | n78 | 0.31 | -0.21 | -0.12 | 0.23 | -0.20 | NA | +/- 0.75 MU, i.e., +/- 2.55 dB |  |  |
| PAD_n78_2 | n78 | 0.70 | 0.17 | 0.01 | -0.36 | -0.54 | NA |  |  |  |
| PAD_n78_3 | n78 | -0.13 | 1.51 | -0.60 | -0.15 | -0.61 | NA |  |  |  |
| Lab alignment conclusion |  | [Pass] | [Pass] | [Pass] | [Pass] | [Pass] | TBA |  |  |  |

Figure 2 presents the TRMS measurement results, and Figure 3 illustrates the deviations between TRMS measurement values and reference values.

Figure 1. TRMS measurement results submitted by 5 labs

Figure 2. Deviations between TRMS measurement values and reference values

Observation 1: The maximum deviations between TRMS measurement values and reference values are 1.54 dB for band n41 and 1.51 dB for band n78, respectively.

Observation 2: 3GPP FR1 MIMO OTA lab alignment among the 5 labs (CAICT, CMCC&BUPT, Huawei, MediaTek, and Xiaomi) can be confirmed based on the following conditions:

- The reference value of each PAD at each band is the linear average (in dB) of the PAD measurement results submitted by the above 5 labs

- The pass/fail limit for lab PAD alignment is +/- 0.75MU, i.e., +/- 2.25 dB for band <3GHz and +/- 2.55 dB for band >3GHz

3	Conclusion

This contribution presents a summary and analysis of the submitted PAD measurement results.

Observation 1: The maximum deviations between TRMS measurement values and reference values are 1.54 dB for band n41 and 1.51 dB for band n78, respectively.

Observation 2: 3GPP FR1 MIMO OTA lab alignment among the 5 labs (CAICT, CMCC&BUPT, Huawei, MediaTek, and Xiaomi) can be confirmed based on the following conditions:

- The reference value of each PAD at each band is the linear average (in dB) of the PAD measurement results submitted by the above 5 labs

- The pass/fail limit for lab PAD alignment is +/- 0.75MU, i.e., +/- 2.25 dB for band <3GHz and +/- 2.55 dB for band >3GHz

References

- R4-2120683, “Updated time plan for FR1 lab alignment and requirement development”, CAICT, OPPO, 3GPP RAN4#101e, Nov. 2021.

- R4-2208320, “CAICT FR1 MIMO OTA lab alignment results”, CAICT, 3GPP RAN4#103e, May 2022.

- R4-2208412, “PAD testing results of CMCC & BUPT joint lab”, CMCC, BUPT, 3GPP RAN4#103e, May 2022.

- R4-2207654, “MIMO OTA lab alignment results”, Huawei, 3GPP RAN4#103e, May 2022.

- R4-2208319, “PAD test results for NR FR1 MIMO OTA lab alignment”, MediaTek, 3GPP RAN4#103e, May 2022.

- R4-2209514, “3GPP NR FR1 MIMO OTA Lab Alignment Result”, Xiaomi, 3GPP RAN4#103e, May 2022.

- R4-2203067, “Framework for FR1 MIMO OTA lab alignment activity”, CAICT, 3GPP RAN4#101-bis-e, Jan. 2022.

- R4-2207300, “WF on NR MIMO OTA”, vivo, CAICT, 3GPP RAN4#102e, Feb. 2022.