## Summary
| Expanded uncertainty (dB) | [6 empty] | A method exceeds agreed value |  |  | Comments
|  | Compact Antenna Test Range |  |  | Proposed value
|  | 52.6<f ≤71.0 GHz |  |  | 52.6<f ≤71.0 GHz |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz | 4.2<f≤6 GHz
EIRP - Normal conditions |  | 5.4305 |  |  | 5.4
EIRP - Extreme conditions |  | 5.9935 |  |  | 6
Power dynamics |  |  |  |  | 0.4 | [5 empty] | same as existing number
EVM (%) |  |  |  |  | 1.1 | [5 empty] | EVM proposed new value 0.1% increase
In-band TRP |  | 5.5402 |  |  | 5.5
ACLR- absolute |  | 5.5733 |  |  | 5.6
ACLR - relative |  | 5.5134 |  |  | 5.5
OBUE |  | 5.5776 |  |  | 5.6
Co-existance EM
Co-location EM
TX IMD
|  | Compact Antenna Test Range |  |  | Proposed value
|  |  | 71GHz<f ≤110GHz | 110GHz<f ≤142GHz |  | 71GHz<f ≤110GHz | 110GHz<f ≤142GHz
OOB EM |  | x | 6.899 | 6.9696 | x | 6.9 | 7
RX EM |  | x | 6.899 | 6.9696 | x | 6.9 | 7

## Change history
Date | Meeting | TDoc | CR | Rev | Cat | Subject/Comment | New version
2022-11-22 | RAN4#105 | R4-2219450 |  |  |  | Proposal, FR2-2 values, replacing original value for 24~29G range to 52.6 ~ 71 G range. Taking excel sheet from 37.941

## TE
Test Equipment errors | [9 empty] | Comments
UID | Uncertainty source | Uncertainty value (dB) |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 52.6<f ≤71.0 GHz |  |  |  | 52.6<f ≤71.0 GHz
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 2.36 |  | Gaussian | 1 | 2.36
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 2.36 |  | Gaussian | 1 | 2.36
C1-8 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - relative (ACLR) | 2.36 |  | Gaussian | 1 | 2.36
C1-9 | RF power measurement equipment standard uncertainty σ (dB) of the absolute level for a time domain wideband measurement for FR2 | 2.61 |  | Gaussian | 1 | 2.61
C1-2 | RF signal generator | 2.38 |  | Gaussian | 1 | 2.38
C1-3 | Uncertainty of the network analyzer | 0.87 |  | Gaussian | 1 | 0.87
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 |  | Rectangular | 1.7321 | 0.3
Test Equipment errors
UID | Uncertainty source | Uncertainty value (dB) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 52.6<f ≤71.0 GHz | 71.0GHz<f ≤110GHz | 110GHz<f ≤142GHz |  |  | 52.6<f ≤71.0 GHz | 71.0GHz<f ≤110GHz | 110GHz<f ≤142GHz
C1-7 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 2.36 | 2.39 | 2.49 | Gaussian | 1 | 2.36 | 2.39 | 2.49 | RAN4#94: RK: 40-60GHz MU is different for different methods: IAC=0.6 CATR=0.7 Reverb=1.15
C1-3 | Uncertainty of the network analyzer | 0.85 | 0.85 | 0.85 | Gaussian | 1 | 0.85 | 0.85 | 0.85
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 | 0.52 | 0.52 | Rectangular | 1.7321 | 0.30022 | 0.30022 | 0.30022
Conducted MU (minus missmatch) - up to 6 GHz
UID | Uncertainty source | Uncertainty value (dB) |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | f<3 GHz | 3<f<4.2 GHz |  |  | f<3 GHz | 3<f<4.2 GHz
23 | DL-RS MU derived from conducted spec | 0.41 | 0.56 | Gaussian | 1 | 0.41 | 0.56
23 | Total power dynamic range conducted unbcertainty | 0.2 | 0.2 | Gaussian | 1 | 0.2 | 0.2
23 | ACLR relative - MU of TE from conducted | 0.41 | 0.56 | Gaussian | 1 | 0.41 | 0.56
| ACLR abb q(conducted minus missmatch) | 1.0108 | 1.2321 | Gaussian | 1 | 1.0108 | 1.2321
| OBUE (conducted minus missmatch) | 0.75239 | 0.85703 | Gaussian | 1 | 0.75239 | 0.85703
27 | Additional (co-location) emissions - condcuted uncertainty (minus missmatch) | 1.02 | 1.28 | Gaussian | 1 | 1.02 | 1.28
| TX IMD - conducted measurement uncertainty | 1 | 1.1 | Gaussian | 1 | 1 | 1.1
Conducted MU (minus missmatch) - whole spectrum
UID | Uncertainty source | Uncertainty value (dB) |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz
27 | OOB emsions - Conducted Uncertainty (minus missmatch) | 1.02 | 2.04 | Gaussian | 1 | 1.02 | 2.04
27 | Rx emisions - Conducted Uncertainty (minus missmatch) | 1 | 1.99 | Gaussian | 1 | 1 | 1.99
Co-location errors
UID | Co-locoation error source | Uncertainty value (dB) |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz |  |  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz
| Uncertainty related to the selection of the CLTA (Note) | 1.5 | 1.5 | Rectangular | 1.7321 | 0.86603 | 0.86603
| Uncertainty related to the placement of the CLTA (Note) | 1.7 | 1.7 | Rectangular | 1.7321 | 0.9815 | 0.9815
| Uncertainty related to measuring close to noise floor - Emisions | 0.68 | 0.68 | Gaussian | 1 | 0.68 | 0.68
| Uncertainty related to measuring close to noise floor Tx OFF | 1 | 1 | Gaussian | 1 | 1 | 1
| Impedance mismatch between feeder cable and CLTA | 0.14 | 0.23 | U-Shaped | 1.4142 | 0.098995 | 0.16263
| Gain variations in LNA | 0.1 | 0.1 | Gaussian | 1 | 0.1 | 0.1
| Gain variations in measurement amplifier | 0.1 | 0.1 | Gaussian | 1 | 0.1 | 0.1

## CATR-Er
CATR | [10 empty] | Comments
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 52.6<f ≤71.0 GHz | 71.0GHz<f ≤110GHz | 110GHz<f ≤142GHz |  |  |  | 52.6<f ≤71.0 GHz | 71.0GHz<f ≤110GHz | 110GHz<f ≤142GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 | 0.2 |  | Exp. normal | 2 | 1 | 0.1 | 0.1 | 0
A2-18a | Misalignment and pointing error of BS (for TRP) | 0.3 | 0.3 | 0.3 | Exp. normal | 2 | 1 | 0.15 | 0.15 | 0.15
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 2.36 | 2.39 | 2.49 | Gaussian | 1 | 1 | 2.36 | 2.39 | 2.49
A2-2a | Standing wave between BS and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.14849 | RK: uncertainty changed to 0.03 dB
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01
A2-4a | QZ ripple experienced by BS | 0.4 | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.4 | RK: uncertainty value scorrect to match std uncertainty
A2-12 | Frequency flatness of test system | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25
A2-13 | QZ ripple with BS (extreme test conditions) | 0.7 | 0.7 |  | Gaussian | 1 | 1 | 0.7 | 0.7 | 0
A2-15 | radome loss variation | 0.5 | 0.5 |  | Gaussian | 1 | 1 | 0.5 | 0.5 | 0
A2-14 | wet radome loss variation | 0.9 | 0.9 |  | Gaussian | 1 | 1 | 0.9 | 0.9 | 0
A2-16 | Change in absorber behaviour | 0.5 | 0.5 |  | Gaussian | 1 | 1 | 0.5 | 0.5 | 0
C1-9 | RF power measurement equipment standard uncertainty σ (dB) of the absolute level for a time domain wideband measurement for FR2 | 2.61 | 0 |  | Gaussian | 1 | 1 | 2.61 | 0 | 0
A2-19 | Uncertainty of the LNA (FR2 only) | 1.5 | 1.5 | 1.5 | Gaussian | 1 | 1 | 1.5 | 1.5 | 1.5
A2-20 | Uncertainty of the mixer (FR2 only) | 2.25 | 2.25 | 2.35 | Gaussian | 1 | 1 | 2.25 | 2.25 | 2.35
A2-10 | Miscellaneous uncertainty | x | x | 0 | Rectangular | 1.7321 | 1 | #VALUE! | #VALUE! | 0
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.85 | 0.85 | 0.87 | Gaussian | 1 | 1 | 0.85 | 0.85 | 0.87
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 1.3718 | 1.6829 | 1.6122 | U-shaped | 1.4142 | 1 | 0.97 | 1.19 | 1.14 | RK: calcualtion corrected - same unceratinty value
A2-5b | Mismatch of receiver chain for low power receiver | 1.3718 | 1.6829 | 1.6122 | U-shaped | 1.4142 | 1 | 0.97 | 1.19 | 1.14 | Nok: these are non-zero in 37.843. can set the same as in 37.843?
A2-6 | Insertion loss of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 |  | RK: uncertainty changed to 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01
A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | 0.29 | U-shaped | 1.4142 | 1 | 0.14849 | 0.20506
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 | 0.52 | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.30022 |  | RK: calcualtion corrected - same unceratinty value
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  | RK: uncertainty changed to 0 Nok: these were non-zero in TRE37.843. Can set the same as in TR37.843?
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  | RK: uncertainty changed to 0 Nok: these were non-zero in TRE37.843. Can set the same as in TR37.843?
A2-9 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009
A2-11 | Switching uncertainty | 0.43 | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.24826 | 0.24826 |  | RK: uncertainty changed to 0.1
Note on C10: Richard Kybett: This is most liely a typo, the std uncertainty is correct just the uncertainty value has FR1 value (0.928) in it.

## EIRP
| Table 9.2.7-2: OTA test system specific measurement uncertainty values for the EIRP accuracy, FR2, Normal test conditions | [9 empty] | Comments
Expanded uncertainty (dB)
52.6<f ≤71.0 GHz
Indoor Anechoic Chamber
| Compact Antenna Test Range | 5.4305
One Dimensional Compact Range Chamber
Near Field Test Range
PWS
| Proposed value | 5.4 | [8 empty] | Checked 28/1/20: same values as were in TR 38.817-02.
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
Table 9.2.3.4-2: CATR MU value derivation for EIRP accuracy measurements, Normal test conditions, FR2 | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 52.6<f ≤71.0 GHz |  |  |  |  | 52.6<f ≤71.0 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 |  | Exp. normal | 2 | 1 | 0.1 |  | 0.01 | 0
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 2.36 |  | Gaussian | 1 | 1 | 2.36 |  | 5.5696 | 0
A2-2a | Standing wave between BS and test range antenna | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 |  | 0.02205 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 |  | 0.0001 | 0
A2-4a | QZ ripple experienced by BS | 0.4 |  | Gaussian | 1 | 1 | 0.4 |  | 0.16 | 0
A2-12 | Frequency flatness of test system | 0.25 |  | Gaussian | 1 | 1 | 0.25 |  | 0.0625 | 0
A2-19 | Uncertainty of the LNA (FR2 only) | 1.5 |  | Gaussian | 1 | 1 | 1.5
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.85 |  | Gaussian | 1 | 1 | 0.85 |  | 0.7225 | 0
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 1.3718 |  | U-shaped | 1.4142 | 1 | 0.97 |  | 0.9409 | 0
A2-6 | Insertion loss of receiver chain | 0.18 |  | Rectangular | 1.7321 | 1 | 0.10392 |  | 0.0108 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 |  | 0.0001 | 0
A2-7 | Influence of the calibration antenna feed cable | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 |  | 0.02205 | 0
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 |  | Rectangular | 1.7321 | 1 | 0.30022 |  | 0.090133 | 0
A2-8 | Misalignment positioning system | 0 |  | Exp. normal | 2 | 1 | 0 |  | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 |  | Exp. normal | 2 | 1 | 0 |  | 0 | 0
A2-9 | Rotary joints | 0 |  | U-shaped | 1.4142 | 1 | 0 |  | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 |  | U-shaped | 1.4142 | 1 | 0.06364 |  | 0.00405 | 0
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 |  | Gaussian | 1 | 1 | 0.009 |  | 8.1e-05 | 0
A2-11 | Switching uncertainty | 0.43 |  | Rectangular | 1.7321 | 1 | 0.24826 |  | 0.061633 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 2.7706 | 0 | 2.7706 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 5.4305 | 0 | 5.4305 | 0
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0

## EIRP ex
| Table 9.3.4-2: Test system specific measurement uncertainty values for the EIRP accuracy in Extreme test conditions, FR2 | [9 empty] | Comments
Expanded uncertainty (dB)
52.6<f ≤71.0 GHz
Indoor anechoic
| CATR | 5.9935
Near field
1D CATR
Plane wave sythesiser
| Proposed value | 6 | [8 empty] | Checked 28/1/20: same values as were in TR 38.817-02.
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
Table 9.3.3.4-1: CATR MU value derivation for EIRP accuracy measurements in Extreme test conditions, FR2 | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 52.6<f ≤71.0 GHz |  |  |  |  | 52.6<f ≤71.0 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 |  | Exp. normal | 2 | 1 | 0.1 |  | 0.01 | 0
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 2.36 |  | Gaussian | 1 | 1 | 2.36 |  | 5.5696 | 0
A2-2a | Standing wave between BS and test range antenna | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 |  | 0.02205 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 |  | 0.0001 | 0
A2-13 | QZ ripple with BS (extreme test conditions) | 0.7 |  | Gaussian | 1 | 1 | 0.7 |  | 0.49 | 0
A2-12 | Frequency flatness of test system | 0.25 |  | Gaussian | 1 | 1 | 0.25 |  | 0.0625 | 0
A2-15 | radome loss variation | 0.5 |  | Gaussian | 1 | 1 | 0.5 |  | 0.25 | 0
A2-14 | wet radome loss variation | 0.9 |  | Gaussian | 1 | 1 | 0.9 |  | 0.81 | 0
A2-16 | Change in absorber behaviour | 0.5 |  | Gaussian | 1 | 1 | 0.5 |  | 0.25 | 0
A2-19 | Uncertainty of the LNA (FR2 only) | 1.5 |  | Gaussian | 1 | 1 | 1.5
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.87 |  | Gaussian | 1 | 1 | 0.87 |  | 0.7569 | 0
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 1.3718 |  | U-shaped | 1.4142 | 1 | 0.97 |  | 0.9409 | 0
A2-6 | Insertion loss of receiver chain | 0.18 |  | Rectangular | 1.7321 | 1 | 0.10392 |  | 0.0108 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 |  | 0.0001 | 0
A2-7 | Influence of the calibration antenna feed cable | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 |  | 0.02205 | 0
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 |  | Rectangular | 1.7321 | 1 | 0.30022 |  | 0.090133 | 0
A2-8 | Misalignment positioning system | 0 |  | Exp. normal | 2 | 1 | 0 |  | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 |  | Exp. normal | 2 | 1 | 0 |  | 0 | 0
A2-9 | Rotary joints | 0 |  | U-shaped | 1.4142 | 1 | 0 |  | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 |  | U-shaped | 1.4142 | 1 | 0.06364 |  | 0.00405 | 0
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 |  | Gaussian | 1 | 1 | 0.009 |  | 8.1e-05 | 0
A2-11 | Switching uncertainty | 0.43 |  | Rectangular | 1.7321 | 1 | 0.24826 |  | 0.061633 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 3.0579 | 0 | 3.0579 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 5.9935 | 0 | 5.9935 | 0
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0

## power dynamics
|  | Expanded uncertainty (dB) | [9 empty] | Comments
52.6<f ≤71.0 GHz
Indoor anechoic
CATR
Near field
1D CATR
Plane wave sythesiser
| Proposed value |  | 0.4 |  | same value proposed already | [6 empty] | Same MU values as for FR1.
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!

## EVM
|  | Expanded uncertainty (dB) |  | Total EVM on top of existing 1% EVM uncertainty [%] | [7 empty] | Comments
|  | 24.25<f ≤29.5GHz |  | 52.6<f ≤71.0 GHz
Indoor anechoic
CATR
Near field
1D CATR
Plane wave sythesiser
| Proposed value |  |  | 1.1 |  | Propose new value | [5 empty] | Values for test equipment only (1% EVM, 12Hz Freq error and 25ns TAE).
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
CATR | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0 | #REF!
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0 | #REF!

## TX OFF
| Table 9.10.3-1: Maximum accepted test system uncertainty values for the EIRP accuracy, FR2 | [9 empty] | Comments
Expanded uncertainty (dB)
52.6<f ≤71.0 GHz
| Indoor anechoic | [9 empty] | Checked 28/1/20 pointing error 0.2 in EIRP and 0.3 in TX OFF - use 0.2 for now. NA is 0.3 in EIRP and 0.2 here, use 0.3 cancels out pointing error Receiver missmatch uncertainty different to EIRP: 0.72 rather than 0.43/0.57. Could be becuase of limiter? New error description. Final values matched those as were in TR 38.817-02
| CATR | 5.8647
Near field
1D CATR
Plane wave sythesiser
| Proposed value | 5.9
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
Table 9.10.2.3-1: CATR uncertainty assessment for EIRP measurements for transmitter OFF power and transmitter transient period | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 52.6<f ≤71.0 GHz |  |  |  |  | 52.6<f ≤71.0 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 |  | Exp. normal | 2 | 1 | 0.1 |  | 0.01 | 0
C1-9 | RF power measurement equipment standard uncertainty σ (dB) of the absolute level for a time domain wideband measurement for FR2 | 2.61 |  | Gaussian | 1 | 1 | 2.61 |  | 6.8121 | 0
A2-2a | Standing wave between BS and test range antenna | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 |  | 0.02205 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 |  | 0.0001 | 0
A2-4a | QZ ripple experienced by BS | 0.4 |  | Gaussian | 1 | 1 | 0.4 |  | 0.16 | 0
A2-12 | Frequency flatness of test system | 0.25 |  | Gaussian | 1 | 1 | 0.25 |  | 0.0625 | 0
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.87 |  | Gaussian | 1 | 1 | 0.87 |  | 0.7569 | 0
A2-5b | Mismatch of receiver chain for low power receiver | 1.3718 |  | U-shaped | 1.4142 | 1 | 0.97 |  | 0.9409 | 0
A2-6 | Insertion loss of receiver chain | 0.18 |  | Rectangular | 1.7321 | 1 | 0.10392 |  | 0.0108 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 |  | 0.0001 | 0
A2-7 | Influence of the calibration antenna feed cable | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 |  | 0.02205 | 0
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 |  | Rectangular | 1.7321 | 1 | 0.3 |  | 0.09 | 0
A2-8 | Misalignment positioning system | 0 |  | Exp. normal | 2 | 1 | 0 |  | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 |  | Exp. normal | 2 | 1 | 0 |  | 0 | 0
A2-9 | Rotary joints | 0 |  | U-shaped | 1.4142 | 1 | 0 |  | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 |  | U-shaped | 1.4142 | 1 | 0.06364 |  | 0.00405 | 0
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 |  | Gaussian | 1 | 1 | 0.009 |  | 8.1e-05 | 0
A2-11 | Switching uncertainty | 0.43 |  | Rectangular | 1.7321 | 1 | 0.24826 |  | 0.061633 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 2.9922 | 0 | 2.9922 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 5.8647 | 0 | 5.8647 | 0
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0

## In-band TRP
| Table 11.2.7-2: Test system specific MU values for the OTA BS output power test, Normal test conditions, FR2 | [9 empty] | Comments
Expanded uncertainty (dB)
52.6<f ≤71.0 GHz
| Indoor anechoic | [9 empty] | Checked 28/1/20: CATR -pointing error different: 0.2 for EIRP, 0.35 here -RX chain missmatch calc is worng (corrected) -Misalignment of positioning system, has calc error (std uncertainty is zero though which matches EIRP) -Misalignment of SGH and pointing system is zero in TRP/TX OFF but is 0.5 here - why? -Switching uncertainty different (uses EIRP uncertainty as std uncertainty) -With all above corrected get 1.74/2.07 compared to 1.78/2.06 - so total is ok Reverb: NA uncertainty shoudl be 0.3 not 0.2, small incraes in total but under agreed value
| CATR | 5.5402
Near field
1D CATR
Plane wave sythesiser
reverb
| Proposed value | 5.5
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
Table 11.2.3.4-1: CATR MU value derivation for OTA BS output power measurement, FR2 | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 52.6<f ≤71.0 GHz |  |  |  |  | 52.6<f ≤71.0 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 |  | Exp. normal | 2 | 1 | 0.1 | 0 | 0.01 | 0
C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - high power (EIRP, TRP) | 2.36 |  | Gaussian | 1 | 1 | 2.36 | 0 | 5.5696 | 0
A2-2a | Standing wave between BS and test range antenna | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 | 0 | 0.02205 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 | 0 | 0.0001 | 0
A2-4a | QZ ripple experienced by BS | 0.4 |  | Gaussian | 1 | 1 | 0.4 | 0 | 0.16 | 0
A2-12 | Frequency flatness of test system | 0.25 |  | Gaussian | 1 | 1 | 0.25 | 0 | 0.0625 | 0
A2-19 | Uncertainty of the LNA (FR2 only) | 1.5 |  | Gaussian | 1 | 1 | 1.5
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.85 |  | Gaussian | 1 | 1 | 0.85 | 0 | 0.7225 | 0
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 1.3718 |  | U-shaped | 1.4142 | 1 | 0.97 | 0 | 0.9409 | 0
A2-6 | Insertion loss of receiver chain | 0.18 |  | Rectangular | 1.7321 | 1 | 0.10392 | 0 | 0.0108 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 | 0 | 0.0001 | 0
A2-7 | Influence of the calibration antenna feed cable | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 | 0 | 0.02205 | 0
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 |  | Rectangular | 1.7321 | 1 | 0.30022 | 0 | 0.090133 | 0
A2-8 | Misalignment positioning system | 0 |  | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 |  | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 |  | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 |  | U-shaped | 1.4142 | 1 | 0.06364 | 0 | 0.00405 | 0
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 |  | Gaussian | 1 | 1 | 0.009 | 0 | 8.1e-05 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 2.7595 | 0 | 2.7595 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 5.4086 | 0 | 5.4086 | 0
TRP Summation error | [6 empty] | 1.2 | 1.2
Total MU | [6 empty] | 5.5402 | 1.2
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
Table 11.2.5.4-1: Reverberation chamber MU value derivation for OTA BS output power, FR2 | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Stage 1: Calibration measurement | [8 empty] | 0 | 0
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | #REF! | #REF! | #REF! | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | #REF! | #REF! | #REF! | #REF!

## ACLR-abs
| Table 11.3.7-3: Test system specific MU values for the absolute OTA ACLR, Normal test conditions, FR2 | [9 empty] | Comments
Expanded uncertainty (dB)
52.6<f ≤71.0 GHz
| Indoor anechoic | [9 empty] | Checked 28/1/20 CATR -pointing error different: 0.2 for EIRP, 0.3 here - missalignment values different to EIRP (where they are 0) - with EIRP values used the final value is 2.41/2.43 rather than 2.37/2.37 - propose using EIRP values Reverb - NA is wrong (0.2 not 0.3dB) - total very similar as was in TR38.817-02 (2.32) smaller than agreed value so OK
| CATR | 5.5733
Near field
1D CATR
Plane wave sythesiser
Reverberation Chamber
| Proposed value | 5.6
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Table 11.3.3.4-1: CATR MU value derivation for the EIRP measurement of the absolute OTA ACLR, FR2 | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 52.6<f ≤71.0 GHz |  |  |  |  | 52.6<f ≤71.0 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 |  | Exp. normal | 2 | 1 | 0.1 | 0 | 0.01 | 0
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 2.36 |  | Gaussian | 1 | 1 | 2.36 | 0 | 5.5696 | 0
A2-2a | Standing wave between BS and test range antenna | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 | 0 | 0.02205 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 | 0 | 0.0001 | 0
A2-4a | QZ ripple experienced by BS | 0.4 |  | Gaussian | 1 | 1 | 0.4 | 0 | 0.16 | 0
A2-12 | Frequency flatness of test system | 0.25 |  | Gaussian | 1 | 1 | 0.25 | 0 | 0.0625 | 0
A2-19 | Uncertainty of the LNA (FR2 only) | 1.5 |  | Gaussian | 1 | 1 | 1.5
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.87 |  | Gaussian | 1 | 1 | 0.87 | 0 | 0.7569 | 0
A2-5b | Mismatch of receiver chain for low power receiver | 1.3718 |  | U-shaped | 1.4142 | 1 | 0.97 | 0 | 0.9409 | 0
A2-6 | Insertion loss of receiver chain | 0.18 |  | Rectangular | 1.7321 | 1 | 0.10392 | 0 | 0.0108 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 | 0 | 0.0001 | 0
A2-7 | Influence of the calibration antenna feed cable | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 | 0 | 0.02205 | 0
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 |  | Rectangular | 1.7321 | 1 | 0.3 | 0 | 0.09 | 0
A2-8 | Misalignment positioning system | 0 |  | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 |  | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 |  | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 |  | U-shaped | 1.4142 | 1 | 0.06364 | 0 | 0.00405 | 0
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 |  | Gaussian | 1 | 1 | 0.009 | 0 | 8.1e-05 | 0
A2-11 | Switching uncertainty | 0.43 |  | Rectangular | 1.7321 | 1 | 0.24826 | 0 | 0.061633 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 2.7768 | 0 | 2.7768 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 5.4426 | 0 | 5.4426 | 0
TRP Summation error | [6 empty] | 1.2 | 1.2 |  |  | RK: corrected after Nokia comment
Total MU | [6 empty] | 5.5733 | 1.2
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Table 11.3.5.4-1: Reverberation chamber MU value derivation for absolute ACLR measurement, FR2 | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 2.36 | 0 | Gaussian | 1 | 1 | 2.36 | 0 | 5.5696 | 0
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Stage 1: Calibration measurement
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
C1-3 | Uncertainty of the network analyzer | 0.87 | 0 | Gaussian | 1 | 1 | 0.87 | 0 | 0.7569 | 0
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | #REF! | #REF! | #REF! | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | #REF! | #REF! | #REF! | #REF!

## ACLR-rel
| Table 11.3.7-4: Test system specific MU values for relative OTA ACLR, Normal conditions, FR2 | [9 empty] | Comments
Expanded uncertainty (dB)
52.6<f ≤71.0 GHz
| Indoor anechoic | [9 empty] | Checked 28/1/20 CATR -pointing error is different again 0.2 for EIRP, 0.3 here - reciever missmatch different to EIRP -reciever insertion loss variation different (EIRP=0, here, 0.18/0.2) - Using EIRP value, final value is close 1.94/2.23 vs. 1.96/2.31 propose using EIRP values Reverb - NA uncertainty shoudl be 0.3 not 0.2, small incraes in total but under agreed value
| CATR | 5.5134
Near field
1D CATR
Plane wave sythesiser
Reverberation Chamber
| Proposed value | 5.5
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 0 | 0 | 0 | 0
TRP Summation error | [6 empty] | 0.75 | 0.75
Total MU | [6 empty] | 0.75 | 0.75
Table 11.3.3.4-2: CATR MU value derivation for the EIRP measurement of the relative OTA ACLR, FR2 | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 52.6<f ≤71.0 GHz |  |  |  |  | 52.6<f ≤71.0 GHz
Stage 2: BS measurement
A2-1a | Misalignment and pointing error of BS (for EIRP) | 0.2 |  | Exp. normal | 2 | 1 | 0.1 | 0 | 0.01 | 0
C1-8 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - relative (ACLR) | 2.36 |  | Gaussian | 1 | 1 | 2.36 | 0 | 5.5696 | 0
A2-4a | QZ ripple experienced by BS | 0.4 |  | Gaussian | 1 | 1 | 0.4 | 0 | 0.16 | 0
A2-12 | Frequency flatness of test system | 0.25 |  | Gaussian | 1 | 1 | 0.25 | 0 | 0.0625 | 0
A2-19 | Uncertainty of the LNA (FR2 only) | 1.5 |  | Gaussian | 1 | 1 | 1.5
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.85 |  | Gaussian | 1 | 1 | 0.85 | 0 | 0.7225 | 0
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 1.3718 |  | U-shaped | 1.4142 | 1 | 0.97 | 0 | 0.9409 | 0
A2-6 | Insertion loss of receiver chain | 0.18 |  | Rectangular | 1.7321 | 1 | 0.10392 | 0 | 0.0108 | 0
A2-11 | Switching uncertainty | 0.43 |  | Rectangular | 1.7321 | 0 | 0.24826 | 0 | 0.061633 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 2.7455 | 0 | 2.7455 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 5.3812 | 0 | 5.3812 | 0
TRP Summation error | [6 empty] | 1.2 | 1.2 |  |  | [rk] corrected
Total MU | [6 empty] | 5.5134 | 1.2
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Table 11.3.5.4-2: Reverberation chamber MU value derivation for relative ACLR measurement, FR2 | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
C1-8 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - relative (ACLR) | 2.36 | 0 | Gaussian | 1 | 1 | 2.36 | 0 | 5.5696 | 0
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Stage 1: Calibration measurement | [8 empty] | 0 | 0
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
C1-3 | Uncertainty of the network analyzer | 0.87 | 0 | Gaussian | 1 | 1 | 0.87 | 0 | 0.7569 | 0
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | #REF! | #REF! | #REF! | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | #REF! | #REF! | #REF! | #REF!

## OBUE
| Table 11.4.7-2: Test system specific MU values for the OTA OBUE measurement, FR2 | [9 empty] | Comments
Expanded uncertainty (dB)
52.6<f ≤71.0 GHz
| Indoor anechoic | [9 empty] | Checked 29/1/20 CATR - NA is wrong -Influence of calibration feed cable is different to ACS abs/EIRP -numbetr of other minor differences - corrected total 2.53 vs 2.37 below agreed value so ok Reverb - NA is wrong (0.2 not 0,.3) total very similar to TR38.817-02 (2.32) smaller than agreed value so OK
| CATR | 5.5776
Near field
1D CATR
Plane wave sythesiser
Reverberation Chamber
| Proposed value | 5.6
Indoor anechoic | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Table 11.4.3.4-1: CATR MU value derivation for OTA OBUE measurement, FR2 | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 52.6<f ≤71.0 GHz |  |  |  |  | 52.6<f ≤71.0 GHz
Stage 2: BS measurement
A2-18a | Misalignment and pointing error of BS (for TRP) | 0.3 |  | Exp. normal | 2 | 1 | 0.15 | 0 | 0.0225 | 0
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 2.36 |  | Gaussian | 1 | 1 | 2.36 | 0 | 5.5696 | 0
A2-2a | Standing wave between BS and test range antenna | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 | 0 | 0.02205 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 | 0 | 0.0001 | 0
A2-4a | QZ ripple experienced by BS | 0.4 |  | Gaussian | 1 | 1 | 0.4 | 0 | 0.16 | 0
A2-12 | Frequency flatness of test system | 0.25 |  | Gaussian | 1 | 1 | 0.25 | 0 | 0.0625 | 0
A2-19 | Uncertainty of the LNA (FR2 only) | 1.5 |  | Gaussian | 1 | 1 | 1.5
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.87 |  | Gaussian | 1 | 1 | 0.87 | 0 | 0.7569 | 0
A2-5b | Mismatch of receiver chain for low power receiver | 1.3718 |  | U-shaped | 1.4142 | 1 | 0.97 | 0 | 0.9409 | 0
A2-6 | Insertion loss of receiver chain | 0.18 |  | Rectangular | 1.7321 | 1 | 0.10392 | 0 | 0.0108 | 0
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 |  | Gaussian | 1 | 1 | 0.01 | 0 | 0.0001 | 0
A2-7 | Influence of the calibration antenna feed cable | 0.21 |  | U-shaped | 1.4142 | 1 | 0.14849 | 0 | 0.02205 | 0
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.51962 |  | Rectangular | 1.7321 | 1 | 0.3 | 0 | 0.09 | 0
A2-8 | Misalignment positioning system | 0 |  | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 |  | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 |  | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 |  | U-shaped | 1.4142 | 1 | 0.06364 | 0 | 0.00405 | 0
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 |  | Gaussian | 1 | 1 | 0.009 | 0 | 8.1e-05 | 0
A2-11 | Switching uncertainty | 0.43 |  | Rectangular | 1.7321 | 1 | 0.24826 | 0 | 0.061633 | 0
Combined standard uncertainty (1σ) (dB) | [6 empty] | 2.7791 | 0 | 2.7791 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 5.447 | 0 | 5.447 | 0
TRP summation error | [6 empty] | 1.2 | 1.2 |  |  | [rk] corrected
Total MU | [6 empty] | 5.5776 | 1.2
Near field | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Combined standard uncertainty (1σ) (dB) | [6 empty] | 0 | 0
One Dimensional Compact Range Chamber | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Plane wave synthesiser | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [8 empty] | 0 | 0
Table 11.4.5.4-1: Reverberation chamber MU value derivation for OTA OBUE measurement, FR2 | [8 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 24.25<f ≤29.5GHz | 37<f ≤40GHz |  |  |  | 24.25<f ≤29.5GHz | 37<f ≤40GHz
Stage 2: BS measurement
C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 2.36 | 0 | Gaussian | 1 | 1 | 2.36 | 0 | 5.5696 | 0
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1 | #REF! | #REF! | #REF! | #REF!
Stage 1: Calibration measurement
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
C1-3 | Uncertainty of the network analyzer | 0.87 | 0 | Gaussian | 1 | 0.87 | 0.87 | 0 | 0.7569 | 0
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Combined standard uncertainty (1σ) (dB) | [6 empty] | #REF! | #REF! | #REF! | #REF!
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | #REF! | #REF! | #REF! | #REF!

## OOB EM
| Table 12.2.5-2: Spurious emissions MU values – FR2 | [9 empty] | Comments
Expanded uncertainty (dB)
|  |  |  | 71GHz<f ≤110GHz | 110GHz<f ≤142GHz
| Indoor Anechoic Chamber |  |  | x
| CATR |  |  | 6.899 | 6.9696
| Reverberation Chamber |  |  | x
| Proposed value |  |  | 6.9 | 7
Table 12.2.3.3-1: CATR value derivation for TX spurious emissions, FR2
UID | Uncertainty source |  |  | Distribution of the probability | Divisor based on distribution shape | ci
|  | 71GHz<f ≤110GHz | 110GHz<f ≤142GHz |  |  |  | 71GHz<f ≤110GHz | 110GHz<f ≤142GHz
Stage 2: BS measurement
A2-18a | Misalignment and pointing error of BS (for TRP) | 0.3 | 0.3 | Exp. normal | 2 | 1 | 0.15 | 0.15 | 0.0225 | 0.0225
C1-7 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) - low power (UEM, absolute ACLR) | 2.39 | 2.49 | Gaussian | 1 | 1 | 2.39 | 2.49 | 5.7121 | 6.2001
A2-19 | Uncertainty of the LNA (FR2 only) | 1.5 | 1.5 | Gaussian | 1 | 1 | 1.5 | 1.5 | 2.25 | 2.25
A2-20 | Uncertainty of the mixer (FR2 only) | 2.25 | 2.35 | Gaussian | 1 | 1 | 2.25 | 2.35 | 5.0625 | 5.5225 | Removed from RSS summing because already included in power measurement
A2-2a | Standing wave between BS and test range antenna | 0.21 | 0.21 | Gaussian | 1 | 1 | 0.21 | 0.21 | 0.0441 | 0.0441
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-4a | QZ ripple experienced by BS | 0.4 | 0.4 | Gaussian | 1 | 1 | 0.4 | 0.4 | 0.16 | 0.16
A2-12 | Frequency flatness of test system | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.0625 | 0.0625
A2-10 | Miscellaneous uncertainty | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0
Stage 1: Calibration measurement | [8 empty] | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.85 | 0.85 | Gaussian | 1 | 1 | 0.85 | 0.85 | 0.7225 | 0.7225
A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver | 1.6829 | 1.6122 | U-shaped | 1.4142 | 1 | 1.19 | 1.14 | 1.4161 | 1.2996
A2-5b | Mismatch of receiver chain for low power receiver | 1.6829 | 1.6122 | U-shaped | 1.4142 | 1 | 1.19 | 1.14 | 1.4161 | 1.2996
A2-6 | Insertion loss of receiver chain | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.0108 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.0001 | 0.0001
A2-7 | Influence of the calibration antenna feed cable | 0.29 | 0.29 | U-shaped | 1.4142 | 1 | 0.20506 | 0.20506 | 0.04205 | 0.04205
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.52 | 0.52 | Rectangular | 1.7321 | 1 | 0.30022 | 0.30022 | 0.090133 | 0.090133
A2-8 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-1b | Misalignment and pointing error of calibration antenna (for EIRP) | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
A2-9 | Rotary joints | 0 | 0 | U-shaped | 1.4142 | 1 | 0 | 0 | 0 | 0
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.00405 | 0.00405
A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.24826 | 0.24826 | 0.061633 | 0.061633
Combined standard uncertainty (1σ) (dB) | [6 empty] | 3.4662 | 3.5028 | 3.4662 | 3.5028
Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) | [6 empty] | 6.7938 | 6.8656 | 6.7938 | 6.8656
TRP Summation error | [6 empty] | 1.2 | 1.2
Total MU | [6 empty] | 6.899 | 6.9696

## RX EM
|  | Expanded uncertainty (dB) | [11 empty] | Comments
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
| Indoor anechoic chamber | x | x | x
| CATR | x | x | x
| Reverberation Chamber | x | x | x
| Agreed value | x | x | 5 | [9 empty] | Checked 29/1/20: same as TX emssisons
Indoor anechoic chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz |  |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz |  |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Reverberation Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB)
|  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz |  |  |  | 18GHz<f≤26.5 GHz | 26.5GHz<f ≤40GHz | 40GHz<f ≤60GHz
Stage 2: BS measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0