3GPP TSG-RAN WG4 Meeting #114 	R4-2502320

Athens, Greece, Feb 17 – 21, 2025

Source: 	Huawei, HiSilicon

Title: 	TP to TR 38.908 – section 6.5 and 6.6

Agenda Item:	7.13.4.1

Document for:	Approval

1 Introduction

In RAN4#113 the TP to TR 38.908 on section 6.5 and 6.6 [1] was agreed and updated in the latest version of the TR [2].

This paper discusses those issues and suggests text to be captured in the TR

TP to TR 38.908

## 6.5	Method of test

### 6.5.1	General

Traditionally, the measurement can be divided into two different stages: Test range calibration and measurement. For an  EEIRP measurement a third stage is introduced.

      - Test range pathloss calibration and measurement receiver absolute power level calibration.

      - EIRP pattern measurements for a set of defined test beam directions.

      - Post-processing to calculate EEIRP for defined elevation bins.

-π≤φ≤π0≤θ≤π2

EIRPkθn,φmθn,φmm=1..Mn=1..N

K

EIRPθn,φm=1Kk=1KEIRPkθn,φm

EIRPkθn,θm

EEIRPi=θHi-θLisinθHi-sinθLiMN∙n=nLinHim=1MEIRPθn,φmcosθn

θLiθHinLinHiθLi,θHiN= nHi-nLi

In the calibration stage the antenna test range transmission loss is characterized with the intention to minimize error sources related to test setup. A common approach used is the substitution method, where a reference antenna with known antenna gain is used to determine the transmission loss in the test range. The test range calibration is described in detail in TR 37.941, subclause 8.3.

In addition, the measurement uncertainty can be improved by calibrating the absolute power measurement uncertainty of the measurement receiver (spectrum analyzer with fairly high absolute power measurement uncertainty) using a power meter (with very low measurement uncertainty). A detailed description of suitable test ranges and corresponding calibration procedures can be found in TR 37.941, subclause 8.3.1.

### 6.5.2	Initial conditions

The initial conditions section contains the test set up initial conditions including:

- the test environment,

- the channels to be tested

- the beams to be tested or the test beam direction set in this case

- Mechanical tilt and steering range conditions for test

Test environment

EEIRP is to be measured under normal conditions

NOTE: Currently, only a few parameters are measured under extreme conditions which is very challenging as the OTA chamber cannot be environmentally controlled. Controlling the temperature of the DUT locally is implemented for limited tests, but this is not practical to EEIRP measurements.

RF channels

EEIRP measurement is a very long measurement, the number of channels tested will increase the test time.

The frequency variation across the EEIRP band is approximately 10%, over this range, the gain of the main beam will slightly change. However, EEIRP estimated value is dominated by the sidelobe behaviour of the antenna pattern rather than the main beam. As such, looking at the EEIRP over the frequency range there is no clear worst case, the performance is no worse at Bottom channel or Top channel than it is at the Middle channel.

As no RF channel is worse than any other it could be argued any channel could be used. Middle channel is the obvious choice.

Test signal and test model

As EEIRP requirement is based on power spectral density (PSD), testing should be carried out with the highest PSD the BS is capable to support.

NOTE: Maximum PSD might be achieved when using the smallest channel bandwidth, despite limited output power.

The performance is dependent on the radiation beam shape, non-linearities or distortions potentially generated by multi-carrier signals will not significantly alter the beam shape of the wanted signal.

AFor this reason single carrier test signal is sufficient for the scope of conformance testing, together with test model NR-FR1-TM1.1.

Beams to be tested.

The test beam direction set is discussed in sub-clause 6.3. The set of beams used for the test should be described or referenced in the initial conditions.

Mechanical tilt and steering range conditions for test

The test beam direction set is applied to the BS steering range and applicable mechanical tilt conditions.

The BS should be compliant to the EEIRP requirement over all operational conditions however it can be seen from figure 6.5.2-1 that the worst case conditions with respect to steering range and mechanical tilt are quite clear.

Figure 6.5.2-1. left: Margin vs mechanical down tilt, right margin vs electrical steering range

In figure 6.5.2-1 on the left hand graph the mechanical tilt is varied for various electrical steering ranges. In all cases it can be seen that as mechanical tilt increases the margin gets larger.

On the right hand graph the electrical steering range is varied (for a fixed mechanical tilt). It is clear that as the steering range increases the margin reduces

The worst-case condition is when mechanical tilt is minimum (closest to horizon) and the electrical steering range is maximum. As such this condition should be declared and tested.

In real products we should allow for that fact that these 2 conditions may not be declared simultaneously as such the following caveat should be added.

The minimum mechanical down tilt and the maximum electrical steering range shall be declared. If the maximum steering range and the minimum mechanical down tilt are not simultaneously supported, the manufacturer shall declare and test the following two instances:

  - Minimum mechanical down tilt and the reduced electrical steering range

  - Maximum electrical steering range and increased mechanical down tilt

The mechanical tilts and associated electrical steering ranges are declared.

### 6.5.3	Test procedure

The test procedure is based on a CATR/IAC where the BS is placed on a positioner and rotated so the required measurement direction is in line with the measurement antenna. The procedure has similarities to existing EIRP and TRP measurement procedures.

The test procedure is as follows:

1. Calibrate the test range, according to calibration described in TR 37.941, subclause 8.3

2. Place the BS at the positioner, such that blocking effect in the vertical domain is minimized.

3. Align the manufacturer declared coordinate system orientation (D.2) of the test object with the coordinate system used by the test system.

4. The measurement device characteristics shall be:

-	Detection mode: True RMS.

5. Set the BS to transmit on both plarizations according to applicable test configuration in TS 38.141-2, subclause 4.8 using the corresponding test model NR-FR1-TM1.1 described in TS 38.141-1, subclause 4.9.2.

6. Orient the positioner (and the BS) 	to the angle for measurement point θn,φm for m=1..M and n=1.. N. Spatial sampling grid is defined in Annex N.3 in TS 38.141-2.

7. Configure test beam(s) and test equipment for measuring test beam k for k=1…K, as defined in clause 4.9.3.2kk=1..K.

8. Measure EIRPkθn,φm=EIRPp1+EIRPp2 where p1 and p2 denotes two orthogonal polarizations.

9. Repeat step 6 and 7 for all K test beams.

10. Repeat step 5, 6, 7 and 8 for all M and N positioner locations.

11. Calculate EEIRPi profile from measured EIRPkθn,φm patterns for bins i=1..7 as described in subclause 6.5.

NOTE: step 9 and step 10 could be exchangeable.

## 6.6	Measurement uncertainty

The EEIRP measurement is very similar to the in-band TRP measurements, as such the in-band emission MU budgets from TR 37.941 can be used as a starting point for the EEIRP budgets.

Additional sources of uncertainty are:

Dynamic range uncertainty

One major difference with EEIRP, compared to TRP, is the necessity to measure complete EIRP pattern with focus on the sidelobe region (i.e. for a real-life site this is the region pointing towards the sky). This means measurement receiver must be considered in such way that the dynamic range can handle the power created by the peak main beam as the maximum level and detecting the low levels in the side lobe region with accurate power. The power level accuracy is typically different for high power with large SNR compared to coming closer to the noise floor of the measurement receiver.In addition, to avoid the impacts (Reflections and refractions of the main beam within the chamber) from main beam to sidelobe beam, the chamber spatial isolation is necessary to remove the effect of the main beam.

The lowest EEIRP requirement is 15dBm/MHz

Currently the dynamic range contribution is 0.51dB, in order to maintain that level of accuracy in the above scenario an isolation of approximately 58dB under the assumption of peak EIRP 63.7dBm for main beam would be required:

Pint=10*log101063.7-5810+101510=15.5dBm/MHz i.e. 0.5dB error

EEIRP sampling Error

The EEIRP sampling error is caused by the fact that coverage steering range is sampled with a limited set of test beams.

The measure EEIRP profile using different granularity of test beams distributed within the setting range will proceed error within EEIRP bin.

According to the agreed test beam pattern, errors between the EEIRP simulation results of test beam and EEIRP simulation results of reference testing beams (e..g 2.4K or 8K or 10K) are provided in the following table.

Table 6.6.-1.Errors per bin between the EEIRP simulation results of test beam and EEIRP simulation results of reference testing beams

| ZTE |  | Bin index | 1 | 2 | 3 | 4 | 5 | 6 | 7 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Ant 4, 12degree | Error for Rect Ref | -0.98 | -0.17 | -0.70 | -0.28 | -0.30 | 0.04 | 0.36 |
|  |  | Error for Ellipse Ref | 0.25 | -0.22 | -0.03 | 0.07 | 0.21 | -0.11 | -0.40 |
|  | Ant 4, 16degree | Error for Rect Ref | -1.51 | -1.38 | -0.78 | 0.64 | -0.42 | -0.20 | 0.54 |
|  |  | Error for Ellipse Ref | -0.09 | -0.97 | 0.38 | -0.03 | 0.20 | -0.05 | -0.54 |
|  | Ant 1, 12degree | Error for Rect Ref | -0.98 | -0.17 | -0.70 | -0.28 | -0.30 | 0.04 | 0.36 |
|  |  | Error for Ellipse Ref | 0.25 | -0.22 | -0.03 | 0.07 | 0.21 | -0.11 | -0.40 |
|  | Ant 1, 16degree | Error for Rect Ref | -1.13 | -0.09 | 0.03 | 0.09 | 0.14 | 0.33 | 0.72 |
|  |  | Error for Ellipse Ref | 0.14 | 0.10 | 0.03 | 0.02 | -0.01 | -0.06 | -0.12 |
| Huawei |  | Bin index | 1 | 2 | 3 | 4 | 5 | 6 | 7 |
|  | Ant 4, 12degree | Error for Rect Ref | -1.25 | -0.32 | -1.15 | -0.81 | -0.46 | -0.26 | 0.36 |
|  |  | Error for Ellipse Ref | 0.28 | 0.07 | 0.01 | 0.07 | 0.17 | 0.10 | -0.40 |
|  | Ant 4, 16degree | Error for Rect Ref | -0.44 | -1.12 | -0.82 | 0.38 | -0.77 | -0.34 | 0.37 |
|  |  | Error for Ellipse Ref | 0.43 | 0.27 | -0.06 | -0.01 | 0.10 | 0.03 | -0.48 |
|  | Ant 1, 12degree | Error for Rect Ref | -1.08 | -0.18 | -0.05 | 0.01 | 0.05 | 0.28 | 0.66 |
|  |  | Error for Ellipse Ref | 0.14 | 0.01 | 0 | 0 | 0.02 | -0.03 | -0.15 |
|  | Ant 1, 16degree | Error for Rect Ref | -1.15 | -0.34 | -0.17 | -0.09 | -0.01 | 0.25 | 0.60 |
|  |  | Error for Ellipse Ref | -0.01 | 0.01 | 0.03 | 0.04 | 0.03 | 0 | -0.18 |
| Ericsson |  | Bin index | 1 | 2 | 3 | 4 | 5 | 6 | 7 |
|  | Ant 4, 16degree | Error for Rect Ref | -1,35 | -0,43 | 1,03 | 0,7 | 0,82 | 0,49 | 1,96 |
|  |  | Error for Ellipse Ref | 0,48 | -0,19 | 0,2 | 0,57 | 0,11 | 0,15 | 1,07 |
|  | Ant 1, 16degree | Error for Rect Ref | -0,66 | -0,311 | 0,04 | 0,05 | 0,09 | 0,23 | 0,63 |
|  |  | Error for Ellipse Ref | 0,41 | -0,3 | 0,1 | 0,17 | 0,06 | 0,02 | 0,06 |
| Nokia |  | Bin index | 1 | 2 | 3 | 4 | 5 | 6 | 7 |
|  | Ant 1, 12degree | Error for Rect Ref | -1.47 | -0.12 | -0.33 | -0.15 | 0.06 | 0.28 | 0.90 |
|  |  | Error for Ellipse Ref | 1.16 | -0.10 | -0.26 | -0.12 | 0.04 | 0.13 | 0.62 |

In the following section, the probability distribution curve for error across different antenna configuration and steering range is provided to check the distribution function. Meanwhile the probability distribution curve for error among all 7 bins across different antenna configuration and steering range is provided to check the distribution function.

These two errors have been added to the MU budget as follows:

| Table 6.6-1: IAC MU value derivation for EEIRP measurement, FR1 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB) |  |  |  |  |  |
|  |  |  |  | 4.2<f≤7.125 GHz |  |  |  |  |  |  |  | 4.2<f≤7.125 GHz |  |
| Stage 2: BS measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A1-1 | Positioning misalignment between the BS and the reference antenna |  |  | 0.03 | Rectangular | 1.73 | 1 |  |  |  |  | 0.02 |  |
| A1-2 | Pointing misalignment between the BS and the receiving antenna |  |  | 0.3 | Rectangular | 1.73 | 1 |  |  |  |  | 0.17 |  |
| A1-3 | Quality of quiet zone |  |  | 0.1 | Gaussian | 1.00 | 1 |  |  |  |  | 0.10 |  |
| A1-4a | Polarization mismatch between the BS and the receiving antenna |  |  | 0.01 | Rectangular | 1.73 | 1 |  |  |  |  | 0.01 |  |
| A1-5a | Mutual coupling between the BS and the receiving antenna |  |  | 0 | Rectangular | 1.73 | 1 |  |  |  |  | 0.00 |  |
| A1-6 | Phase curvature |  |  | 0.05 | Gaussian | 1.00 | 1 |  |  |  |  | 0.05 |  |
| C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) |  |  | 0.26 | Gaussian | 1.00 | 1 |  |  |  |  | 0.26 |  |
| A1-7 | Impedance mismatch in the receiving chain |  |  | 0.33 | U-shaped | 1.41 | 1 |  |  |  |  | 0.23 |  |
| A1-8 | Random uncertainty |  |  | 0.1 | Rectangular | 1.73 | 1 |  |  |  |  | 0.06 |  |
| A2-21 | Measurement system dynamic range uncertainty |  |  | [0.50] | Rectangular | 1.73 | 1 |  |  |  |  | [0.30] |  |
| *new | [EEIRP sampling error] |  |  | TBD | [TBD] | 1 | 1 |  |  |  |  | TBD |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A1-9 | Impedance mismatch between the receiving antenna and the network analyzer |  |  | 0.05 | U-shaped | 1.41 | 1 |  |  |  |  | 0.04 |  |
| A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna |  |  | 0.01 | Rectangular | 1.73 | 1 |  |  |  |  | 0.01 |  |
| A1-11 | Impedance mismatch between the reference antenna and the network analyzer |  |  | 0.05 | U-shaped | 1.41 | 1 |  |  |  |  | 0.04 |  |
| A1-3 | Quality of quiet zone (normal test conditions) |  |  | 0.10 | Gaussian | 1.00 | 1 |  |  |  |  | 0.10 |  |
| A1-4b | Polarization mismatch between the reference antenna and the receiving antenna |  |  | 0.01 | Rectangular | 1.73 | 1 |  |  |  |  | 0.01 |  |
| A1-5b | Mutual coupling between the reference antenna and the receiving antenna |  |  | 0.00 | Rectangular | 1.73 | 1 |  |  |  |  | 0.00 |  |
| A1-6 | Phase curvature |  |  | 0.05 | Gaussian | 1.00 | 1 |  |  |  |  | 0.05 |  |
| C1-3 | Uncertainty of the network analyzer |  |  | [0.20 | Gaussian | 1.00 | 1 |  |  |  |  | [0.20] |  |
| A1-12 | Influence of the reference antenna feed cable |  |  | 0.05 | Rectangular | 1.73 | 1 |  |  |  |  | 0.03 |  |
| A1-13 | Reference antenna feed cable loss measurement uncertainty |  |  | 0.06 | Gaussian | 1.00 | 1 |  |  |  |  | 0.06 |  |
| A1-14 | Influence of the receiving antenna feed cable |  |  | 0.05 | Rectangular | 1.73 | 1 |  |  |  |  | 0.03 |  |
| C1-4 | Uncertainty of the absolute gain of the reference antenna |  |  | 0.43 | Rectangular | 1.73 | 1 |  |  |  |  | 0.25 |  |
| A1-15 | Uncertainty of the absolute gain of the receiving antenna |  |  | 0.00 | Rectangular | 1.73 | 1 |  |  |  |  | 0.00 |  |
| Combined standard uncertainty (1σ) (dB) |  |  |  |  |  |  |  |  |  |  |  |  | TBD |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) |  |  |  |  |  |  |  |  |  |  |  |  | TBD |
| Table 6.6-2: CATR MU value derivation for EEIRP measurement, FR1 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB) |  |  |  |  |  |
|  |  |  |  | 4.2<f≤7.125 GHz |  |  |  |  |  |  |  | 4.2<f≤7.125 GHz |  |
| Stage 2: BS measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A2-18a | Misalignment and pointing error of BS (for TRP) |  |  | 0.30 | Rectangular | 1.73 | 1 |  |  |  |  | 0.173 |  |
| C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) |  |  | 0.26 | Gaussian | 1.00 | 1 |  |  |  |  | 0.260 |  |
| A2-2a | Standing wave between BS and test range antenna |  |  | 0.21 | U-shaped | 1.41 | 1 |  |  |  |  | 0.148 |  |
| A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) |  |  | 0.00 | Gaussian | 1.00 | 1 |  |  |  |  | 0.001 |  |
| A2-4a | QZ ripple experienced by BS |  |  | 0.09 | Gaussian | 1.00 | 1 |  |  |  |  | 0.093 |  |
| A2-12 | Frequency flatness of test system |  |  | 0.25 | Gaussian | 1.00 | 1 |  |  |  |  | 0.250 |  |
| A2-17 | Measurement system dynamic range uncertainty |  |  | [0.50] | Rectangular | 1.73 | 1 |  |  |  |  | [0.30] |  |
| *new | [EEIRP sampling error] |  |  | TBD | [TBD] | 1 | 1 |  |  |  |  | TBD |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Uncertainty of the network analyzer |  |  | 0.20 | Gaussian | 1.00 | 1 |  |  |  |  | 0.20 |  |
| A2-5a | Mismatch of receiver chain between receiving antenna and measurement receiver |  |  | 0.33 | U-shaped | 1.41 | 1 |  |  |  |  | 0.23 |  |
| A2-6 | Insertion loss of receiver chain |  |  | 0.18 | Rectangular | 1.73 | 1 |  |  |  |  | 0.10 |  |
| A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) |  |  | 0.00 | Gaussian | 1.00 | 1 |  |  |  |  | 0.00 |  |
| A2-7 | Influence of the calibration antenna feed cable |  |  | 0.02 | U-shaped | 1.41 | 1 |  |  |  |  | 0.02 |  |
| C1-4 | Uncertainty of the absolute gain of the reference antenna |  |  | 0.43 | Rectangular | 1.73 | 1 |  |  |  |  | 0.25 |  |
| A2-8 | Misalignment positioning system |  |  | 0.00 | Exp. normal | 2.00 | 1 |  |  |  |  | 0.00 |  |
| A2-18b | Misalignment and pointing error of calibration antenna (for TRP) |  |  | 0.50 | Exp. normal | 2.00 | 1 |  |  |  |  | 0.25 |  |
| A2-9 | Rotary joints |  |  | 0.05 | U-shaped | 1.41 | 1 |  |  |  |  | 0.03 |  |
| A2-2b | Standing wave between calibration antenna and test range antenna |  |  | 0.09 | U-shaped | 1.41 | 1 |  |  |  |  | 0.06 |  |
| A2-4b | QZ ripple experienced by calibration antenna (normal test conditions) |  |  | 0.01 | Gaussian | 1.00 | 1 |  |  |  |  | 0.01 |  |
| A2-11 | Switching uncertainty |  |  | 0.26 | Rectangular | 1.73 | 1 |  |  |  |  | 0.15 |  |
| Combined standard uncertainty (1σ) (dB) |  |  |  |  |  |  |  |  |  |  |  |  | TBD |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) |  |  |  |  |  |  |  |  |  |  |  |  | TBD |

| Table 6.6-3: PWS MU value derivation for EEIRP measurement, FR1 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty source | Uncertainty value (dB) |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui (dB) |  |  |  |  |  |
|  |  |  |  | 4.2<f≤7.125 GHz |  |  |  |  |  |  |  | 4.2<f≤7.125 GHz |  |
| Stage 2: BS measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A7-1a | Misalignment and pointing error of BS |  |  | 0.10 | Rectangular | 1.73 | 1 |  |  |  |  | 0.06 |  |
| C1-1 | Uncertainty of the RF power measurement equipment (e.g. spectrum analyzer, power meter) |  |  | 0.50 | Gaussian | 1.00 | 1 |  |  |  |  | 0.50 |  |
| A7-2a | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for BS antenna |  |  | 0.20 | Rectangular | 1.73 | 1 |  |  |  |  | 0.12 |  |
| A7-3 | RF leakage (calibration antenna connector terminated) |  |  | 0.09 | Gaussian | 1.00 | 1 |  |  |  |  | 0.09 |  |
| A7-4a | QZ ripple experienced by BS |  |  | 0.77 | Rectangular | 1.73 | 1 |  |  |  |  | 0.44 |  |
| A7-5 | Miscellaneous uncertainty |  |  | 0.00 | Gaussian | 1.00 | 1 |  |  |  |  | 0.00 |  |
| A7-14 | System non-linearity |  |  | 0.22 | Rectangular | 1.73 | 1 |  |  |  |  | 0.13 |  |
| A7-13 | Frequency flatness of test system |  |  | 0.13 | Rectangular | 1.73 | 1 |  |  |  |  | 0.08 |  |
| A2-21 | Measurement system dynamic range uncertainty |  |  | [0.50] | Rectangular | 1.73 | 1 |  |  |  |  | [0.30] |  |
| *new | [EEIRP sampling error] |  |  | TBD | [TBD] | 1 | 1 |  |  |  |  | TBD |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Uncertainty of network analyzer |  |  | 0.20 | Gaussian | 1.00 | 1 |  |  |  |  | 0.20 |  |
| A7-6 | Mismatch (i.e. reference antenna, network analyser and reference cable) |  |  | 0.45 | U-shaped | 1.41 | 1 |  |  |  |  | 0.32 |  |
| A7-7 | Insertion loss of receiver chain |  |  | 0.18 | Rectangular | 1.73 | 1 |  |  |  |  | 0.10 |  |
| A7-3 | RF leakage (calibration antenna connector terminated) |  |  | 0.09 | Gaussian | 1.00 | 1 |  |  |  |  | 0.09 |  |
| A7-8 | Influence of the calibration antenna feed cable |  |  | 0.10 | Rectangular | 1.73 | 1 |  |  |  |  | 0.06 |  |
| C1-4 | Uncertainty of the absolute gain of the reference antenna |  |  | 0.43 | Rectangular | 1.73 | 1 |  |  |  |  | 0.25 |  |
| A7-9 | Misalignment of positioning system |  |  | 0.00 | Exp. normal | 2.00 | 1 |  |  |  |  | 0.00 |  |
| A7-1b | Misalignment and pointing error of calibration antenna |  |  | 0.05 | Rectangular | 1.73 | 1 |  |  |  |  | 0.03 |  |
| A7-10 | Rotary joints |  |  | 0.00 | U-shaped | 1.73 | 1 |  |  |  |  | 0.00 |  |
| A7-2b | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna |  |  | 0.18 | Rectangular | 1.73 | 1 |  |  |  |  | 0.10 |  |
| A7-4b | QZ ripple experienced by calibration antenna |  |  | 0.20 | Rectangular | 1.73 | 1 |  |  |  |  | 0.12 |  |
| A7-11 | Switching uncertainty |  |  | 0.02 | Rectangular | 1.73 | 1 |  |  |  |  | 0.01 |  |
| A7-12 | Field repeatability |  |  | 0.20 | Gaussian | 1.00 | 1 |  |  |  |  | 0.20 |  |
| Combined standard uncertainty (1σ) (dB) |  |  |  |  |  |  |  |  |  |  |  |  | TBD |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) (dB) |  |  |  |  |  |  |  |  |  |  |  |  | TBD |

References

[1]	R4-2419822	TP to TR 38.908 – section 6.5– Huawei

[2] R4-2419207 TR 38.908 v1.2.0 Protection of fixed satellite service (FSS) UL within 6425 to 7125 MHz - ZTE