3GPP TSG-RAN4 Meeting #116	R4-2509916

# Bengaluru, India, 25th - 29th August, 2025

Agenda item:		7.10.3.3

Source:				Keysight Technologies UK Ltd, Spirent Communications, CAICT, ETS-Lindgren

Title:					On Preliminary MU Budget for DynMIMO OTA TCs

Document for:		Approval

# Introduction

This contribution is reviewing options to define the preliminary MU budget for the dynamic MIMO OTA test cases in RAN4 with the understanding that RAN5 will perform a more through review and finalize the MU/MTSU.

# Discussion

The novel dynamic MIMO OTA test cases, part of the Rel-19 WI on FR1 TRP (Total Radiated Power), TRS (Total Radiated Sensitivity) and MIMO OTA (Over the Air) testing enhancements [1] and defined in [2], have various distinct differences when compared to the existing LTE [3]/FR1 [4][5] MIMO OTA TCs as outlined in Table 1.

Table 1: Key differences between existing LTE/FR1 MMO OTA and FR1 Dynamic MIMO OTA TCs

| Parameter | LTE/FR1 MIMO OTA | FR1 Dynamic MIMO |
| --- | --- | --- |
| Performance Metric based on | Sensitivity/Power [dBm] | Throughput [Mbps] |
| Measurement Channel | Fixed RMCs | Variable RMC (VRMC)/Dynamic Link Adaptation |
| Channel Models | Static CDL (FR1)/SCME (LTE) CMs | Dynamic CMs with time-varying CDL CMs and route |

The following sections further individually discuss the impact these parameters have on the MU budget.

# MU Impact: Performance Metric

While the LTE and FR1 MIMO OTA test plans [3], [4], [5] all focused on radiated sensitivities, i.e., Total Radiated Multi-antenna Sensitivity (TRMS) in dBm at fixed TP levels, the FR1 Dynamic MIMO OTA test plan [6] is solely based on absolute Throughput (TP), specifically CDF of the Total Measured Throughput (CTMT), in Mbps.

One could therefore argue that the MU/MTSU could/should be based on TP as well. However, such an approach would be rather novel and correspond to significant complexities, e.g., end-to-end assessment of MTSU, assessment of a lab’s MU, golden device requirements to assess the MU, etc. It should be noted that demodulation/performance conformance testing for LTE [7] and FR1 [8] is also commonly based on average TP measurements and the corresponding MUs/MTSUs for TCs including fading are defined in terms of powers [9], [10], i.e., dB, e.g., in Clause 7.1.1 of [10].

### F.1.1.2 Measurement of Demod Performance requirements This clause defines the maximum test system uncertainty for Demod Performance requirements. The maximum test system uncertainty allowed for the measurement uncertainty contributors are defined in Table F.1.1.2-1. Table F.1.1.2-1: Maximum measurement uncertainty values for the test system for FR1 (up to 7.125 GHz) and Channel BW ≤ 40 MHz MU contributor; Unit; Value; Comment. AWGN flatness and signal flatness, max deviation for any Resource Block, relative to average over BWconfig; dB; ±2.0; Same as in LTE. Signal to noise ratio uncertainty; dB; ±0.3; Same as in LTE. Signal to noise ratio variation; dB; ±0.5; Same as in LTE. Fading profile power uncertainty for 1Tx; dB; ±0.5; Same as in LTE. Fading profile power uncertainty for 2Tx; dB; ±0.7; Same as in LTE. Fading profile power uncertainty for 4Tx; dB; ±0.7; Same as in LTE. NOTE 1: MU and TT for >6GHz are working assumption pending confirmation from TE vendors.. The maximum test system uncertainty for test cases defined in section 5 is defined in Table F.1.1.2-2. …

Given the precedent of demod conformance test cases and to avoid significant delays and complexities finalizing the dynamic MIMO OTA test plan, it proposed to continue to define the preliminary MU budget for dynamic MIMO OTA TCs in a similar fashion as for the LTE/FR1 MIMO OTA test cases, i.e., in terms of powers and dB.

Proposal 1: Continue to define the preliminary MU budget for dynamic MIMO OTA test cases in a similar fashion as for the LTE/FR1 MIMO OTA test cases, i.e., in terms of powers and dB.

For certification organizations of RAN5/CTIA Dynamic MIMO OTA TCs, the definition of the MTSU in RAN5/CTIA in dB has no additional impact, i.e., the determination whether a lab’s test system meets the MTSU is based on the assessment of individual MU elements and comparing the total expanded MU with the corresponding MTSU. However, the discrepancy of the test metric based on TP [Mbps] and the MU/MTSU based on power [dB] will have an impact in RAN4, specifically the definition of pass/fail requirements. Prior to defining performance requirements using a dedicated measurement campaign, volunteer labs must undergo a separate alignment campaign using reference devices sent around between these volunteer labs. If the results of a lab deviate by at most by the preliminary MU/MTSU from the average measured by all labs, the lab is considered aligned and can submit measurement results for the measurement campaign; labs that are not aligned, are not permitted to participate in the measurement campaign. As such, if the MU/MTSU is not aligned with the final metric, e.g., dB vs Mbps, the assessment of whether a lab is aligned or not is not possible and a different approach is needed in RAN4 if it is decided to perform a Dynamic MIMO OTA lab alignment activity in future releases.

Proposal 2: Given the discrepancy of final metric based on TP [Mpbs] and the preliminary MU based on power [dB], RAN4 to determine a new/different lab alignment approach if it is decided to perform a Dynamic MIMO OTA lab alignment activity in future releases.

# MU Impact: Measurement Channel

The question is whether the move from fixed RMCs, used for existing LTE/FR1 MIMO OTA TCs [3], [4], [5], to VRMCs/link adaptation, adopted by FR1 Dynamic MIMO OTA TCs [6], is adding any additional MU. Extensive campaigns could be performed to assess whether different gNBs and their respective schedulers yield differences in the final metric. However, the MU of demodulation test cases with fixed and VRMC including fading has been assessed for FR1 already and it was concluded that the MU is the same as captured in Clause F.1 of [10]. For instance, the MTSU in Table F.1.1.2-2 for TC 5.6.2.1.1_1 (2Rx FDD FR1 PDSCH performance with link adaptation - 2x2 MIMO for both SA and NSA) is 1dB

which matches the MTSU for the corresponding TC with fixed RMC, i.e., TC 5.2.2.1.1_1 (2Rx FDD FR1 PDSCH mapping Type A performance - 2x2 MIMO with baseline receiver for both SA and NSA) in Table F.1.1.2-2

Given this precedent of demod conformance test cases, it is proposed to assume that dynamic link adaptation/VRMC for dynamic MIMO OTA TCs does not add any additional MU when compared to the static MIMO OTA TCs with fixed RMCs.

Proposal 3: Assume that dynamic link adaptation/VRMC for dynamic MIMO OTA TCs does not add any additional MU when compared to the static MIMO OTA TCs with fixed RMCs.

# MU Impact: Channel Models

The channel models are emulated by the Channel Emulator typically placed between the gNB and the OTA measurement probes. Its MU is captured in the MU budget accordingly, e.g., UID 5 of Table B.1-1 in [11] for the static MIMO OTA TCs.

Table B.1-1: Preliminary measurement uncertainty budget for FR1 MPAC system UID; Description of uncertainty contribution; Example value (410 MHz <f ≤ 3 GHz) [dB]; Example value (3 GHz < f ≤ 7.125GHz) [dB]; Distribution of the probability; Std Uncertainty (410 MHz < f ≤ 3 GHz) [dB]; Std Uncertainty (3 GHz < f ≤ 7.125GHz) [dB]. Stage 2: DUT measurement. 1; Mismatch for measurement process; 0; 0; U-Shaped; 0; 0. 2; Measure distance uncertainty; 0; 0; Normal; 0; 0. 3; Quality of quiet zone; 0.5; 0.5; Actual; 0.5; 0.5. 4; Base Station simulator; 1.3; 1.3; Normal; 0.65; 0.65. 5; Channel Emulator - absolute output power - output signal stability - output stability with temperature; 1.5 0.5 0.4; 1.5 0.5 0.4; Actual (Normal - power; rect-stability); 0.84; 0.84. 6; Amplifier uncertainties; 0.7; 0.7; Rectangular; 0.40; 0.40. 7; Random uncertainty; 0.4; 0.4; Normal; 0.20; 0.20. 8; Throughput measurement: output level step resolution; 0; 0; Rectangular; 0; 0. 9; Signal flatness; 0; 0; Normal; 0; 0. Stage 1: Calibration measurement. 10; Mismatch for calibration process - loopback cable path - system input path - reference antenna; 0.2; 0.2; U-Shaped; 0.14; 0.14. 11; Reference antenna positioning misalignment; 0; 0; Normal; 0; 0. 12; Quality of quiet zone; 0.5; 0.5; Rectangular; 0.29; 0.20. 13; Total uncertainty of the Network Analyzer; 0.2; 0.5; Normal; 0.10; 0.25. 14; Uncertainty of an absolute gain of the calibration antenna; 0.58; 0.58; Normal; 0.29; 0.29. 15; Offset of the Phase Center of the Reference Antenna; 0; 0; Normal; 0; 0. Total Expanded Uncertainty, U, with 95 % Confidence Interval; 2.61; 2.65.

The dynamic MIMO OTA TCs add the temporal variation of CDL models and the traversing of the UE on a synthesized route, i.e., an additional MU could be applicable due to the dynamic MIMO OTA features when compared to static MIMO OTA. It is therefore proposed to introduce an additional MU element ‘Dynamic Channel Modelling’ to the preliminary RAN4 MU budget. At this point, it is suggested to assign a fixed [0.75] dB value to this new MU element and let RAN5 finalize the definition of this MU element. It should be noted that no technical justification is available yet to justify this value.

Proposal 4: Introduce an additional MU element ‘Dynamic Channel Modelling’ to the preliminary RAN4 MU budget to take into account the temporal variation of CDL models and the traversing of the UE on a synthesized route.

Proposal 5: Assign a fixed [0.75] dB value to MU element ‘Dynamic Channel Modelling’ and let RAN5 finalize the definition of this MU element.

# Preliminary RAN4 MU Budget for FR1 Dynamic MIMO

With the proposed MU aspects for the FR1 dynamic MIMO TCs presented in this contribution, the preliminary RAN4 MU budget is tabulated in Table 2, i.e., the only difference between the preliminary MU for dynamic vs static MIMO OTA TCs is the addition of the fixed MU element ‘Dynamic Channel Modelling’ (UID 6)

Table 2: Preliminary measurement uncertainty budget for FR1 MPAC system and Dynamic MIMO OTA

| UID | Description of uncertainty contribution | Example value (410 MHz <f ≤ 3 GHz) [dB] | Example value (3 GHz < f ≤ 7.125GHz) [dB] | Distribution of the probability | Std Uncertainty (410 MHz < f ≤ 3 GHz) [dB] | Std Uncertainty (3 GHz < f ≤ 7.125GHz) [dB] |
| --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |
| 1 | Mismatch for measurement process | 0 | 0 | U-Shaped | 0 | 0 |
| 2 | Measure distance uncertainty | 0 | 0 | Normal | 0 | 0 |
| 3 | Quality of quiet zone | 0.5 | 0.5 | Actual | 0.50 | 0.50 |
| 4 | Base Station simulator | 1.3 | 1.3 | Normal | 0.65 | 0.65 |
| 5 | Channel Emulator |  |  | Actual | 0.84 | 0.84 |
|  | - absolute output power | 1.5 | 1.5 | (Normal - power; |  |  |
|  | - output signal stability | 0.5 | 0.5 | rect-stability) |  |  |
|  | - output stability with temperature | 0.4 | 0.4 |  |  |  |
| 6 | Dynamic Channel Modelling (Note 1) | [0.75] | [0.75] | Normal | [0.38] | [0.38] |
| 7 | Amplifier uncertainties | 0.7 | 0.7 | Rectangular | 0.40 | 0.40 |
| 8 | Random uncertainty | 0.4 | 0.4 | Normal | 0.2 | 0.2 |
| 9 | Throughput measurement: output level step resolution | 0 | 0 | Rectangular | 0 | 0 |
| 10 | Signal flatness | 0 | 0 | Normal | 0 | 0 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |
| 11 | Mismatch for calibration process | 0.2 | 0.2 | U-Shaped | 0.10 | 0.10 |
|  | - loopback cable path |  |  |  |  |  |
|  | - system input path |  |  |  |  |  |
|  | - reference antenna |  |  |  |  |  |
| 12 | Reference antenna positioning misalignment | 0 | 0 | Normal | 0 | 0 |
| 13 | Quality of quiet zone | 0.5 | 0.5 | Rectangular | 0.29 | 0.29 |
| 14 | Total uncertainty of the Network Analyzer | 0.2 | 0.5 | Normal | 0.1 | 0.25 |
| 15 | Uncertainty of an absolute gain of the calibration antenna | 0.58 | 0.58 | Normal | 0.29 | 0.29 |
| 16 | Offset of the Phase Centre of the Reference Antenna | 0 | 0 | Normal | 0 | 0 |
| Total Expanded Uncertainty, U, with 95 % Confidence Interval |  |  |  |  | [2.70] | [2.74] |
| Note 1: This MU value will need to be addressed/finalized by RAN5. |  |  |  |  |  |  |

Proposal 6: Endorse the preliminary MU budget for Dynamic MIMO OTA TCs (FR1 MPAC system) in Table 2.

# Conclusion

The following observations and conclusions were made in this contribution.

Proposal 1: Continue to define the preliminary MU budget for dynamic MIMO OTA test cases in a similar fashion as for the LTE/FR1 MIMO OTA test cases, i.e., in terms of powers and dB.

Proposal 2: Given the discrepancy of final metric based on TP [Mpbs] and the preliminary MU based on power [dB], RAN4 to determine a new/different lab alignment approach if it is decided to perform a Dynamic MIMO OTA lab alignment activity in future releases.

Proposal 3: Assume that dynamic link adaptation/VRMC for dynamic MIMO OTA TCs does not add any additional MU when compared to the static MIMO OTA TCs with fixed RMCs.

Proposal 4: Introduce an additional MU element ‘Dynamic Channel Modelling’ to the preliminary RAN4 MU budget to take into account the temporal variation of CDL models and the traversing of the UE on a synthesized route.

Proposal 5: Assign a fixed [0.75] dB value to MU element ‘Dynamic Channel Modelling’ and let RAN5 finalize the definition of this MU element.

Proposal 6: Endorse the preliminary MU budget for Dynamic MIMO OTA TCs (FR1 MPAC system) in Table 2.

# References

- RP-241623, WI on FR1 TRP (Total Radiated Power), TRS (Total Radiated Sensitivity) and MIMO OTA (Over the Air) testing enhancement Phase 3, vivo, CAICT, 3GPP TSG RAN Meeting #104, June 2024

- TR 38.762, Multiple Input Multiple Output (MIMO) Over-the-Air (OTA) dynamic test methodology for FR1 UEs

- TR 37.977, Verification of radiated multi-antenna reception performance of User Equipment (UE)

- TS 38.151, Multiple Input Multiple Output (MIMO) Over-the-Air (OTA) performance requirements

- TR 38.761, Measurements of Multiple Input Multiple Output (MIMO) Overthe-Air (OTA) performance of User Equipment (UE)

- TR 38.762, Multiple Input Multiple Output (MIMO) Over-the-Air (OTA) dynamic test methodology for FR1 UEs

- TS 36.101, User Equipment (UE) radio transmission and reception

- TS 38.101-4, User Equipment (UE) radio transmission and reception; Part 4: Performance requirements

- TS 36.521-1, User Equipment (UE) conformance specification;Radio transmission and reception; Part 1: Conformance Testing

- TS 38.521-4: User Equipment (UE) conformance specification; Radio transmission and reception; Part 4: Performance requirements

- TS 38.551, UE Conformance - Multiple Input Multiple Output (MIMO) Over-the-Air (OTA) requirements for NR UEs