3GPP TSG-RAN WG4 Meeting #116	R4-2510482

Bengaluru, India, August 25th – 29th, 2025

Source:	GIST, NRRA

Title:	Discussion on OTA TRP Measurement Using Simultaneously Active Probes

Agenda Item:		7.10.2

Document for:	Discussion

Introduction

With the increasing demand for faster and more efficient OTA testing methodologies, especially for modulated signals and beamforming-based systems, the use of simultaneously active probes has emerged as a practical approach worthy of further exploration. This contribution presents a practical demonstration of TRP measurement using an 11-element active probe array developed under the joint collaboration of the National Radio Research Agency (RRA) of Korea, Chosun University (CSU), Gwangju Institute of Science and Technology (GIST), and Dabin Systems, known as the RRA-MAPS system. The focus is to highlight the accuracy and efficiency benefits compared to conventional mechanically scanned approaches.

Measurement Setup

We performed OTA TRP measurements using a calibrated near-field to far-field transform (NFTF) system equipped with simultaneously active probes. The Device Under Test (DUT) was a horn antenna (SCHWARZBECK BBHA 9120 C), operating across 2 – 18 GHz. It was fed with a 5G NR signal generated at 3.55 GHz with the following signal parameters:

- Channel bandwidth: 100 MHz

- Modulation: QPSK

- Subcarrier spacing: 30 kHz

- Number of Resource Blocks: 273

- Conforming to TS 38.141: G-FR1-A1-2 test model

The probe arc array consisted of 11 probes located at a radius of 0.8 m. Probe elevation angles were: 7.5°, 22.5°, and in 15° steps up to 172.5°. All probes operated in V/H dual polarization and captured amplitude and phase data simultaneously. For full-sphere coverage, azimuth angle (AZ) was rotated in 5° steps (72 positions) and elevation arc goniometer moved ±5° around its center, yielding 3 effective elevation scans. This resulted in 216 mechanical positions (3 elevation × 72 azimuth) and 2,376 total measurement points (11 probes × 3 EL positions × 72 AZ rotations).

Synchronization across probes was ensured using rubidium atomic clock reference distribution, and the system was calibrated for both amplitude and phase using a known-gain source.

Figure: RRA-MAPS with 11 Probes and Measured V and H Patterns

TRP Calculation and Results

TRP was calculated using Equation (5.1) from TR 38.870, based on integrated power over the full 3D reconstructed pattern from near-field measurements.

Key measurements:

- Signal Generator Output: −20.58 dBm

- Measured TRP using simultaneously active probes: −22.19 dBm

- Error: 1.61 dB

The current result was obtained using a prototype system in an early-stage test environment. As the chamber tuning and probe calibration procedures are further refined, additional accuracy improvements are anticipated.

With a mechanical rotation speed of approximately 2 degrees per second, the full TRP measurement process can be completed in about 10 minutes.

Conclusion

This contribution demonstrates that the system developed under the RRA-MAPS project, using simultaneously active probes, can be successfully used for OTA TRP measurement of 5G NR signals. The results are within 1.61 dB of the known source power, while providing substantial time savings, good accuracy, and mechanical efficiency.

The system setup is particularly suitable for:

- Rapid production-line diagnostics

- Testing wideband modulated signals

- Evaluation of small or mid-size antennas in a controlled near-field chamber

- Reference

- TR 38.870: "Enhanced Over-the-Air (OTA) test methods for NR FR1 Total Radiated Power (TRP) and Total Radiated Sensitivity (TRS)"

- TR 38.810: "NR; Study on test methods"

- R4-2307504: "CR for TR38.810 Integrating simultaneously active probe concept in the NFTF method"

- R4-2400727: "Discussion on NFTF System for UE Transmitter Characteristics"

- R4-2407599: "Discussion on the Operation of the Probe Module Array"

- R4-2411308: "UE RF Testing Procedure under NFTF Measurement Setup"

- R4-2507095: "Discussion on NFTF Method for OTA TRP Measurement"