3GPP TSG-RAN WG4 Meeting #86bis	R4-1805385

Melbourne, Vic, Australia, 16th – 20th April 2018

Source:	Nokia, Nokia Shanghai Bell

Title:	On pre-scan, peak EIRP and fast TRP measurements for TX spurious and EMC emissions of NR BS type 2-O

Agenda item:	7.7.4.3

Document for:	Discussion

# 1	Introduction

During the last RAN4#86 meeting, the WF on OTA TRP conformance requirements [1] was approved, which outlined a list of open issues to be addressed in forthcoming meetings.

In [2], we quantitively showed that the TRP measurement duration of spurious emission for mmWave NR base stations using a comprehensive TRP measurement method is impractical long (i.e., in the order of months), which was attributed to a very wide range of spurious frequencies, the number of TRP measurement points and the measurement bandwidth. Consequently, we have been advocating pre-scan measurements [3]-[6] during spurious and EMC emission testing in view of reducing the TRP measurement duration to a reasonable time for the last few meetings.

Based on the pre-scan concept described in [3], we outline the detailed pre-scan methodology for spurious and EMC emission measurement of NR BS type 2-O in a typical indoor Anechoic Chamber system set up. To further reduce measurement duration, we also advocate peak EIRP and fast TRP measurements to be performed on a limited number of spurious frequencies with the strongest unwanted emission power levels identified in the pre-scan. As compared with peak EIRP and fast TRP measurement methods, the comprehensive TRP measurement method would yield the least possible radiated power levels. Therefore, if the measured peak EIRP and fast TRP can meet the unwanted emission limit then it is not necessary to perform comprehensive TRP measurements.

The proposed pre-scan, peak EIRP and TRP measurement methods can be used for testing spurious and EMC emissions ranging from 30 MHz to the second harmonic of the upper frequency edge of the downlink operating band or [60 GHz] in a single test set up. The proposed methods are applicable to other test systems such as Compact Antenna Test Range.

# Discussion

## 2.1	Radiated spurious and EMC emissions

According to TR 38.817-02, radiated spurious emissions can no longer be differentiated from EMC emissions for over the air testing. Consequently, one unified OTA emission requirement is specified for both spurious and EMC for NR BS type 2-O. As a result, one single OTA test system set up can be used for measuring both the spurious and EMC emissions.

The unified OTA requirement for both Tx spurious and EMC emissions is shown in Figure 1. The frequency range and measurement bandwidth for EMC is the same as the spurious emissions.

2nd harmonic/ [60 GHz]

ΔfOBUE

ΔfOBUE

DL

30 MHz

EMC/Spurious emission limits

EMC/Spurious emission limits

Figure 1: Radiated spurious and EMC emission frequency range

## 2.2	Indoor Anechoic Chamber test set up for spurious and EMC emissions

TR 37.842 outlines four OTA test systems which are used for testing the EIRP accuracy requirement:

- Indoor Anechoic Chamber

- Compact Antenna Test Range

- One Dimensional Compact Range Chamber

- Near Field Test Range

Figure 2 shows one example of a typical indoor Anechoic Chamber test system set up.

A

Measurement Distance  r          Type equation here.   Type equation here.

Measurement Antenna

DUT

h3          Type equation here.   Type equation here.

Vertical movement

Styrofoam table

h1          Type equation here.   Type equation here.

h2          Type equation here.   Type equation here.

ϕ

Turntable

Ground Plane

D

PC Computer

Positioner Controller

LNA

Spectrum Analyser

Filter

Figure 1: A typical indoor Anechoic Chamber system set up for spurious and EMC emission measurements

Stage 1: Calibration:

- The measurement antenna should be placed at a fixed distance of r meters from the periphery of the DUT such that the measurement is performed in the far field of the transmitting DUT antenna; the far field conditions are

far field conditions  r ≫ 2D2λr ≫ D         r ≫ λ

where D is the largest dimension of the DUT or measurement antenna and λ is the wavelength at the frequency of measurement.

For the lower chunks of the spurious frequency spectrum (30 – 100 MHz), the far field distance is impractically large, which prohibits far field measurements in legacy test chambers. As such, the far field conditions cannot be met at the configured distance r. We carried out experiments [7] to validate the feasibility of such measurements and the experimental results concluded that the minimum distance r = 3 m.

- Place a SGH antenna in the position of the DUT.

- Calculate the calibration value between A and D, LA-D.

Stage 2: Measurement:

- Place NR BS type 2-O (i.e., DUT) on an RF transparent turn-table at a nominal height [h1 = 1.5 m] above the ground plane.

- Configure all NR BS type 2-O TX branches according to a maximum power requirement for a single carrier.

- If the DUT cannot be configured as described in Step 2 because it can only support multiple beams, then set the DUT to generate several beams at the maximum transmit power levels utilizing all the BS TX branches.

- Configure the DUT to transmit the test signal according to FFS.

- RF filtering (e.g., a notch, low pass or high pass filter) is necessary to increase the dynamic range of measurement.

- Typically, power cables attached to the DUT may not radiate in the spurious and EMC frequency range. Therefore, for emission measurements, the measurement antenna’s half-power beamwidth need not cover the cables supplying electrical power and GPS cables do not wrap around the DUT.

- Set the spectrum analyser start frequency (fstart) and stop frequency (fstop) to the lowest (fl) and highest (fh) operating frequency of the measurement antenna, where the measurement antenna bandwidth = fh-fl.

- Set the spectrum analyser resolution bandwidth in accordance with the MBW requirement as specified in TS 37.105 or TS 38.104.

2a: Pre-scan:

- Select the spectrum analyser power detection mode to peak detector.

- Configure the sweep time of the spectrum analyser to auto-couple.

- Rotate the DUT on the turn-table through 360° in the azimuth (ϕ) direction in a discrete step of [∆ϕ= 45°]. The total number of ∆ϕ-sectors, J= 360°∆ϕ, and let j denote the jth ∆ϕ-sector. For spurious frequencies above 18 GHz, the 3 dB beamwidth of the measurement antenna may not completely cover the DUT, then  ∆ϕ should be reduced.

- For the jth ∆ϕ-sector, raise and lower the measurement antenna in the vertical direction from [h2 = 1 m] to [h3 = 4 m] to measure emission power levels.

- Record unwanted emission power levels (Pmeas), spurious emission frequency, ∆ϕ-sector j and antenna height h for each measurement bandwidth in list L.

- Rotate the DUT to the next ∆ϕ-sector and repeat Steps 12 to 13 until all J sectors are completed.

- Entries whose Pmeas ≈ system noise floor are discarded from L.

2b: Peak EIRP measurements:

- Select the spectrum analyser power detection mode to power averaging (RMS).

- Configure the sweep time of the spectrum analyser to less than the transmission duration; for non-continuous transmission, employ the spectrum analyser gated sweep (i.e., gate triggered in such a way that the analyser only sweeps when the DUT is transmitting).

- If L ≠ {} then select the spurious frequency with the highest unwanted emission power level recorded in L.

- Repeat Steps 7-8 if a different measurement antenna is used.

- Rotate the turn-table to the jth ∆ϕ-sector where the highest unwanted emission power level was recorded in L.

- Raise and lower the measurement antenna in the vertical direction from [h2 = 1 m] to [h3 = 4 m] to measure the maximum power level (Pmax,j).

- Measurements should be made with the measurement antenna oriented in both orthogonal polarizations (Pmax,p1,j, Pmax,p2,j). If the measurement antenna supports single polarization, then it should be oriented in one polarization at a time.

- Calculate EIRPmax,j = EIRPp1,j + EIRPp2,j (in linear units), where EIRPp1,j = Pmax,p1,j + LA-D (in dB) and EIRPp2,j = Pmax,p2,j + LA-D (in dB). If EIRPmax,j < limitMBW, then the spurious frequency under measurement is compliant and no further measurements are needed.

2c: Fast TRP measurements:

- If EIRPmax,j ≥limitMBW,  then repeat Steps 20 to 23 for the remaining J – 1 ∆ϕ-sectors and calculate

TRPfast≈1Jj=1JEIRPmax,j,                ∀ EIRPmax,j>limitMBW-[20 dB]

If TRPfast< limitMBW, then the spurious frequency under measurement is compliant and no further

measurements are needed.

- Repeat Steps 18 to 24 for a total number of [10] spurious frequencies.

# 3	Conclusions

In this document, we have outlined a comprehensive test methodology for spurious and EMC emission measurements of NR BS type 2-O. It comprises three phases:

- Pre-scan – fast measurements are performed over the entire spurious frequency range to identify spurious frequencies with useful emissions; it is not meant to measure the actual emission power levels.

- Peak EIRP measurements – the maximum EIRP is measured for the top [10] spurious frequencies with the strongest emission power levels; the measured maximum EIRP is applied to the unwanted spurious/EMC emission requirement.

- Fast TRP measurements – this method is carried out if the above maximum EIRP does not meet the emission limit by computing the average of EIRP in which each EIRP is the maximum power level measured in different azimuth sectors. The fast TRP is applied to the unwanted spurious/EMC requirement.

The pre-scan, peak EIRP and fast TRP measurement methods are key to reduce spurious and EMC emission measurement time. As compared with the peak EIRP and fast TRP measurements, the comprehensive TRP measurement method would yield the least emission power levels. Therefore, if the measured peak EIRP and fast TRP can meet the unwanted emission requirement then it is not necessary to perform comprehensive TRP measurements. Hence, the comprehensive TRP measurement method acts as a fall back only.

# References

- R4-1803392, WF on OTA TRP conformance requirements, Ericsson

- R4-1713793, On TRP measurement duration of NR base stations, Nokia, Nokia Shanghai Bell

- R4-1803009, On OTA pre-scanning methods for spurious measurements of NR BS type 2-O, Nokia, Nokia Shanghai Bell

- R4-1800912, On OTA methods for spurious measurements of NR BS type 2-O, Nokia, Nokia Shanghai Bell

- R4-1713796, Spurious emission measurement procedures for NR BS type 2-O, Nokia, Nokia Shanghai Bell

- Nokia’s contribution for equivalent isotropically radiated power (EIRP), total radiated power (TRP), and channel edge measurements, 47CFR30

- D Moongilan, Wireless MIMO Base Station Near and Far Field Measurement Challenges, IEEE International Symposium on Electromagnetic Compatibility & Signal/Power Integrity, Washington, DC, 2017