3GPP TSG-RAN4 Meeting #88	R4- 1810508

Gothenburg, Sweden, 20th – 24th August 2018

Source:	Ericsson

Title:	MU framework and values for OBUE for anechoic chamber for FR2

Agenda item:	7.9.4.3.3.1

Document for:	Discussion

# Introduction

This contribution proposes a measurement uncertainty framework and assessment for OTA operating band unwanted emissions measurement in a compact antenna test range which is one possible method for conformance testing of AAS base stations. We believe that similarly to FR1, the OBUE MU can be based upon the EIRP accuracy MU budget.

# Discussion

Unwanted emissions are specified with TRP as the metric. If TRP is measured by integration of individual EIRP measurements, then the basic MU budget can be derived from the EIRP accuracy MU budget. Additional uncertainty relating to the so-called “systematic error” may need to be added; this is the subject of a separate discussion.

Two issues are under discussion in relation to the EIRP accuracy MU budget; one is the instrument uncertainties. The uncertainties quoted here relate to experience in the lab with carefully calibrated instruments. Instrument uncertainty may be the subject of separate generic discussion.

A second discussion point is on the need for an LNA in the measurement setup. The lowest emissions mask level is -13dBm for a WA BS. Assuming around e.g. 80-90dB pathloss in the far field (including test antenna gain and cable losses), then the measured TRP level would be around -90 to -100 dBm/MHz. There may not be so much DUT antenna pattern variation for unwanted emissions, and these levels can be below the noise floor of a spectrum analyser. An LNA may well be needed. It is not clear that including an LNA would have a significant impact on the MU budget assuming careful calibration, however some more discussion could be had on any additional MU margin.

In addition, the used EIRP uncertainty is assuming long enough sweeping time. Possibly, to speed up the measurements, a shorter sweeping time could affect the proposed measurement uncertainty. Such a change in accuracy may be counteracted by an averaging effect arising from taking multiple independent measurements around the sphere.

Based on radiated transmit power in MU, framework on uncertainty budget have been reused.

Indoor anechoic chamber uncertainty contributions for AAS BS OTA operating band unwanted emissions measurement

| UID | Description of uncertainty contribution | Details in annex of [1] |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | B1-1 |
| 2 | Pointing misalignment between the AAS BS and the receiving antenna. | B1-2 |
| 3 | Quality of quiet zone | B1-3 |
| 4 | Polarization mismatch between the AAS BS and the receiving antenna | B1-4 |
| 5 | Mutual coupling between the AAS BS and the receiving antenna | B1-5 |
| 6 | Phase curvature | B1-6 |
| 7 | Uncertainty of the RF power measurement equipment | E |
| 8 | Impedance mismatch in the receiving chain | B1-8 |
| 9 | Random uncertainty | B1-9 |
| Stage 1: Calibration measurement |  |  |
| 10 | Impedance mismatch between the receiving antenna and the network analyzer | B1-10 |
| 11 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | B1-11 |
| 12 | Impedance mismatch between the reference antenna and the network analyzer | B1-12 |
| 13 | Quality of quiet zone | B1-3 |
| 14 | Polarization mismatch for reference antenna | B1-4 |
| 15 | Mutual coupling between the reference antenna and the receiving antenna | B1-5 |
| 16 | Phase curvature | B1-6 |
| 17 | Uncertainty of the Network Analyzer | E |
| 18 | Influence of the reference antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | B1-14 |
| 19 | Reference antenna feed cable loss measurement uncertainty | B1-15 |
| 20 | Influence of the receiving antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | B1-16 |
| 21 | Uncertainty of the absolute gain of the reference antenna | E |
| 22 | Uncertainty of the absolute gain of the receiving antenna | B1-18 |

# Compact antenna test range uncertainty assessment for AAS BS OTA operating band unwanted emissions measurement

| UID | Uncertainty source | Uncertainty value 24.25<f <29.5GHz | Uncertainty value 37<f <40GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] 24.25<f <29.5GHz | Standard uncertainty ui [dB] 37<f <40GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Positioning misalignment between the AAS BS and the reference antenna | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 3 | Quality of quiet zone | 0.10 | 0.10 | Gaussian | 1 | 1 | 0.10 | 0.10 |
| 4 | Polarization mismatch between the AAS BS and the receiving antenna | 0.018 | 0.018 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 5 | Mutual coupling between the AAS BS and the receiving antenna | 0.00 | 0.00 | Rectangular | √3 | 1 | 0.00 | 0.00 |
| 6 | Phase curvature | 0.07 | 0.07 | Gaussian | 1 |  | 0.07 | 0.07 |
| 7 | Uncertainty of the RF Power Measurement Equipment | 0.50 | 0.70 | Gaussian | 1 | 1 | 0.50 | 0.70 |
| 8 | Impedance mismatch in the receiving chain | 0.42 | 0.42 | U-shaped | √2 | 1 | 0.30 | 0.30 |
| 9 | Random uncertainty | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 10 | Impedance mismatch between the receiving antenna and the network analyzer | 0.43 | 0.57 | U-shaped | √2 | 1 | 0.30 | 0.40 |
| 11 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.43 | 0.43 | Rectangular | √3 | 1 | 0.25 | 0.25 |
| 12 | Impedance mismatch between the reference antenna and the network analyzer. | 0.43 | 0.57 | U-shaped | √2 | 1 | 0.30 | 0.40 |
| 13 | Quality of quiet zone | 0.10 | 0.10 | Gaussian | 1 | 1 | 0.10 | 0.10 |
| 14 | Polarization mismatch for reference antenna | 0.018 | 0.018 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 15 | Mutual coupling between the reference antenna and the receiving antenna | 0.00 | 0.00 | Rectangular | √3 | 1 | 0.00 | 0.00 |
| 16 | Phase curvature | 0.07 | 0.07 | Gaussian | 1 | 1 | 0.07 | 0.07 |
| 17 | Uncertainty of the network analyzer | 0.30 | 0.30 | Gaussian | 1 | 1 | 0.30 | 0.30 |
| 18 | Influence of the reference antenna feed cable | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 19 | Reference antenna feed cable loss measurement uncertainty | 0.10 | 0.10 | Gaussian | 1 | 1 | 0.10 | 0.10 |
| 20 | Influence of the receiving antenna feed cable | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 21 | Uncertainty of the absolute gain of the reference antenna | 0.52 | 0.52 | Rectangular | √3 | 1 | 0.30 | 0.30 |
| 22 | Uncertainty of the absolute gain of the receiving antenna | 0.00 | 0.00 | Rectangular | √3 | 1 | 0.00 | 0.00 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | [0.92] | [1.11] |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | [1.81] | [2.17] |

# Conclusion

This contribution presents an example of the measurement uncertainty framework and budget for OBUE measurements. The basic MU is 1.81dB for 28GHz and 2.17dB for 39GHz; SE arising from the TRP calculation may need to be combined with these values and potentially any additional margin for residual LNA uncertainty.

# References

TS 37.842, Evolved Universal Terrestrial Radio Access (E-UTRA) and Universal Terrestrial Radio Access (UTRA; Radio Frequency (RF) requirement background for Active Antenna System (AAS) Base Station (BS)

Huawei, R4-1808500, “Draft CR for 37.843”, RAN4#87, May 2018