## Slide 1: WF on MU and TT for Rx requirements for FR2
Nokia, Nokia Shanghai Bell
3GPP TSG-RAN WG4 Meeting#88
Gothenburg, Sweden, 20 – 24 August 2018
Agenda Item: 6.1.1
R4-1811634

## Slide 2: Proposals for MU and TT for NR FR2 OTA receiver requirements provided in RAN4#88
[1] R4-1809708, “Proposals on MU and TT for NR FR2 OTA receiver requirements”, Nokia, Nokia Shanghai Bell.
[2] R4-1809884, “NR FR2 OTA measurement uncertainty value for CATR”, Keysight Technologies, NSI-MI Technologies
[3] R4-1810353, “Test tolerance for FR1 and FR2 RX regulatory requirements”, NTT DOCOMO, INC.
[4] R4-1810511, “Measurement uncertainties for EIS for FR2 in indoor anechoic chamber”, Ericsson.
[5] R4-1810512, “Measurement uncertainties for EIS for FR2 in CATR,” Ericsson.
[6] R4-1811264, “Proposed MU budgets for OT EIRP and EIS measurements of BS at mmWave in CATR”, MVG Industries.

## Slide 3: Summary for EIS MU for CATR
Proposals:
  Nokia 3.2
  MVG 3.33
  Keysight/NSI-MI 3.24
  Ericsson by IAC 1.81 (for 37<f<40G)
Way forward for EIS MU:
  3.2
UID | Uncertainty Source | Uncertainty value MVG | Uncertainty value Keysight/NSI-MI | Uncertainty value Ericsson 24.25<f <29.5GHz | Uncertainty value Ericsson 37<f <40GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] MVG | Standard uncertainty ui [dB] Keysight/NSI-MI | Standard uncertainty ui [dB] Ericsson 24.25<f <29.5GHz | Standard uncertainty ui [dB] Ericsson 37<f <40GHz
 |  | FR2 | FR2 | FR2 | FR2 |  |  |  | FR2 | FR2 | FR2 | FR2
Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  | 
1 | Misalignment DUT & pointing error | 0.1 | 0.1 | 0.2 | 0.2 | Exp. normal | 2 | 1 | 0.05 | 0.05 | 0.10 | 0.10
2 | Standing wave between DUT and test range antenna | 0.1 | 0.03 | 0.21 | 0.21 | U-shaped | √2 | 1 | 0.07 | 0.02 | 0.15 | 0.15
3 | RF Signal Generator | 1.23 | 1.23 | 0.57 | 0.57 | Gaussian | 1 | 1 | 1.23 | 1.23 | 0.57 | 0.57
4 | RF leakage & dynamic range, test range antenna cable connector terminated. | 0.1 | 0.1 | 0.01 | 0.01 | Normal | 1 | 1 | 0.1000 | 0.1000 | 0.0100 | 0.0100
5 | QZ ripple with DUT | 0.4 | 0.081 | 0.0928 | 0.0928 | Normal | 1 | 1 | 0.4000 | 0.0810 | 0.0928 | 0.0928
6 | Miscellaneous Uncertainty | 0 | 0 | 0 | 0 | Normal | 1 | 1 | 0 | 0 | 0 | 0
Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  | 
7 | Network Analyzer | 0.4 | 0.2 | 0.3 | 0.3 | Normal | 1 | 1 | 0.40 | 0.20 | 0.30 | 0.30
8 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.5 | 0.74 | 0.43 | 0.57 | U-shaped | √2 | 1 | 0.35 | 0.52 | 0.30 | 0.40
9 | Insertion loss of transmitter chain | 0 | 0.12 | 0.26 | 0.26 | Rectangular | √3 | 1 | 0.00 | 0.07 | 0.15 | 0.15
10 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.1 | 0.1 | 0.01 | 0.01 | Normal | 1 | 1 | 0.10000 | 0.10000 | 0.01000 | 0.01000
11 | Flexing cables & connector repeatability | 0 | 0 | 0.21 | 0.21 | U-shaped | √2 | 1 | 0.000 | 0.000 | 0.148 | 0.148
12 | SGH Calibration uncertainty | 1.6 | 1.6 | 0.52 | 0.52 | Rectangular | √3 | 1 | 0.92 | 0.92 | 0.30 | 0.30
13 | Misalignment positioning system | 0.1 | 0.1 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
14 | QZ ripple with SGH | 0.2 | 0.072 | 0.1 | 0.1 | Normal | 1 | 1 | 0.20 | 0.072 | 0.10 | 0.100
15 | Rotary joints |  |  | 0 | 0 | U-shaped | √2 | 1 | 0.000 | 0.000 | 0.000 | 0.000
16 | Misalignment SGH and pointing error |  |  | 0.3 | 0.3 | Exp. normal | 2 | 1 | 0.00 | 0.00 | 0.15 | 0.15
17 | Misalignment calibration system |  |  | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0
18 | Standing wave between SGH and test range antenna | 0.1 | 0.03 | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.07 | 0.02 | 0.06 | 0.06
19 | Switching uncertainty |  | 0.1 | 0.43 | 0.43 | Rectangular | √3 | 1 | 0.00 | 0.06 | 0.25 | 0.25
Combined standard uncertainty (1σ) [dB |  |  |  |  |  |  |  |  | 1.70 | 1.65 | 0.88 | 0.92
1, 0, 27, 6, 1, 127, 28, 4 1, 0, 28, 14, 1, 120, 28, 60 Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  | 3.33 | 3.24 | 1.73 | 1.81

## Slide 4: Summary for EIS MU for IAC
Proposals:
  Nokia 3.1
  Ericsson by IAC 2.04 (for 37<f<40G)
Way forward for EIS MU:
  3.1
UID | Uncertainty source | Uncertainty value | Uncertainty value | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] | Standard uncertainty ui [dB]
 |  | 24.25<f | 37<f |  |  |  | 24.25<f | 37<f
 |  | <29.5GHz | <40GHz |  |  |  | <29.5GHz | <40GHz
Stage 2: DUT measurement |  |  |  |  |  |  |  | 
1 | Positioning misalignment between the AAS BS and the reference antenna | 0.18 | 0.18 | Rectangular | 1.73 | 1 | 0.1 | 0.1
2 | Pointing misalignment between the AAS BS and the transmitting antenna | 0.18 | 0.18 | Rectangular | 1.73 | 1 | 0.1 | 0.1
3 | Quality of quiet zone | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1
4 | Polarization mismatch between the AAS BS and the transmitting antenna | 0.018 | 0.018 | Rectangular | 1.73 | 1 | 0.01 | 0.01
5 | Mutual coupling between the AAS BS and the transmitting antenna | 0 | 0 | Rectangular | 1.73 | 1 | 0 | 0
6 | Phase curvature | 0.07 | 0.07 | Gaussian | 1 | 1 | 0.07 | 0.07
7 | Uncertainty of the RF signal generator | 0.57 | 0.57 | Gaussian | 1 | 1 | 0.57 | 0.57
8 | Impedance mismatch in the transmitting chain | 0.42 | 0.42 | U-shaped | 1.41 | 1 | 0.3 | 0.3
9 | Random uncertainty | 0.18 | 0.25 | Rectangular | 1.73 | 1 | 0.1 | 0.15
Stage 1: Calibration measurement |  |  |  |  |  |  |  | 
10 | Impedance mismatch between the transmitting antenna and the network analyzer | 0.43 | 0.57 | U-shaped | 1.41 | 1 | 0.3 | 0.4
11 | Positioning and pointing misalignment between the reference antenna and the transmitting antenna | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.25 | 0.25
12 | Impedance mismatch between the reference antenna and network analyzer | 0.43 | 0.57 | U-shaped | 1.41 | 1 | 0.3 | 0.4
13 | Quality of quiet zone | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1
14 | Polarization mismatch for reference antenna | 0.018 | 0.018 | Rectangular | 1.73 | 1 | 0.01 | 0.01
15 | Mutual coupling between the reference antenna and the transmitting antenna | 0 | 0 | Rectangular | 1.73 | 1 | 0 | 0
16 | Phase curvature | 0.07 | 0.07 | Gaussian | 1 | 1 | 0.07 | 0.07
17 | Uncertainty of the network analyzer | 0.3 | 0.3 | Gaussian | 1 | 1 | 0.3 | 0.3
18 | Influence of the reference antenna feed cable | 0.18 | 0.18 | Rectangular | 1.73 | 1 | 0.1 | 0.1
19 | Reference antenna feed cable loss measurement uncertainty | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1
20 | Influence of the transmitting antenna feed cable | 0.18 | 0.18 | Rectangular | 1.73 | 1 | 0.1 | 0.1
21 | Uncertainty of the absolute gain of the reference antenna | 0.52 | 0.52 | Rectangular | 1.73 | 1 | 0.3 | 0.3
22 | Uncertainty of the absolute gain of the transmitting antenna | 0 | 0 | Rectangular | 1.73 | 1 | 0 | 0
Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.97 | 1.04
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.89 | 2.04

## Slide 5: Way forward for MU and TT for NR FR2 OTA receiver requirements
For reference sensitivity: MU_EISFR2 = 3.2 dB
For other receiver requirements (except receiver spurious emission): MUFR2 = MUFR1_6GHz + (3.2 – 1.6) dB
Where MUFR2 is the MU for FR2 receiver requirement
And MUFR1_6GHz is the MU for corresponding FR1 (4.2-6GHz) receiver requirement
For all receiver requirements: TT = MU