3GPP TSG-RAN4 Meeting #90bis	R4-1905166

# Xian, China, 8th – 12th April 2019

| CR-Form-v11.4 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 37.843 | CR | 0016 | rev | - | Current version: | 15.3.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network | X | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | CR to TR 37.843: Addition of DBCR MU evaluation and calibration & test procedure in subclause 10.2.2, 10.2.3, 10.3.2, 10.4.1, and Annex E. |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | CAICT SRTC |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R4 |  |  |  |  |  |  |  |  |  |
| Work item code: | NR_newRAT-Perf |  |  |  |  |  | Date: |  |  | 2019-04-10 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-15 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) Rel-12 (Release 12) Rel-13 (Release 13) Rel-14 (Release 14) Rel-15 (Release 15) Rel-16 (Release 16) |  |
| Reason for change: |  | At last meeting the general description for DBCR was given, with this CR the MU evalaution and calibration & test procedure are added in subclause 10.2.2, 10.2.3, 10.3.2, 10.4.1, and Annex E. |  |  |  |  |  |  |  |  |
| Summary of change: |  | The DBCR test procedure and MU table are added in TR 37.873. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  |  |  |  |  |  |  |  |  |  |
| Clauses affected: |  | subclause 10.2.2, 10.2.3, 10.3.2, 10.4.1 and Annex E. |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | x | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | x | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | x | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  |  |  |  |  |  |  |  |  |  |

<START OF CHANGE>

10.2.2.1	General

The EIRP accuracy uncertainty budget was re-assessed and adjusted slightly according as detailed in the following tables. The MU budget for the CATR and DBCR are shown as it is this MU that determines the agreed MU and TT.

10.2.2.3	DBCR

10.2.2.3.1	MU assessment

10.2.2.3.1.1	MU budget

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | E6-1 |
| 2 | RF power measurement equipment | F1 |
| 3 | Standing wave between DUT and test range antenna | E6-2 |
| 4 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | E6-19 |
| 5 | QZ ripple DUT | E6-3 |
| 19 | Miscellaneous uncertainty | E6-2 |
| 23 | MU of TE derived from conducted specification (conducted MU / 1.96) | F2 |
| Stage 1: Calibration measurement |  |  |
| 6 | Uncertainty of network analyser | F1 |
| 7 | Mismatch of receiver chain | E6-7 |
| 8 | Insertion loss variation of receiver chain | E6-8 |
| 9 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | E6-4 |
| 10 | Influence of the calibration antenna feed cable: a) Flexing cables, adapters, attenuators, connector repeatability | E6-9 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | F1 |
| 12 | Misalignment positioning system | E6-11 |
| 13 | Misalignment of calibration antenna and test range antenna | E6-1 |
| 14 | Rotary Joints | E6-12 |
| 15 | Standing wave between reference calibration antenna and test range antenna | E6-3 |
| 16 | Quality of quiet zone | E6-5 |
| 20 | Switching uncertainty | E6-15 |

10.2.2.3.1.2	MU value

| EIRP uncertainty budget |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 2 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 3 | Standing wave between DUT and test range antenna | 0.21 | 0.21 | U-shaped | √2 | 1 | 0.15 | 0.15 |
| 4 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 5 | QZ ripple with DUT | 0.20 | 0.20 | Normal | 1 | 1 | 0.20 | 0.20 |
| 23 | Frequency flatness | 0,23 | 0,23 | Normal | 1 | 1 | 0.23 | 0.23 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 6 | Network Analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.2 |
| 7 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 8 | Insertion loss variation in receiver chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.1 | 0.1 |
| 9 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 10 | Influence of the calibration antenna feed cable | 0.022 | 0.022 | U-shaped | √2 | 1 | 0.015 | 0.015 |
| 11 | SGH Calibration uncertainty | 0.50 | 0.433 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 13 | Misalignment SGH and pointing error | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 |
| 14 | Rotary joints | 0.048 | 0.048 | U-shaped | √2 | 1 | 0.034 | 0.034 |
| 15 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.06 | 0.06 |
| 16 | QZ ripple with SGH | 0.0325 | 0.0325 | Normal | 1 | 1 | 0.0325 | 0.0325 |
| 17 | Switching uncertainty | 0.26 | 0.26 | Rectangular | √3 | 1 | 0.15 | 0.15 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.587 | 0.67 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.15 | 1.30 |

10.2.3.5	DBCR

10.2.3.5.1	General

The Dyamic Beamforming Compact Range (DBCR) uses the DUT which radiates a wave front to a plane wave converter.  As showen in Figure 10.2.3.5.1-1. The probe antenna is set up of a two-dimensional array of antenna modules, connected by a special receive network. Optimized amplitude and phase settings in the receive network result in a plane wave like the DUT transmitted. As no near field to far field transformation has to be carried out, it is not required to measure a whole set of azimuth and elevation angles, just like the compact range test range.

Figure 10.2.3.5.1-1: DBCR measurement system setup for OTA E-UTRA DL RS power

Test Method limitations and scope

The maximum size of the DUT is a chamber restriction that would affect the quality of the quiet zone. For larger DUT sizes larger size chambers should be considered such that the uncertainty of the quiet zone is taken into account.

10.2.3.5.2	Calibration

The calibration measurement is done by using a reference antenna (SGH used in figure 10.2.3.5.2 -1) with known efficiency or gain values. In the calibration measurement the reference antenna is measured in the same place as the DUT, and the attenuation of the complete transmission path (C↔A, as in figure 10.2.3.5.2 -1) from the DUT to the measurement receiver is calibrated out. Figure 10.2.3.5.2 -1 presents a setup of a typical DBCR test range.

Figure 10.2.3.5.2-1: DBCR calibration system setup for OTA E-UTRA DL RS power

1)	Path loss calibration C→A:

a)	Measure SGH (or other calibrated reference antenna) reflection coefficient separately at the antenna's connector with a network analyser (or equivalent measurement equipment) to obtain ΓSGH.

b)	Measure cable loss from point C to input of SGH, call this LC↔SGH which is the equivalent of 20log|S21| from the use of a network analyser.

c)	Calculate the combined total path loss from C→A by using the following expression:

-	LSGHcal = LC,SGH + 10log(1 - |ΓSGH|2) - GSGH;

-	where  10log(1 - |ΓSGH|2)  is the compensation for SGH connector return loss, GSGH  is the known gain of the reference SGH.

2)	Connect SGH and C↔A cable.

3)	To remove polarization(s) mismatch between range antenna (labelled as feeder antenna in diagram) and SGH use positions to position the SGH in the boresight of range antenna.

4)	Measure path loss C→B with network analyzer LC→B = 20log|S21|.

5)	Calculate the test path loss compensation factor.  This is the total path loss between A↔B using the results from step 1c and 4. L = LSGHcal LSGHcal  - LC→B.

Where ΓSGH is the reflection coefficient (or mismatch) seen at the SGH connector (S11 with a network analyzer).

10.2.3.5.3	Procedure

1)	Set up AAS BS in place of SGH from calibration stage.  Align DUT with boresight of range antenna.

2)	Rotate the AAS BS to make the testing direction aligned with the direction of the receiving antenna.

3)	Configure TX branch and carrier according to maximum power requirement and test configuration.

4)	Set the AAS BS to transmit the test signal.

5)	Measure the appropriate test parameter

6)	Repeat steps 2-6 for all conformance test beam direction pairs and test conditions.

The appropriate test parameter in step 6 for DL RS power is PDL_RS the measured signal power of DL RS EIRP (in the beam peak direction), calculation of EIRP EIRPDL_RS using the following equation:

EIRPDL_RS_p(x) = PDL_RS_meas + LA→B.

And

EIRPDL_RS = EIRPDL_RS_p1 + EIRPDL_RS_p2 where the declared beam is the measured signal at port 1 (p1) and port 2 (p2).

10.2.3.5.4	MU assessment

10.2.3.5.4.1	MU Budget

As the DL RS power is an absolute measurement most of the uncertainties form the EIRP accuracy remain the same. Also it can be noted that the measured signal is a wanted signal and hence will be beam formed in the same way as the wanted signal, hence any errors which may be dependent on the beam shape will be the same as for the EIRP accuracy measurement.

Table 10.2.3.5.4.1-1: DBCR uncertainty budget format for OTA E-UTRA DL RS power measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | E6-1 |
| 2 | RF power measurement equipment | F1 |
| 3 | Standing wave between DUT and test range antenna | E6-2 |
| 4 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | E6-19 |
| 5 | QZ ripple DUT | E6-3 |
| 19 | Miscellaneous uncertainty | E6-2 |
| 23 | MU of TE derived from conducted specification (conducted MU / 1.96) | F2 |
| Stage 1: Calibration measurement |  |  |
| 6 | Uncertainty of network analyser | F1 |
| 7 | Mismatch of receiver chain | E6-7 |
| 8 | Insertion loss variation of receiver chain | E6-8 |
| 9 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | E6-4 |
| 10 | Influence of the calibration antenna feed cable: a) Flexing cables, adapters, attenuators, connector repeatability | E6-9 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | F1 |
| 12 | Misalignment positioning system | E6-11 |
| 13 | Misalignment of calibration antenna and test range antenna | E6-1 |
| 14 | Rotary Joints | E6-12 |
| 15 | Standing wave between reference calibration antenna and test range antenna | E6-3 |
| 16 | Quality of quiet zone | E6-5 |
| 20 | Switching uncertainty | E6-15 |

10.2.3.5.4.2	MU Value

Table 10.2.3.5.4.2-1: DBCR uncertainty assessment for OTA E-UTRA DL RS power measurement

| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 23 | MU of TE derived from conducted specification | 0.41 | 0.56 | Gaussian | 1 | 1 | 0.41 | 0.56 |
| 3 | Standing wave between DUT and test range antenna | 0.21 | 0.21 | U-shaped | √2 | 1 | 0.15 | 0.15 |
| 4 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 5 | QZ ripple with DUT | 0.20 | 0.20 | Normal | 1 | 1 | 0.20 | 0.20 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 6 | Network Analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 7 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 8 | Insertion loss variation in receiver chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 9 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 10 | Influence of the calibration antenna feed cable | 0.022 | 0.022 | U-shaped | √2 | 1 | 0.015 | 0.015 |
| 11 | SGH Calibration uncertainty | 0.50 | 0.433 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 13 | Misalignment SGH and pointing error | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 |
| 14 | Rotary joints | 0.048 | 0.048 | U-shaped | √2 | 1 | 0.034 | 0.034 |
| 15 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.06 | 0.06 |
| 16 | QZ ripple with SGH | 0.0325 | 0.0325 | Normal | 1 | 1 | 0.0325 | 0.0325 |
| 17 | Switching uncertainty | 0.26 | 0.26 | Rectangular | √3 | 1 | 0.15 | 0.15 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.66 | 0.80 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.30 | 1.55 |

10.2.3.6	Summary

{editors note: Last section at this level summarises MU estimates for all chambers and concludes final agreed MU}

The different OTA test chamber MU are summarised below:

Table 10.2.3.6-1: Test system specific measurement uncertainty values for the OTA E-UTRA DL RS power test

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 1.15 | 1.43 |
| Compact Antenna Test Range | 1.25 | 1.52 |
| Near Field chamber | 1.17 | 1.39 |
| Dyamic Beamforming Compact Range | 1.31 | 1.55 |
| … |  |  |
| Common maximum accepted test system uncertainty | 1.3 | 1.5 |

10.3.2	OTA sensitivity

10.3.2.1	General

The EIRP accuracy uncertainty budget was re-assessed and adjusted slightly according as detailed in the following tables. The MU budget for the CATR and DBCR are shown as it is this MU that determines the agreed MU and TT.

The OTA sensitivity MU was derived in TR 37.842 [4] as follows:

Table 10.3.2-1: Test system specific measurement uncertainty values for the EIS test

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 1.22 | 1.25 |
| Compact Antenna Test Range | 1.33 | 1.40 |
| One Dimensional Compact Range Chamber | 1.29 | 1.43 |
| Near Field Test Range | 1.24 | 1.24 |
| Dyamic Beamforming Compact Range | 1.24 | 1.42 |
| Common maximum accepted test system uncertainty | 1.3 | 1.4 |

10.3.2.2	DBCR

10.3.2.2.1	MU assessment

10.3.2.2.1.1	MU budget

| UID | Description of uncertainty contribution | Details in anex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | E6-1 |
| 2 | Standing wave between DUT and test range antenna | E6-2 |
| 3 | Signal Generator uncertainty | E |
| 4 | RF leakage & dynamic range, a) feeder package terminated b) DUT terminated | E6-19 |
| 5 | QZ ripple DUT | E6-3 |
| 6 | Miscellaneous uncertainty | E6-13 |
| Stage 1: Calibration measurement |  |  |
| 7 | Uncertainty of network analyser and/or signal generator | E |
| 8 | Mismatch of transmitter chain | E6-6 |
| 9 | Insertion loss of transmitter chain | E6-7 |
| 10 | RF leakage & dynamic range | E6-8 |
| 11 | Influence of the calibration antenna feed cable a) Flexing cables, adapters, attenuators & connector repeatability | E6-9 |
| 12 | Uncertainty of the absolute gain of calibration antenna | E |
| 13 | Uncertainty due to antenna mounting apparatus or rotary joints | E6-11 |
| 14 | Quality of quiet zone | E6-4 |
| 15 | Rotary Joints | E6-14 |
| 16 | Misalignment positioning system | E6-15 |
| 17 | Misalignment calibration antenna | E6-1 |
| 18 | Standing wave between DUT and test range antenna | E6-16 |
| 19 | Switching uncertainty | E6-17 |

10.3.2.2.1.2	MU value

| EIS uncertainty budget |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz |  | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  |
| 2 | Standing wave between DUT and test range antenna | 0.21 | 0.21 | U-shaped | √2 | 1 | 0.15 | 0.15 |  |
| 3 | RF signal generator | 0.46 | 0.46 | Gaussian | 1 | 1 | 0.46 | 0.46 |  |
| 4 | RF leakage & dynamic range, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |  |
| 5 | QZ ripple with DUT | 0.20 | 0.20 | Normal | 1 | 1 | 0.20 | 0.20 |  |
| 6 | Miscellaneous uncertainty | 0 | 0 | Normal | 1 | 1 | 0 | 0 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |
| 7 | Network Analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |  |
| 8 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |  |
| 9 | Insertion loss of transmitter chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |  |
| 10 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.00156 | 0.00156 | Normal | 1 | 1 | 0.00156 | 0.00156 |  |
| 11 | Flexing cables & connector repeatability | 0.022 | 0.022 | U-shaped | √2 | 1 | 0.015 | 0.015 |  |
| 12 | SGH Calibration uncertainty | 0.502 | 0.433 | Rectangular | √3 | 1 | 0.29 | 0.25 |  |
| 13 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  |
| 14 | QZ ripple with SGH | 0.0325 | 0.0325 | Normal | 1 | 1 | 0.0325 | 0.0325 |  |
| 15 | Rotary joints | 0.048 | 0.048 | U-shaped | √2 | 1 | 0.034 | 0.034 |  |
| 16 | Misalignment SGH and pointing error | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 |  |
| 17 | Misalignment calibration system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |  |
| 18 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.06 | 0.06 |  |
| 19 | Switching uncertainty | 0.26 | 0.26 | Rectangular | √3 | 1 | 0.15 | 0.15 |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.696 | 0.73 |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.36 | 1.43 |  |

10.4.1.5	DBCR

10.4.1.5.1	General

The Dyamic Beamforming Compact Range (DBCR) is showen in Figure 10.4.1.5.1-1. The probe antenna is set up of a two-dimensional array of antenna modules, it optimized amplitude and phase settings of the probe array which transmit the spherical wave to plane wave in the area of the DUT. So as the DUT is being tested as in a far field test range.

Figure 10.4.1.5.1-1: DBCR measurement system setup for OTA BS output power

Test Method limitations and scope

The maximum size of the DUT is a chamber restriction that would affect the quality of the quiet zone. For larger DUT sizes larger size chambers should be considered such that the uncertainty of the quiet zone is taken into account.

10.4.1.5.2	Calibration

[Note: This stage may be omitted provided calibration stage has been performed already during output power measurement]

Calibration should be carried out using the same procedure as in 10.2.3.5.2.

10.4.1.5.3	Procedure

1)	Set up AAS BS in place of SGH from calibration stage.  Align the coordinates system of the DUT with that of the test system.

2)	Set the AAS BS to transmit the test signal at the maximum power according to the appropriate test model Measure the appropriate test parameter

3)	Rotate the AAS BS to make the testing direction aligned with the direction of the receiving antenna.

4)	Measure the appropriate test parameter.

5)	Repeat the above steps 3)~5) at a number of discrete directions around the sphere according to the chosen measurement grid, see subclause 10.8.3.

6)	Calculate TRPs from power density, as shown in subclause 10.8.3.

The appropriate test parameter in step 5 is PR_AAS_desired, B, measured mean power within the desired signal channel bandwidth for each carrier arriving at the measurement equipment connector B in figure 10.4.1.5.1-1, and  calculation of power powerd using the following equitation:

PowerD = PR_AAS_desired, B+ LTX_cal, A→B

NOTE: If the test facility only supports single polarization, then measure power with the test facility's test antenna/probe polarization matched to the AAS BS, then measure and sum the power on both polarizations. If the test facility supports dual polarization then measure total power for two orthogonal polarizations and calculate total radiated transmit power as the sum over both polarizations.

10.4.1.5.4	MU assessment

10.4.1.5.4.1	MU Budget

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | E6-1 |
| 2 | RF power measurement equipment | F.1 |
| 3 | Standing wave between DUT and test range antenna | E6-3 |
| 4 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | E6-4 |
| 5 | QZ ripple DUT | E6-5 |
| 19 | Miscellaneous uncertainty | E6-14 |
|  | Test system frequency flatness | E6-16 |
| Stage 1: Calibration measurement |  |  |
| 6 | Uncertainty of network analyser | E6-6, F.1 |
| 7 | Mismatch of receiver chain | E6-7 |
| 8 | Insertion loss variation of receiver chain | E6-8 |
| 9 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | E6-4 |
| 10 | Influence of the calibration antenna feed cable: a) Flexing cables, adapters, attenuators, connector repeatability | E6-9 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | E |
| 12 | Misalignment positioning system | E6-11 |
| 13 | Misalignment of calibration antenna and test range antenna | E6-1 |
| 14 | Rotary Joints | E6-12 |
| 15 | Standing wave between reference calibration antenna and test range antenna | E6-3 |
| 16 | Quality of quiet zone | E6-5 |
| 20 | Switching uncertainty | E6-15 |

10.4.1.5.4.2	MU Value

| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.3 | 0.3 | Rectangular | √3 | 1 | 0.174 | 0.174 |
| 2 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 3 | Standing wave between DUT and test range antenna | 0.21 | 0.21 | U-shaped | √2 | 1 | 0.15 | 0.15 |
| 4 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 5 | QZ ripple with DUT | 0.20 | 0.20 | Normal | 1 | 1 | 0.20 | 0.20 |
|  | Test system frequency flatness | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 6 | Network Analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 7 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 8 | Insertion loss variation in receiver chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 9 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 10 | Influence of the calibration antenna feed cable | 0.022 | 0.022 | U-shaped | √2 | 1 | 0.015 | 0.015 |
| 11 | SGH Calibration uncertainty | 0.50 | 0.433 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 13 | Misalignment SGH and pointing error | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 |
| 14 | Rotary joints | 0.048 | 0.048 | U-shaped | √2 | 1 | 0.034 | 0.034 |
| 15 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.06 | 0.06 |
| 16 | QZ ripple with SGH | 0.0325 | 0.0325 | Normal | 1 | 1 | 0.0325 | 0.0325 |
| 17 | Switching uncertainty | 0.26 | 0.26 | Rectangular | √3 | 1 | 0.15 | 0.15 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.62 | 0.69 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.21 | 1.36 |

10.4.1.6	Summary

Table 10.4.1.6-1: Test system specific measurement uncertainty values for the OTA BS output power test

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 1,1 | 1,3 |
| Compact Antenna Test Range | 1,4 | 1,5 |
| Near Field chamber |  |  |
| Dyamic Beamforming Compact Range | 1.21 | 1.36 |
| … |  |  |
| Common maximum accepted test system uncertainty | 1.4 | 1.5 |

E.6	Dynamic Beamforming Compact Range

E6-1 Misalignment DUT/calibration antenna & pointing error

This contribution denotes uncertainty in DUT/calibration antenna alignment and DUT/calibration antenna pointing error.  In this measurement the DUT/calibration antenna is aligned to maximum, also allowing for a zero contribution for polarization mismatch uncertainty. By adjusting for maximums to align, this contribution can be a small contribution. The calibration antenna's phase centre and polarization purity changes slightly according to the frequency. Therefore, there should be some uncertainty reserved for this. To ensure that the point error is at a minimal, this contribution should be captured using the antenna pattern cut which is broadest (in the case of the DUT this would most likely be in the azimuth domain).

E6-2 Measurement Receiver: uncertainty of the absolute level

The receiving device used to measure the received signal level in the EIRP tests either as an absolute level or as a relative level. These receiving devices to name a few are spectrum analyzers, network analyzers or power meter. These devices will have an uncertainty contribution of their own; this value declared by the test gear vendor should be recorded as this uncertainty contribution. If a power meter is used then both measurement uncertainty and out of band noise is considered as part of the contribution. This uncertainty value can be found in Annex E and was a result of compromised value in order to align all test methods having this uncertainty contribution.

E6-3 Standing wave between DUT and test range antenna

This value is extracting the uncertainty value and standard deviation of gain ripple coming from standing waves between DUT and test range antenna. This value can be captured by moving the DUT towards the test range antenna as the standing waves go in and out of phase causing a ripple in measured gain.

E6-4 RF leakage (SGH connector terminated & test range antenna connector cable terminated)

This contribution denotes noise leaking in to connector and cable(s) between test range antenna and receiving equipment.  The contribution also includes the noise leakage between the connector and cable(s) between SGH/reference antenna and transmitting equipment.

E6-5	 QZ ripple DUT/calibration antenna

This is the quiet zone (QZ) ripple experienced by the AAS BS/calibration antenna during the measurement or calibration phase. The purpose of this component is to capture the contributions that the reflections from the walls, roof and floor that will add to the EIRP measurement. The sum of all these reflections from the walls, probe array (if applicable), roof and floor will give the overall value for the QZ ripple. In other words, the uncertainty component from the wall will not be separated from the roof or the floor. The purpose of this uncertainty component is to capture the overall reflections from the chamber walls experienced by the AAS BS/calibration antenna. To capture the full effect of the QZ ripple a distance of  1λ should be measured from each of the AAS BS/calibration antenna physical aperture edges, i.e. total QZ distance = physical aperture length +2 λ, to ensure the full volume of the QZ is captured in the uncertainty measurement.

E6-6 Uncertainty of network analyser

a)	drift (temp, oscillators, filters, etc.) start-to-end time of  measurements.

This uncertainty includes all the uncertainties involved in the S21 measurement (including drift and frequency flatness) with a network analyzer, and will be calculated from the manufacturer's data in logs. This uncertainty also includes analyzer uncertainty for multi-polarization (2 or more ports) measured simultaneously. This uncertainty value can be found in Annex E and was a result of compromised value in order to align all test methods having this uncertainty contribution.

E6-7 Mismatch of receiver chain (i.e. between receiving antenna and measurement receiver)

This uncertainty is the residual uncertainty contribution coming from multiple reflections between the receiving antenna and the test receiver equipment. This value can be captured through measurement by measuring the S11 towards the receive antenna and also towards the test receiver. The mismatch between the antenna reflection and the receiver reflection can also be calculated. If the same cable is used for calibration Stage 1, this can be considered systematic and negligible.

E6-8 Insertion loss of receiver chain

This uncertainty is the residual uncertainty contribution coming from introducing an antenna at the end of the cable.  If this cable does not change/move between the calibration Stage 1 and the measurement Stage 2, the uncertainty is assumed to be systematic and negligible during the measurement stage. Alternatively, the insertion loss can also be calculated by taking the measurement of the cable where port 2 is the end of the cable connected to the AAS BS or calibration antenna.

IL = -20log10|S21| dB

E6-9 Influence of the calibration antenna feed cable

a)	Flexing cables, adapters, attenuators, extra pathloss cable & connector repeatability.

During the calibration phase this cable is used to feed the calibration antenna and any influence it may have upon the measurements is captured. This is assessed by repeated measurements while flexing the cables and rotary joints. The largest difference between the results is recorded as the uncertainty.

E6-10 Uncertainty of the absolute gain of the calibration antenna

This uncertainty consists of the uncertainty of the gain value associated with the gain value denoted from the antenna calibration. This uncertainty value can be found in Annex E and was a result of compromised value in order to align all test methods having this uncertainty contribution.

E6-11 Misalignment positioning system

This contribution originates from uncertainty in sliding position and turn table angle accuracy. If the calibration antenna is aligned to maximum this contribution can be considered negligible and therefore set to zero.

E6-12 Rotary Joints

If applicable the contribution of this uncertainty the accuracy in changing from azimuth to vertical measurements.

E6-14 Miscellaneous Uncertainty

The term 'miscellaneous uncertainty' is used to define all the unknown, unquantifiable, etc. uncertainties associated with EIRP measurements. This term should include truly random effects as well as systematic uncertainties, such as that arising from dissimilarity between the patterns of the reference antenna (SGH) and the DUT.

E6-15 Switching Uncertainty

The purpose of the switching unit is to switch electromechanically different RF path to different measurement instruments of different measurement modes. The electromechanical switching clearly reduces the errors arising from manual switching work. Switching is also used to measure the path loss values of each polarization component.  Even though the electromechanical switching is preferable during path loss and antenna performance measurements, some minor uncertainties can occur when the switch states are programmed to change their polarity.

E6-16 Frequency flatness

This uncertainty contribution to account for the frequency interpolation error caused by a finite frequency resolution during the calibration stage.

E6-17 Quality of quiet zone (extreme)

This contribution is related to the ambient quality of the quiet zone (C1-3, 3GPP TR 37.842 [4]) which originates from a reflectivity level of an anechoic chamber. The reflectivity level is determined from the average standard deviation of the electric field in the quiet zone. As the environmental enclosure is larger than the DUT and the material of the environmental chamber may cause some reflection and refraction the quite zone flatness will be effected. The quality of the quiet zone for the extreme test is therefore larger tan that for the ambient due to the environmental enclosures effect.

E6-18 Wet radome loss variation

The environmental chamber radome will be an ineffective thermal isolator and will have extreme temperatures on the inside and the OTA chamber ambient temperature on the outside. In such conditions condensation is inevitable. This uncertainty is due to the variation in the radome loss due to condensation on the environmental chamber radome.

E6-19 Radome loss variation

The environmental chamber radome will affect the path between the DUT and the test antenna due to both its insertion loss and also reflection and refraction from the materials surface. The loss is dependent on the material as well as its proximity to the DUT. The uncertainty is the residual uncertainly of the total loss after calibration

E6-20 Change in absorber behaviour

The environmental chamber will contain RF absorptive material to prevent reflections within the chamber. This RF absorptive material will be subjected to the extreme temperatures inside the environmental chamber and hence its properties will change. This uncertainty accounts for the effect of that change in behavior.

<END OF CHANGE>